Development of automated drug delivery system for an ultra-short-acting β-blocker, landiolol, to stably reduce myocardial oxygen consumption without inducing circulatory collapse in a canine heart failure model

Supplemental Material

Automated drug delivery system

We developed an automated drug delivery system that administers landiolol, dextran and furosemide based on negative feedback control of hemodynamics. The scheme of the system is described in the main manuscript. First, the user sets the target values of arterial pressure (AP*) and left atrial pressure (PLA*). The system continuously measures the subject’s arterial pressure (AP), left atrial pressure (PLA), right atrial pressure (PRA), and cardiac output (CO). Based on the circulatory equilibrium framework [1–3], the system calculates the subject’s hemodynamic parameters comprising systemic vascular resistance (R), stressed blood volume (V) and the slopes of the Flank-Starling curves for the left ventricle (SL) and right ventricle (SR) from the measured hemodynamic variables using the following equations.

R=AP-PRACO#1

V=0.129×CO+19.61×PRA+3.49×PLA#2

SL=COlnPLA-2.03+0.8#3

SR=COlnPRA-2.13+1.9#4

From the hemodynamic parameters of AP*, and PLA*, the system calculates target SL (SL*) and target V (V*) by the algorithm shown below. Assuming that R and the ratio of SL to SR are constant, we can describe target CO (CO*), target PRA (PRA*), target V (V*), target SL (SL*), target SR (SR*) by the following equations.

AP*=CO*×R+PRA*#5

V*=0.129×CO*+19.61×PRA*+3.49×PLA*#6

CO*=SL*lnPLA*-2.03+0.8#7

CO*=SR*lnPRA*-2.13+1.9#8

SL*SR*=SLSR#9

By solving equations (5)－(9), the system determines CO*, PRA*, V*, SL*, and SR* from SL, SR, R, AP*, and PLA*. The system repeats these calculations to determine SL* and V* every 10 s.

To minimize the difference between SL* and SL, a proportional-integral (PI) feedback controller determines the infusion rate of landiolol by the following equations.

ut=Kpet+Ki0teτdτ#10

where t is the time increment, u(t) is the infusion rate of landiolol, e(t) is the error signal (SL*-SL), Kp is the proportional gain, and Ki is the integral gain. We determine the parameters of the PI controller (Kp = -1.5 (unitless), Ki = 0.01 s-1) based on the step response of landiolol (data not shown). When u(t) is negative, the system stops landiolol infusion and keeps the integral term constant until u(t) becomes positive.

To minimize the difference between V* and V, a nonlinear feedback controller determines the infusion rate of dextran or injection of furosemide by the following “if-then” rule.

If V* - V > 1 ml·kg-1, then infuse dextran 15 ml·kg-1

If V* - V < -2 ml·kg-1, then inject furosemide 10 mg

After furosemide injection, the system temporarily suspends the injection of furosemide for 10 min.

References

1. 	Uemura K, Kamiya A, Hidaka I, et al (2006) Automated drug delivery system to control systemic arterial pressure, cardiac output, and left heart filling pressure in acute decompensated heart failure. J Appl Physiol 100:1278–86. https://doi.org/10.1152/japplphysiol.01206.2005

2. 	Uemura K, Sugimachi M, Kawada T, et al (2004) A novel framework of circulatory equilibrium. Am J Physiol Heart Circ Physiol 286:H2376–H2385. https://doi.org/10.1152/ajpheart.00654.2003

3. 	Uemura K, Kawada T, Kamiya A, et al (2005) Prediction of circulatory equilibrium in response to changes in stressed blood volume. Am J Physiol Heart Circ Physiol 289:H301–H307. https://doi.org/10.1152/ajpheart.01237.2004