Table 1	Arsenic concentration range and arsenic (as As) detection by FTKs

| Arsenic (as As) in µg/L |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Concentration Points | AsKit-1 |  |  | AsKit-2 |  |  | AsKit-3 |  |  | AsKit-4 |  |  |
|  | DM | M1 | M2 | DM | M1 | M2 | DM | M1 | M2 | DM | M1 | M2 |
| Blank | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |
| Blank | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |
| Blank | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |
| Blank | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND | ND |
| 5 | CNS | CNS | CNS | CNS | CNS | CNS | ~10 | ~10 | ~10 | >5 & <10 | >5 & <10 | >5 & <10 |
| 10 | ~10 | ~10 | >10 & <20 | >10 & <25 | >10 & <25 | >10 & <25 | ~10 | >10 & <25 | >10 & <25 | >10 & <20 | >10 & <20 | >10 & <20 |
| 20 | ~20 | ~20 | 20 | CNS | CNS | CNS | CNS | CNS | CNS | ~20 | ~20 | ~20 |
| 25 | CNS | CNS | CNS | ~25 | ~25 | ~25 | ~25 | ~25 | ~25 | CNS | CNS | CNS |
| 30 | ~30 | ~30 | ~30 | CNS | CNS | CNS | CNS | CNS | CNS | ~30 | ~30 | ~30 |
| 40 | ~40 | ~40 | ~40 | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS |
| 50 | ~50 | ~50 | ~50 | ~50 | ~50 | ~50 | ~50 | ~50 | ~50 | ~50 | ~50 | ~50 |
| 60 | ~60 | ~60 | ~60 | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS |
| 70 | ~70 | ~70 | ~70 | CNS | CNS | CNS | CNS | CNS | CNS | ~70 | ~70 | ~70 |
| 80 | ~80 | ~80 | ~80 | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS |
| 90 | ~90 | ~90 | ~90 | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS | CNS |
| 100 | ~100 | ~100 | ~100 | ~100 | ~100 | ~100 | ~100 | ~100 | ~100 | ~100 | ~100 | 100 |
| 250 | CNS | CNS | CNS | CNS | CNS | CNS | ~250 | ~250 | ~250 | CNS | CNS | CNS |
| 500 | CNS | CNS | CNS | ~500 | ~500 | ~500 | ~500 | ~500 | ~500 | CNS | CNS | CNS |

ND	: Not detected

CNS	: Concentration point not specified in the FTK

DM	: Demineralized water

M1	: Matrix-1 (Water with TDS 550 mg/L)

M2	: Matrix-2 (Water with TDS 1200 mg/L)

Table 2 	Sensitivity, specificity and overall testing efficiency of the FTK evaluation in the laboratory

| FTK Code | Matrix | No. of observations (n) | (a) | (b) | (c) | (d) | SEArsenic | SPArsenic | OTEArsenic |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| AsKit-1 | DM Water | 14 | 10 | 0 | 0 | 4 | 100.0 | 100.0 | 100.0 |
|  | M1 water | 14 | 10 | 0 | 0 | 4 | 100.0 | 100.0 | 100.0 |
|  | M2 water | 14 | 9 | 1 | 0 | 4 | 100.0 | 80.0 | 92.9 |
|  | Matrixpooled | 42 | 29 | 1 | 0 | 12 | 100.0 | 92.3 | 97.6 |
| AsKit-2 | DM Water | 9 | 4 | 1 | 0 | 4 | 100.0 | 80.0 | 88.9 |
|  | M1 water | 9 | 4 | 1 | 0 | 4 | 100.0 | 80.0 | 88.9 |
|  | M2 water | 9 | 4 | 1 | 0 | 4 | 100.0 | 80.0 | 88.9 |
|  | Matrixpooled | 27 | 12 | 3 | 0 | 12 | 100.0 | 80.0 | 88.9 |
| AsKit-3 | DM Water | 11 | 6 | 1 | 0 | 4 | 100.0 | 80.0 | 90.9 |
|  | M1 water | 11 | 5 | 2 | 0 | 4 | 100.0 | 66.7 | 81.8 |
|  | M2 water | 11 | 5 | 2 | 0 | 4 | 100.0 | 66.7 | 81.8 |
|  | Matrixpooled | 33 | 16 | 5 | 0 | 12 | 100.0 | 70.6 | 84.8 |
| AsKit-4 | DM Water | 11 | 5 | 2 | 0 | 4 | 100.0 | 66.7 | 81.8 |
|  | M1 water | 11 | 5 | 2 | 0 | 4 | 100.0 | 66.7 | 81.8 |
|  | M2 water | 11 | 5 | 2 | 0 | 4 | 100.0 | 66.7 | 81.8 |
|  | Matrixpooled | 33 | 15 | 6 | 0 | 12 | 100.0 | 66.7 | 81.8 |

- : true positive; (b): false positive; (c): false negative; (d): true negative

Table 3	Indicators for FTK Adequacy, Safety and Information

| Indicators for FTK Adequacy, Safety and Information |  | FTK Code |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | AsKit-1 | AsKit-2 | AsKit-3 | AsKit-4 |
| FTK Adequacy (KAdq) indicators |  |  |  |  |  |
| (A1) Ease of handling | Observation | C | C | C | NC |
|  | Weight | 2 | 2 | 2 | 0 |
| (A2) Adequacy of consumables | Observation | C | C | C | C |
|  | Weight | 2 | 2 | 2 | 2 |
| (A3) Portability of FTK | Observation | C | C | C | C |
|  | Weight | 2 | 2 | 2 | 2 |
| (A4) Adequacy of main apparatus, accessories and spares | Observation | C | C | C | C |
|  | Weight | 2 | 2 | 2 | 2 |
| (A5) Adequacy with respect to space provided for placement of chemicals, apparatus etc. | Observation | C | C | C | NC |
|  | Weight | 2 | 2 | 2 | 0 |
| Adequacy weight, KAdq =i=5nAi |  | 10 | 10 | 10 | 6 |
| FTK Safety (KSafe) indicators |  |  |  |  |  |
| (S1) No Hazard involved in using kit | Observation | C | C | C | NC |
|  | Weight | 2 | 2 | 2 | 0 |
| (S2) No Leakage of liquid chemical (s) | Observation | C | C | C | C |
|  | Weight | 2 | 2 | 2 | 2 |
| (S3) No Leakage of gas (if test is based on gas formation) | Observation | C | C | C | C |
|  | Weight | 2 | 2 | 2 | 2 |
| (S4) Good Packaging of chemicals | Observation | C | C | C | C |
|  | Weight | 2 | 2 | 2 | 2 |
| (S5) Kit does not have any fragile item | Observation | C | C | C | NC |
|  | Weight | 2 | 2 | 2 | 0 |
| Safety weight, KSafe =i=5nSi |  | 10 | 10 | 10 | 6 |
| FTK Information (KInf) indicators |  |  |  |  |  |
| (I1) Instructions for operations are adequate and simple. | Observation | C | C | C | C |
|  | Weight | 2 | 2 | 2 | 2 |
| (I2) Precautions for handling corrosive and reactive chemicals are given. | Observation | C | C | C | NC |
|  | Weight | 2 | 2 | 2 | 0 |
| (I3) Instructions for disposal of used material and wastes are given | Observation | NC | NC | NC | NC |
|  | Weight | 0 | 0 | 0 | 0 |
| (I4) Specifications of chemicals in “public domain” are available. | Observation | NC | NC | NC | NC |
|  | Weight | 0 | 0 | 0 | 0 |
| (I5) Instructions for remedial measures are given in case the chemical comes in contact with skin or any other injury occurred during the use of kit. | Observation | NC | NC | NC | NC |
|  | Weight | 0 | 0 | 0 | 0 |
| Information weight, KInf =i=5nIi |  | 4 | 4 | 4 | 2 |

Compliance (C): ‘2’ points; 	Partial Compliance (PC): ‘1’ point; 	Non-compliance (NC): ‘0’ point

Table 4	Statistical weights (individual parameters and cumulative parameters)

| Parameters |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | AsKit-1 | AsKit-2 | AsKit-3 | AsKit-4 |
| Individual parameter weight (1-10 scale) |  |  |  |  |  |
| Technical efficiency | OTEArsenic | 97.6 | 88.9 | 84.8 | 81.8 |
|  | Technical efficiency weight (KTe) | 9.76 | 8.89 | 8.48 | 8.18 |
| Adequacy weight (KAdq) |  | 10.0 | 10.0 | 10.0 | 6.0 |
| Safety weight (KSafe) |  | 10.0 | 10.0 | 10.0 | 6.0 |
| Information weight (KInf) |  | 4.0 | 4.0 | 4.0 | 2 |
| Individual parameter weight with the application of importance factor |  |  |  |  |  |
| Technical efficiency weight (6KTe) |  | 58.56 | 53.34 | 50.88 | 49.08 |
| Adequacy weight (KAdq) |  | 10.0 | 10.0 | 10.0 | 6.0 |
| Safety weight (2KSafe) |  | 20.0 | 20.0 | 20.0 | 12.0 |
| Information weight (KInf) |  | 4.0 | 4.0 | 4.0 | 2 |
| Cumulative parameter weight (CPW) (1-100 scale) |  | 92.56 | 87.34 | 84.88 | 69.08 |

Table 5	Arsenic test results: AsKit-1 in the field and AAS-HG in the laboratory

| Arsenic (as As) Concentration Range, µg/L |  | Number of sample |  | AsKit-1 results as compared to ±5% of AAS-HG Results |  |  |  |  | AsKit-1 results as compared to ±10% of AAS-HG Results |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | AsKit-1 | AAS-HG |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  | Positive |  | Negative |  | Total | Positive |  | Negative |  | Total |
|  |  |  |  | True | False | False | True |  | True | False | False | True |  |
| <10 | <5 | 28 | 27 | 0 | 0 | 0 | 28 | 28 | 0 | 0 | 0 | 28 | 28 |
|  | ≥5<10 |  | 3 |  |  |  |  |  |  |  |  |  |  |
| ≥10<20 |  | 19 | 21 | 11 | 3 | 5 | 0 | 19 | 18 | 0 | 1 | 0 | 19 |
| ≥20<30 |  | 6 | 2 | 0 | 4 | 2 | 0 | 6 | 5 | 1 | 0 | 0 | 6 |
| ≥30<40 |  | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| ≥40<50 |  | 4 | 4 | 0 | 1 | 3 | 0 | 4 | 4 | 0 | 0 | 0 | 4 |
| ≥50<60 |  | 6 | 7 | 3 | 2 | 1 | 0 | 6 | 4 | 1 | 1 | 0 | 6 |
| ≥60<70 |  | 4 | 3 | 0 | 3 | 1 | 0 | 4 | 4 | 0 | 0 | 0 | 4 |
| ≥70<80 |  | 3 | 0 | 0 | 1 | 2 | 0 | 3 | 0 | 1 | 2 | 0 | 3 |
| ≥80<90 |  | 4 | 9 | 0 | 1 | 3 | 0 | 4 | 3 | 0 | 1 | 0 | 4 |
| ≥90<100 |  | 5 | 3 | 2 | 3 | 0 | 0 | 5 | 5 | 0 | 0 | 0 | 5 |
| ≥100<110 |  | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 |
|  |  | 80 | 80 | 16 | 19 | 17 | 28 | 80 | 44 | 3 | 5 | 28 | 80 |
| SEArsenic |  |  |  | 48.5% |  |  |  |  | 89.8% |  |  |  |  |
| SPArsenic |  |  |  | 59.6% |  |  |  |  | 90.3% |  |  |  |  |
| OTEArsenic |  |  |  | 55.0% |  |  |  |  | 90.0% |  |  |  |  |

Table 6	Goodness-of-fit Chi (χ2) square model inputs

| Concentration (µg/L) Range | AsKit-1 results (No. of samples) |  | AAS-HG Results (No. of samples) |  |
| --- | --- | --- | --- | --- |
|  | Observed Frequency (O) | Expected Value (E) | Observed Frequency (O) | Expected Value (E) |
| <10 | 28 | 28.82 | 30 | 29.18 |
| ≥10<20 | 19 | 19.87 | 21 | 20.13 |
| ≥20<50 | 10 | 8.45 | 7 | 8.55 |
| ≥50<100 | 22 | 21.86 | 22 | 22.14 |

Number of categories (k) = 8

Sample size (n) = 159

Degree of freedom (df)=7

Chi (χ2) Square =  i=1n(Oi-EiEi = 0.6891

Phi (φ) effect = χ2n = 0.06583