2nd  round of review:

Reviewer 1:

The authors addressed all my concerns but I recommend publication before having this manuscript edited by  a native speaker.

Response:   Reviewer  has raised several important points, including the need for method comparisons and clarification of the core concept of our NGS-based approach. We appreciate Reviewer 1's insightful feedback and acceptance of our approach.

Now as per your suggestion  we have  edited the manuscript  to further improve clarity and readability with correct English

Reviewer 3:

Point 1: In the revised manuscript, some issues have been addressed. However, overall, the core content of the paper focuses on using NGS to elucidate the molecular basis of 14 rare thalassemia cases that conventional testing could not conclusively diagnose. Although this method has demonstrated strong detection capability in these limited 14 samples, it has not been validated with a larger and more diverse range of mutation types, particularly in the detection of CNVs. Therefore, the authors' claims regarding the superiority and reliability of this method seem to lack support from clinical practice, which is difficult to accept. Would it be possible to consider focusing the paper's theme on the hematological and molecular characteristic of these 14 complex/rare thalassemia cases?

Response:   I appreciate the reviewer has acknowledged that our developed method and algorithm demonstrate strong detection capability.  In addition, all detected CNVs have been further validated using the split-read method, as suggested by Reviewer 1. Further, the mutation deetd in this NGS method and panel subsequently submitted to the NCBI ClinVar. The ClinVar Team independently  validated and published with accession number and given blow:

As per  the suggestion of the reviewers for  we have modified the discussion line no 318 to 326.

Table S2: Identified different length of deletion in different samples as find out trough split read and subsequently by depth plot analysis, which were finally submitted and published in the clinvar database.

| Sample Id | Identified CNV | Gene involved | Probable Coordinate [GRCh38] | ClinVar, accession number |
| --- | --- | --- | --- | --- |
| S7, S12 | HBD-HBB del | HBB; HBD | Chr11:5235201-5225911 | VCV003255363.1 |
| S8 | HBBP1-HBB del | HBB; HBD | Chr11:5225911-5243699 | VCV003255364.1 |
| S14, S9 | HBG1-HBB del | HBB; HBD; HBG1 | Chr11:5223699-5250078 | VCV003255362.1 |
| S12 | HBE1-HBB del | HBB; HBD; HBG1; HBG2; HBE1 | Chr11:5235201-5269140 | VCV003255365.1 |
| S4, S10, S11 | Entire target region del | HBB; HBD; HBG1; HBG2; HBE1, OR51B4 | Chr11:5223780-5304186 | VCV003255366.1 |

Point 2. Additionally, in the supplementary materials, Table A indicates that the Sanger sequencing result for S7 is homozygous for Hb E, which is inconsistent with the HPLC results. However, based on the data provided in Table 1 (43% Hb E and 5.2% Hb A2), this phenotype is consistent with the βE/βE genotype. Although the NGS results suggest that the actual genotype is HBD-HBB del//βE , it indicates that the phenotypes of these two genotypes are similar.

Response: It appears that our method has accurately detected the Hb E/β genotype as HBD-HBB del//βE, consistent with the phenotype showing 43% Hb E and 5.2% Hb A2. I would like to clarify that for the βE/βE (Hb E homozygous) genotype, Hb E levels  of HPLC would be  above 80%. Here HbE level is 43% , so it would not be  βE/βE (Hb E homozygous) , instead it would be Hb E/β genotype. Thus   It seems the reviewer may have misinterpreted the HPLC values, which could have led to confusion regarding the phenotype associated with Hb E/β versus βE/βE genotypes. For HPLC value of HbE for Hb E homozygous  ( βE/βE genotype) and  Hb E/β genotype, I kindly refer the reviewer to the following references

1.Leroy, P., & Giordano, P.C. (2005). "Hemoglobin E Trait and Homozygous Hemoglobin E in Immigrants from Southeast Asia: A Comparison of Clinical and Hematologic Findings" in Journal of Pediatric Hematology/Oncology, Volume 27, Issue 5, pp. 253-258.

2."Hemoglobin E Disorders" in Hematology: Basic Principles and Practice
This widely-used hematology textbook discusses Hb E disorders, including typical Hb E and Hb A2 percentages in Hb E/β-thalassemia and Hb E homozygous (HbE/HbE) genotypes.

3.Weatherall, D.J., & Clegg, J.B. (2001). The Thalassaemia Syndromes (4th Edition)

1st round of review
Reviewer 1

In this study, Pal et al. presented a targeted next-generation sequencing panel covering all coding genes and cis-elements that potentially caused the β-hemoglobinopathies and tested it by sequencing 14 individuals with different genotypes in HBB. They found that this panel is capable of detecting multiple disease-causing point mutations and large deletions.

Suggestions to the authors:
Major:

1. The major defect of this NGS panel is the lack of capturing α-globin genes. The authors should at least discuss the prevalence of α-thalassemia mutations in their local population, which is crucial for the practicability of this panel in thalassemia screening. If α-thalassemia mutations, especially the duplications of the α-globin genes highly occurred in the population, then the application of this panel would lead to a leaky scanning of those thalassemia patients compound heterozygous in β-thalassemia mutations and α-duplications.

Response: Thank you for your valuable suggestion. The primary aim of this study to develop a NGS based genomic panel and methodology for detection of any SNVs and CNVs only in beta globin gene cluster. The developed technology can lower the turn around time (TAT) or test time, utilizing most common and tested short ready based NGS platform by – Illumina.

In Indian context prevalence of alpha Thalassemia, loss of 3 alpha allele or HbH disease is very rare in this population. The alpha carrier or loss of 2 allele remain clinical inert for the presence of 4 alpha allele. A study by Nadkarni A. et al. (2008) [PMID: 18447636, DOI: 10.1089/gte.2007.0080] from India, revealed that 10.5% of individual carried a heterozygous alpha-globin gene deletion, while 2.31% had a deletion of two alpha genes. Additionally, 1.1% showed alpha gene triplication. Thus, Thalassemia due to Alpha mutation or combination with beta rarely present as case. Thus, in this population, detection of mutation of HBB gene cluster is very crucial.

As per your suggestion, we include a discussion on the prevalence and significance of α-thalassemia mutations in our local population in the revised manuscript [main document , line number: 345 to 355, page number 13], and also change the tile of the article from “A next generation sequencing ………… rapid   diagnosis of thalassemia”. to “A next generation sequencing ……… for rapid   diagnosis of β-Thalassemia”.

In view of the importance of alpha mutations detection through single gene panel, we have planned to develop a long read based combined beta and alpha gene panel using Oxford Nanopore technology where we include both HBB and HBA genes in a single panel.

2. The authors selected 14 individuals for NGS primarily from 600 individuals, but it is not clearly written regarding the genotypes of these 600 individuals and how the rest samples are filtered. In addition, the authors should discuss the rate of consistency between the hematological indices and their thalassaemic genotypes among these 600 individuals. This is important as it could provide population evidence on why the NGS-based carrier screening is advantageous compared to conventional phenotype-to-genotype screening pipeline.

Response: Thank you for your thoughtful comments and suggestions. To clarify the selection of 14 individuals, out of 600 beta thalassemia patients, a step-by-step subject selection process has been given in the method section (Figure 1) in the revised manuscript.

As per your suggestion, we have included two tables S5 and table S6 in the supplementary data, providing the demographic and clinicopathological condition of different beta thalassemia genotype categories (B0/B0, B+/B+, and B0/B+) enrolled in this study.

3. The authors claimed in their abstract that “Additionally, 169 rare deletions and 11 fusions mutations from the global population were confirmed, with their coordinates obtained from the HbVar database”. This is an unprecise statement as the authors only presented the NGS results in 14 selected samples without testing most of the rare mutations and fusions recorded in the database, though I understand the authors suppose the panel theoretically covers the regions of these known variants. I strongly recommend the authors reorganize this part.

Response: I appreciate your suggestion regarding this issue, and positive criticism. Indeed, the evaluation is based purely on the theoretical coverage of the coordinates. Regarding the detection of copy number variants (CNVs), in HBVar database, we agree that theoretical coverage needs to be substantiated by actual physical method of testing. For this purpose, we need actual blood samples of diverse population and could better way to testify the method. But, we only able to get Asian pupation in this locality, and in this limited scope of the study, thus theoretically validate the coordinate of CNVs in detecting these variations

4. More detailed clinical characteristics of the 14 samples subjected to NGS should be presented in the manuscript.

Response: A table showing the detail haematological data along with clinical characteristic included in the revised manuscript as Table 1.

5. One of the limitations of this study is that there is only 1 control sample recruited, which is probably incapable of evaluating the accuracy of CNV detection, especially the volatility of the read depth in the regions of interest. The authors should include more control samples to optimize their normalization procedures regarding the CNV detection.

Response: Thank you for highlighting this important limitation regarding the use of control samples for evaluating CNV detection accuracy. We acknowledge that having only one control sample is insufficient to thoroughly evaluate the accuracy of CNV detection and assess the volatility of read depth in the regions of interest. Although the sequencing was performed at 100x coverage that generate sufficient amount of data for depth based CNV calling. To accurately identify CNV we have used two different methods for CNV analysis 1. CNVkit & 2. Depth plot analysis.

We are actively seeking additional funding and resources to expand our beta thalassemia as well as control sample cohort and conduct more comprehensive analyses.

6. The authors stated in the “Study setting” session of Materials and Methods that “Interestingly, one of the proband's parents exhibited a completely normal genotype via Sanger sequencing, resulting in a mismatched mutation inheritance pattern (S7-S11)”. This is misleading as S7-S11 are all patients with large deletions in HBB, which means that only the intact allele has been sequenced by Sanger. Therefore, it cannot be defined as “normal genotype”. Moreover, the carriers of large deletions in HBB (HPFH) often showed high levels of HbF, which can be routinely screened by their haematological indices. The authors should also provide more hematological phenotypes to support their conclusions.

Response: Thank you for your valuable comment. The manuscript has been revised according to your suggestion. The statement is revised ( line no 108-109)  and the haematological parameters have been provided in the Table no. 1.

7. In Table 2, the column of “allele 2” is filled with “N” or “NA”. According to the description of the authors, they are carriers of β-thalassemia. Therefore, I suppose this should be “Normal” instead of “NA”

Response: Changed according to your suggestion in the Table no 3 in the  revised manuscript.

8. Again, in Table 2, the CNVs identified in the 14 samples should be highlighted with their HGVS nomenclature and their population frequencies (Reported CNVs or novel ones). In addition, I recommend that authors extract the split reads from the raw sequencing data to present the precise breakpoints of these CNVs.

Response: Thank you for your valuable suggestion regarding the extraction of split reads to present precise breakpoints of CNVs. As per your suggestion we have performed Split read analysis to identify the breakpoint within the target region.  For split read analysis we took FASTQ file, aligned through BWA MEM aligner and convert into SAM file and converted to sorted BAM file and indexed. Finally, each BAM file was visualized through IGV.  IGV view displaying paired orientation of large insert size marked by red lines, indicating structural variation for getting split red profile.

Accordingly for every sample we have compared the split read profile along with depth plot analysis to confirm the presence of large deletion and probable break points within the target region has been submitted to the ClinVar database and published with accession number, presented in the revised supplementary file table: S2,

For your ready information, A figure panel is provided below to show the split reads from the raw sequencing data and compared depth plot of same sample with ClinVar accession number VCV003255363.1

Figure 1: HBD-HBB genomic deletion [ approximately 9.2 KB] identification by split read method in sample S7 & S13 visualize in IGV. A. Representative depth plot showing HBD-HBB genomic deletion in sample S7 and S13. Upper panel showing variant position and their zygosity. Continuous stretch of homozygous variants in HBD-HBB genomic region are indicated in black box. B. IGV view displaying paired orientation of large insert size marked by red lines. Indicating structural variation. C. Grouping of abnormal large insertation in pair orientation, Low coverage region indicating heterozygous deletion indicated in black arrow, the deletion coordinate identified by brake point indicated in box. D. Squished view in IGV illustrating coverage across the target region.

I appreciate your suggestion for population frequencies and HGVS nomenclature, for the identified variants, but present study was not aiming for variant discovery type. As it is well established that  beta globin gene cluster possess several large deletion ( even more than 200KB)  and fusion with different break points, and accordingly for every cases it is not possible to compare such large  deletion with other earlier reported deletions. Thus in most of the data bases like Ithanet ( for Hemoglobinopathy) or ClinVar, for such CNVs, lack of population frequencies and even HGVS nomenclature. Accordingly, we have submitted ClinVar database without HGVS nomenclatures as exact break point as base resolution may not be appropriate.

9. The 14 samples with different genotypes, especially those with deletions in HBB globin gene clusters, conventional approaches like gap-PCR and MLPA should be shown and the breakpoints should be validated by Sanger sequencing to evaluate the accuracy of the NGS panel in detecting the CNVs within the regions of interest.

Response: I appreciate your suggestion. In this approach, the main objective was not to detect deletions at the base-pair resolution; instead, we are identifying both gross gene deletions and rare pathogenic point mutations in a single assay. Detecting every single CNV with unknown breakpoints is labour-intensive and time-consuming. Therefore, the genomic panel designed to covers the gene region and its regulatory LCR crucial for hemoglobin synthesis. To confirm CNVs, we employed two methods: CNVkit and depth plot-based analysis. Additionally, we confirmed the deletions visually using IGV.

Minor:

1. Table S2, Sample S8: HbA2%: 4.3%； HbF%: 101.3%. This does not make sense.
Response: We appreciate your diligence in reviewing our data. There was an inadvertent error, and     relevant data has been updated in the Table 1 of revised manuscript

2. Figure 1 is hard to read and it should be edited with higher resolution.

Response:  The figure 1 has been updated to higher resolution file in the revised figure

3. In Figure 4, the diagram should be marked with specific sample IDs.

Response: As per your suggestion, type of deletion has been marked in the figure and corresponding sample number has been mentioned in the Figure legend.

Reviewer:2

This study aimed to develop and validate a targeted next-generation-sequencing (NGS) panel along with a data analysis algorithm capable of detecting single nucleotide variants (SNVs) and copy number variations (CNVs) within the beta-globin gene cluster. The author devised a targeted NGS panel spanning an 80.4 kb region on chromosome 11, encompassing beta-globin gene cluster and 5' locus control region (LCR), and also developed an advance data analysis algorithm consisting with variant calling and depth plot analysis that facilitates simultaneous detection of SNVs and CNVs in a single run. This study makes a lot of sense.
The main problems are as follows:

1. In the first paragraph, the author describes that "Although HPLC is the gold standard for carrier screening, it has some limitations. However, HPLC is not a gold standard for carrier screening. CE performed better than HPLC.

Response: Thank you for your comment, The statement has been updated as per suggestion.

2. Table 1 and Table 2 should use a three-line table.

Response: Table 1, 2 and 3 updated to three-line table in the Revised manuscript.

3. About table 2. How do you know if two variants are cis structures when they are in HBB and HBD, e.g. S5 and S6? Although whether it is a cis structure or not at the individual level has no effect on the clinical phenotype, it still influences genetic counselling.

Response: I appreciate your concerns, and clinical difference of Cis and Tranas-condition of the HBB and HBD mutation, We have recruited S5 and S6 as they are carrier as per the pedigree of the index case (i.e both of them are respective parents of the index case), but showing ambiguous HPLC with low A2 or borderline A2 and elevated  HBF with low Hb. Thus, after NGS, interestingly we have found both of them possess mutation in the HBB and HBD genes. We could have checked whether this HBB and HBD mutations are in cis condition or trans-condition through their respective prenatal genotype, i.e grandparents of the index cases. But they are not alive.

4. Since the variants in the beta globin gene cluster are relatively less complex, it is obvious that NGS can be used to detect HBB-HBD variants. Several organizations around the world have also developed NGS-based testing solutions for thalassemia gene variants. However, the difficult part of NGS testing should be the variation in the alpha globin gene cluster, how do the authors consider this? If it is used purely for detecting HBB, it is of limited innovativeness.

Response: Thank you for your insightful comments regarding the detection complexities of the α-globin gene cluster. We fully agree that the α-globin gene cluster presents more significant challenges due to its complex structure. In our ongoing projects, we are addressing this complexity by developing a long-read targeted sequencing panel using Oxford Nanopore Technology. This panel includes the α-globin gene cluster and its regulatory regions. Long-read sequencing offers distinct advantages over short-read sequencing, particularly in detecting copy number variations (CNVs), which are prevalent in α-globin gene mutations. Moreover, frequency of alpha globin mutation is very less in clinical phenotype in this populations.

In this regard, I would like to mentioned that A study by Nadkarni A. et al. (2008) [PMID: 18447636, DOI: 10.1089/gte.2007.0080] from India, revealed that 10.5% of individual carried a heterozygous alpha-globin gene deletion, while 2.31% had a deletion of two alpha genes. Additionally, 1.1% showed alpha gene triplication. Thus, Thalassemia due to Alpha globin gene mutation or combination with beta mutation rarely present as case. Thus, in the context of Indian population, detection of mutation of HBB gene cluster is very crucial.

For this study, our focus was on developing and testing a gene panel and unique data analysis algorithm specifically designed to identify β-thalassemia-causing mutations, including single nucleotide variants (SNVs) and CNVs. While this panel may not encompass the complexities of the α-globin gene cluster, it represents a significant step forward in the detection of β-thalassemia mutations, which was the primary goal for clinical setting.

We appreciate your suggestion and are committed to incorporating the detection of α-globin gene variants in future studies, leveraging the advancements in long-read sequencing technologies. Thus, we are in the process of development of long read based NGS gene panel, combined alpha and beta mutation detection.

Reviewer:3

The authors used an amplicon sequencing method to sequence the entire β-globin gene cluster. The authors claim that this method can simultaneously detect single nucleotide variants (SNVs) and copy number variants and can replace conventional analytical methods such as ARMS, Sanger sequencing, and MLPA. However, there are some issues that need to be clarified.

1. Is there a typo with the title?   '' ..... one stop detection of copy number alterations and single nucleotide variations......'', should be '' ..... one step detection of copy number alterations and single nucleotide variations......''?

Response: The term "one stop detection" is not any typo. We used "one stop" to emphasize that our sequencing panel provides a comprehensive solution for detecting both copy number alterations and single nucleotide variations in a single, consolidated NGS run and single centre and method. This phrase is meant to convey the convenience and efficiency of our panel as a singular solution. We have used “One stop” for this purpose.

2. To cover the ~80 Kb target region, the authors designed 336 pairs of primers for amplification. How many PCR tubes would be required to perform the amplification with 336 primer pairs? If it requires too many tubes, this could be a time-consuming, laborious, and complicated operation for clinical applications.

Response: In the Illumina Ampliseq method, all 336 primers are indeed amplified in a single pool and tube based library preparation. So, there is no use of multiple tube for this 336-primers panel.

3. Before a new method is applied in clinical practice, its reliability needs to be thoroughly validated through comparative detection of a sufficient number of various genotype samples. However, based on the research content disclosed by the authors, the actual number of samples they tested was only 14, containing a small number of mutation types. This may not be enough to demonstrate the high reliability of the new method. Furthermore, the authors claim that their method can detect a large number of copy number variants in the HbVar database, which seems to be based on the theoretical possibility of the method's coverage range, without clear verification by testing actual samples with these loci. Therefore, the authors may need to expand the sample size for further validation, and include as many mutation types as possible (at least the common mutation types in the Indian region). Especially for some smaller intragenic copy number deletions, it is unclear whether the authors can detect the changes in sequencing depth.

Response: We acknowledge that the sample size in our study is relatively small and your comments.   In this study, we sequenced 14 samples with diverse genotypes, including different (SNVs) in the HBB gene and HBD genes as well as CNVs  to test the robustness of the method, as per qualitative purpose. Accordingly,  the method has able to detect ,  the different SNVs which are demonstrating the capability of our designed gene panel and algorithm  to detect a wide range of mutations of clinical importance.

For CNV detection we used 3 methods to confirm the deletion in beta globin gene cluster 1. Variant mapping over the target region, continuous stretch of homozygous variants in the target region indicating possible large deletion. This possible large deletion within the genome further verified by 3. Depth plot analysis method which detect the coverage variation across the target region. and CNVkit tool

Regarding the detection of copy number variants (CNVs), in HbVar database, we agree that theoretical coverage needs to be substantiated by actual physical method of testing. For this purpose, we need actual blood samples of diverse population and could better way to testify the method. But, we only able to get Asian pupation in this locality, and in this limited scope of the study, thus theoretically validate the coordinate of CNVs in detecting these variations.

4. For the copy number variations detected in the study (such as HBE1-HBB), were they verified by other methods (such as MLPA)? Were they known deletion types, or if they were unreported deletions, could the authors provide the specific breakpoints of the deletions and the detailed hematological phenotype data (such as blood counts, electrophoresis, etc.) of those cases? This additional information could help improve the significance of the paper.

Response: We appreciate your comments and suggestion. For lack of resources, we have not done the MPLA yet. Although some time it is not possible to get exact breakpoint through MLPA. As per suggestion of Reviewer 1, were analysis with the split read method to further confirmation of long deletion and probable breakpoint.

As per your suggestions We have updated the detailed haematological/ phenotype data in the Table 1.

5. Table 1 and Table 2 can be combined into one table, remove the information of "Functional position, consequences, MAF (1000genome), ExAC, GenomeAD, CADD phread score" in Table 1, and add the clinical blood phenotype (blood routine, electrophoresis, etc.) information of the each case.

Response:  As per your suggestions and other reviewers, we have added Table 1 with haematological and clinical data and we have modified the different Tables in the revised manuscript.

6. in Table S1，sample with an ID of S8，the Hb A2 is 4.3%，the Hb F is 101.3. Is there an issue with this (Hb A2+ Hb F>100%)?

Response: We appreciate your diligence in reviewing our data. There was an inadvertent error, and     relevant data has been updated in the Table 1 of revised manuscript

Reviewer - 4.

1.Gene symbols should be styled in italic. Please consult the HUGO Gene Nomenclature Committee’s website (http://www.genenames.org) for current acceptable gene names and symbols. Please confirm or correct as necessary in your text, tables, and figures.

Response: We have changed the Gene Nomenclature accordingly.

2. Microdeletion and microduplication are the shortcomings of NGS. Is the sample size of microdeletion and microduplication tested in this article sufficient, and is the detection ability for microdeletion and microduplication sufficient? Is 14 samples sufficient for verification?

Response: Thank you for raising this important concern regarding the detection of microdeletions and microduplications using NGS. In modern time most of the cases NGS are being used for frequently for detection of CNVs with advance data analysis method. We have attempted the NGS based method to detect both SNV and CNV in a single run and data analysis algorithm. Yes, 14 samples may not be sufficient in number . But we have recruited these 14 samples having different genotype proposition, i.e, SNVs in HBB, HBD, LCR, CNVs with different length and types, like HBB -HBD, HBB-HBG1, HBB-HBE1, entire HBB gene cluster of 80KB, combination of Homozygous and heterozygous CNVs, combination of CNV and SNVs

3. Does the NA in "allele 2" in Table 2 refer to missing or no mutation in the test results? If the latter can be described in a more appropriate way, please correct any other formatting errors in the table.

Response:  In the revised manuscript in Table 3, “NA” has been changed  to “Normal”

4. The detection range of the reagent kit design includes the LCR regulation area. Is there a corresponding sample for experimentation?

Response: Sample (S4, S10, S11) have deletion encompassing entire target region which include both genomic region as well as LCR region

Editorial comments:

1. Please have a native English speaker edit your manuscript or have it professionally edited to improve the quality of the language and grammar in the article.

Response: Thank you for your suggestion, we have corrected any grammatical error, changes are highlighted in yellow in the revised manuscript.

2. Please make sure that you include the most recent literature available.

Response: We have updated the reference no. 6, 10, 21, 25, 27, 29, 30 and 40 with the most recent references available