| Table 1 Material parameters |  |  |
| --- | --- | --- |
| Part | Elastic modulus | Poisson’s ratio |
| cortical bone | 17000 MPa | 0.3 |
| cancellous bone | 445 MPa | 0.2 |
| Titanium alloy | 113800 MPa | 0.342 |

| Table 2 Mesh convergence of the femur and internal fixation analysis at anatomical reduction with the medial wall intact |  |  |  |
| --- | --- | --- | --- |
| Mesh size(mm） | Number of elements | Maximum equivalent stress（MPa） | Difference from the previous group(%） |
| Femur |  |  |  |
| 4.0 | 84961 | 23.786 | - |
| 3.5 | 123969 | 24.498 | 2.906 |
| 3.0 | 193451 | 24.589 | 0.358 |
| Internal fixation (Main nail) |  |  |  |
| 2.5 | 10281 | 56.742 | - |
| 2.0 | 20924 | 70.319 | 19.288 |
| 1.8 | 27091 | 70.761 | 0.625 |

| Table 3 The maximum displacement of the entire assembly（mm） |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Negative 20% | Negative 15% | Negative 10% | Negative 5% | Neutral | Positive 5% | Positive 10% | Positive 15% | Positive 20% |
| Medial wall intact | 5.096 | 4.942 | 4.779 | 4.622 | 4.471 | 4.324 | 4.18 | 4.34 | 5.83 |
| Medial wall defect | 6.2 | 5.168 | 4.964 | 4.797 | 4.567 | 4.409 | 4.247 | 4.333 | 4.434 |

| Table 4 The peak Von Mises Stress (VMS) values (MPa) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Main nail |  | Pressure nail |  | Tension nail |  | Fracture surface |  |
|  | Medial wall intact | Medial wall defect | Medial wall intact | Medial wall defect | Medial wall intact | Medial wall defect | Medial wall intact | Medial wall defect |
| Negative 20% | 85.7 | 188.52 | 16.736 | 75.278 | 34.39 | 59.239 | 27.702 | 23.138 |
| Negative 15% | 82.722 | 173.5 | 17.881 | 67.191 | 33.255 | 58.371 | 21.172 | 21.189 |
| Negative 10% | 79.648 | 166.14 | 19.001 | 65.899 | 31.873 | 52.833 | 18.382 | 16.974 |
| Negative 5% | 76.247 | 165.13 | 19.685 | 64.904 | 30.946 | 52.529 | 14.448 | 16.616 |
| Neutral | 70.761 | 165.02 | 44.23 | 69.271 | 49.547 | 63.267 | 10.25 | 16.957 |
| Positive5% | 69.963 | 164.83 | 21.481 | 53.448 | 27.895 | 47.463 | 8.895 | 15.564 |
| Positive10% | 68.978 | 163.14 | 20.414 | 53.41 | 24.486 | 46.615 | 6.388 | 12.299 |
| Positive15% | 117.67 | 202.53 | 39.564 | 67.843 | 45.955 | 55.035 | 8.098 | 14.912 |
| Positive20% | 135.8 | 211.21 | 41.498 | 70.681 | 48.883 | 57.59 | 15.722 | 17.666 |