## Transition List
Compound Method |  |  | Quantifier Ion |  |  | Qualifier Ion 1 |  |  | Qualifier Ion 2
Name | RT | Ion Polarity | Transition | Collision Energy | Fragmentor | Transition | Collision Energy | Fragmentor | Transition | Collision Energy | Fragmentor
ISTD Picolinic Acid | 1.709 | Positive | 128.0 -> 82.0 | 20 | 82
Picolinic Acid | 1.7263 | Positive | 124.0 -> 78.0 | 20 | 82 | 124.0 -> 106.0 | 8 | 82
ISTD Nicotinic Acid | 1.9066 | Positive | 130.0 -> 85.0 | 24 | 124
Nicotinic Acid | 1.9086 | Positive | 124.0 -> 80.1 | 20 | 80 | 124.0 -> 78.1 | 30 | 80
ISTD Nicotinamide | 2.1477 | Positive | 129.0 -> 85.0 | 24 | 124
Nicotinamide | 2.1687 | Positive | 123.1 -> 80.0 | 24 | 105 | 123.1 -> 53.0 | 36 | 105
ISTD Quinolinic Acid | 2.5791 | Positive | 171.0 -> 81.0 | 24 | 82
Quinolinic Acid | 2.58 | Positive | 168.0 -> 78.0 | 24 | 82 | 168.0 -> 124.0 | 8 | 82
ISTD 3-Hyroxykynurenine | 2.7592 | Positive | 228.0 -> 110.0 | 16 | 92
3-Hyroxykynurenine | 2.7606 | Positive | 225.0 -> 162.0 | 20 | 92 | 225.0 -> 110.0 | 16 | 92
ISTD Serotonin | 5.2195 | Positive | 181.0 -> 136.0 | 22 | 82
Serotonin | 5.3286 | Positive | 177.1 -> 115.0 | 40 | 82 | 177.1 -> 77.1 | 60 | 82
ISTD Kynurenine | 6.0286 | Positive | 213.0 -> 98.0 | 12 | 92
Kynurenine | 6.1717 | Positive | 209.0 -> 146.0 | 20 | 92 | 209.0 -> 94.0 | 12 | 92 | 209.0 -> 118.0 | 30 | 80
ISTD 3-hydroxyanthranillic acid | 6.8975 | Positive | 157.0 -> 83.0 | 24 | 114
3-hydroxyanthranillic acid | 6.9482 | Positive | 154.0 -> 154.0 | 0 | 114 | 154.0 -> 108.0 | 20 | 114 | 154.0 -> 80.0 | 24 | 114
ISTD Tryptophan | 7.4326 | Positive | 208.0 -> 147.0 | 16 | 82
Tryptophan | 7.4417 | Positive | 205.0 -> 146.0 | 16 | 82 | 205.0 -> 118.0 | 32 | 82
ISTD Kynurenic Acid | 7.8834 | Positive | 195.0 -> 149.0 | 20 | 104
Kynurenic Acid | 7.9091 | Positive | 190.0 -> 144.0 | 20 | 104 | 190.0 -> 116.0 | 36 | 104
ISTD Anthranilic acid | 8.9199 | Positive | 138.9 -> 65.0 | 20 | 80 | 138.9 -> 93.1 | 10 | 80
Anthranillic acid | 8.9323 | Positive | 138.0 -> 65.0 | 30 | 80 | 138.0 -> 92.0 | 20 | 80

## LC parameters
Agilent 1290 Infinity II UHPLC
Multisampler G7167B |  | Quaternary pump G7104A | [8 empty] | Column Compartment G7116B
Sampling Speed
Draw Speed | 100.0 µL/min | Flow | 0.400 mL/min | [7 empty] | Left Temperature Control
Eject Speed | 50.0 µL/min | Use Solvent Types | Yes | [7 empty] | Temperature Control Mode | Temperature Set
Wait Time After Drawing | 2.0 s | Low Pressure Limit | 0.00 bar | [7 empty] | Temperature | 50.0 °C
Injection |  | High Pressure Limit | 1200.00 bar | [7 empty] | Enable Analysis Left Temperature
Needle Wash Mode | Standard Wash | Max. Flow Ramp Up | 100.000 mL/min² | [7 empty] | Enable Analysis Left Temperature On | No
Injection Volume | 5.00 µL | Primary Channel | Automatic | [7 empty] | Enable Analysis Left Temperature Value
|  | Max. Flow Ramp Down | 100.000 mL/min² | [7 empty] | Right Temperature Control
Standard Needle Wash |  | Stroke | [8 empty] | Right temperature Control Mode | Combined
Needle Wash Mode | Flush Port | Automatic Stroke Calculation | Yes | [7 empty] | Enable Analysis Right Temperature
Duration | 10 s | [9 empty] | Enable Analysis Right Temperature On | No
|  | Mixer Selection | [8 empty] | Enforce column for run enabled | No
High Throughput |  | Selected Mixer | Do not use Mixer | [7 empty] | Stop Time
Injection Valve to Bypass for Delay Volume Reduction | Yes | Blend Assist | [8 empty] | Stoptime Mode | As pump/injector
Sample Flush-Out Factor | 20.0 | Enabled | No | [7 empty] | Post Time
|  | Stop Time | [8 empty] | Posttime Mode | Off
Overlapped Injection |  | Stoptime Mode | Time set | [7 empty] | Valve Position | Position 2 (Port 1 -> 2)
Overlap Injection Enabled | No | Stoptime | 20.00 min | [7 empty] | Ready when front door open | Yes
|  | Post Time | [8 empty] | Position Switch After Run | Do not switch
Needle Height Position |  | Posttime Mode | Off
Draw Position Offset | 5.0 mm
Use Vial/Well Bottom Sensing | Yes | Solvent Composition | LC Solvents
|  |  | Channel | Solven Composition |  |  | Used
Stop Time |  |  | A | 100 Aqueous w/ 0.1% Formic Acid |  |  | Yes
Stoptime Mode | No Limit |  | B | 90% Acetonitrile, 10% H2O, w/ 0.1% Formic Acid |  |  | Yes
|  |  | C | 100% Aqueous |  |  | No
Post Time |  |  | D | 100% Isopropanol |  |  | No
Posttime Mode | Off
LC gradient program
|  |  |  | Time (min) | A (%) | B (%) | C (%) | D (%) | Flow (mL/min) | Pressure (bar)
|  |  | 1 | 3.50 min | 100.00 % | 0.00 % | 0.00 % | 0.00 % | --- mL/min | --- bar
|  |  | 2 | 16.00 min | 1.00 % | 99.00 % | 0.00 % | 0.00 % | --- mL/min | --- bar
|  |  | 3 | 17.00 min | 1.00 % | 99.00 % | 0.00 % | 0.00 % | 0.400 mL/min | --- bar
|  |  | 4 | 17.10 min | 100.00 % | 0.00 % | 0.00 % | 0.00 % | 0.600 mL/min | --- bar
|  |  | 5 | 19.00 min | 100.00 % | 0.00 % | 0.00 % | 0.00 % | 0.600 mL/min | --- bar
|  |  | 6 | 19.10 min | 100.00 % | 0.00 % | 0.00 % | 0.00 % | 0.400 mL/min | --- bar
|  |  | 7 | 20.00 min | 100.00 % | 0.00 % | 0.00 % | 0.00 % | --- mL/min | --- bar

## MS-QQQ Parameters
Mass Spectrometer | Agilent 6470
Ion Source | AJS ESI
Polarity | Value (+) | Value (-)
Gas Temp (°C) | 200 | 200
Gas Flow (l/min) | 13 | 13
Nebulizer (psi) | 45 | 45
SheathGasHeater | 325 | 325
SheathGasFlow | 12 | 12
Capillary (V) | 3000 | 2000
VCharging | 500 | 500