Point-to-Point Response Letter:

Manuscript ID: EJRE-D-24-00350R2

Title: A Fully Automated, Expert-Perceptive Image Quality Assessment System for Whole-Body [18F]FDG PET/CT

We sincerely thank the reviewer for their constructive feedback, which has significantly improved the clarity and rigor of our manuscript. Below, we provide point-by-point responses to all comments, with revisions highlighted yellow in the updated manuscript.

1. Reviewer Comment:
In my opinion, the comment number 12 has not been sufficiently addressed. Although, you provide extensive reasoning on expert consensus, no explanation is given about the three point difference. For example, why did you not choose two or four point difference.

Response:

Thank you for your insightful question regarding the rationale for selecting a three-point difference as the threshold for expert disagreement. We appreciate the opportunity to clarify this decision.

In our study, expert consensus plays a crucial role in image quality assessment. Subjective evaluation is conducted using the widely adopted 5-point Likert scale, as referenced in the literature ([1] Qi et al., An artificial intelligence-driven image quality assessment system for whole-body [(18)F]FDG PET/CT.  Eur J Nucl Med Mol Imaging; [2] Zhang et al., Deep learning model for automatic image quality assessment in PET. BMC Med Imaging; [3] Obuchowicz et al., Interobserver variability in quality assessment of magnetic resonance images. BMC Med Imaging). This scale establishes predefined thresholds for diagnostically sufficient image quality.

Unlike binary decision-making approaches commonly used in other studies, our study employs multi-class cases to enable finer differentiation across varying quality levels. Given that 15 observers independently assessed images at the slice level, disagreements were inevitable. Inspired by Küstner et al. ([4] Küstner et al. A Machine-learning Framework for Automatic Reference-free Quality Assessment in MRI. Magnetic Resonance Imaging.), we adopted a conflict resolution strategy to address discrepancies. Experts re-evaluated the images and reached a final consensus through discussion, ensuring the validity of the final scores.

Our threshold, established through expert consensus, ensures rigorous quality assessment. Following extensive discussions, we determined that a three-point or greater discrepancy indicates substantial divergence, necessitating expert review. The decision to use a three-point threshold instead of two or four was based on the following considerations:

- A two-point difference (e.g., 1 vs. 3) would have triggered excessive arbitration, increasing the burden on experts and reducing evaluation efficiency.

- A four-point difference (e.g., 1 vs. 5) might fail to capture significant, though less extreme, disagreements (e.g., 2 vs. 5), potentially compromising the consistency of quality assessments.

- A three-point threshold achieves an optimal balance, minimizing unnecessary arbitration while ensuring major discrepancies (e.g., 1 vs. 4, 2 vs. 5, 1 vs. 5) are properly addressed.

This approach enhances the expert-driven nature of our PET/CT IQA system, reducing labeling noise and improving model reliability.

We appreciate your valuable feedback and hope this revision fully addresses your concern.

2. Reviewer Comment:

What systems were used? Although, you have the information in table 1, no reference is provided. Please add here that system characteristics are provided in table 1.

Response:

We have inserted a direct reference to Table 1 in the revised text (Page 6, Line 113-114):

"All PET/CT scanners underwent routine quality assurance (QA) and quality control (QC) procedures, with their system characteristics detailed in Table 1."

3. Reviewer Comment:
What are the key axial slices? Please add a short explanation of key axial slices and refer to the Supplement.

Response:
We have expanded the definition and added a supplementary reference (Page 6, Lines 125-130):

" We developed customized algorithms to automatically localize seven key axial slices of the whole body that capture critical anatomical landmarks (including the basal ganglia, cerebellar vermis, aortic arch, tracheal carina, liver, pancreas, and iliac bone). These slices were selected for their anatomical significance and reproducibility across patients, with a full description of our segmentation and localization methodologies provided in Supplementary Section 1.1 and 1.2."

4. Reviewer Comment:
Please change "A full exposition of our ..." to "A full description of our ...", which is more natural.

Response:
We have revised this word as suggested (see point 3 for details).

5. Reviewer Comment:

This reference can be provided already in the previous section (see point 3).

Response:

We have moved supplement reference earlier as suggested (see point 3 for details).

6. Reviewer Comment:
What threshholds are being used for the conversion into grayscale and binary images? Do the OpenCV modules this automatically or do you need to set the parameters.

Response:

In the OpenCV modules, the cv2.threshold function requires parameter settings, where maxval=255 defines the white pixel value after binarization. The threshold type is specified as cv2.THRESH_BINARY + cv2.THRESH_OTSU, enabling automatic threshold selection via the OTSU algorithm. Given the grayscale distribution characteristics of CT images, this algorithm automatically determines the global optimal threshold by maximizing inter-class variance, effectively distinguishing HU value differences between target organs and the background.

7. Reviewer Comment:
Please change "... function was utilized to compare ..." to "... function was applied to compare ...", which is more natural.

Response:

Revised as suggested (Page 8, Line 153):

"...function was applied to compare..."

8. Reviewer Comment:
This is a very long sentence, please consider separating into two or three sentences.

Response:
We restructured the sentence for clarity (Page 8, Lines 156-159):

"To better align with the varying characteristics of different anatomical regions, we tailored the localization scheme by incorporating selection logic. For example, we customized the automatic recognition function for the aortic arch in the thorax and implemented tailored evaluation criteria for this slice."

9.Reviewer Comment:
Please change "Details ... was summarized in ..." to "Details ... are summarized in ...

Response:
Revised as suggested (Page 8, Line 162):

"Details... are summarized..."

10. Reviewer Comment:
Please remove the comma in " ... consistent ranges, as shown ..."

Response:
Corrected (Page 8, Line 167):

"...consistent ranges as shown..."

11. Reviewer Comment:
Please add 'the' in the following sentence "... landmarks such as condyle ..." (i.e. "... landmarks such as the condyle ..."

Response:
Revised (Page 8, Line 175):

"...landmarks such as the condyle..."

12. Reviewer Comment:
Please change 'demonstrated' to 'demonstrates'

Response:
Revised (Page 13, Line 268):

"Fig. 6 demonstrates..."

13. Reviewer Comment:
Please review the references. For example, some use acronyms for journal names and some do not. Please be consistent.

Response:

The references have been revised as suggested.

14. Reviewer Comment:
Please change '... from Different Centres ...' to '... from Different Centers ...'. Throughout the manuscript US English is used, please be consistent. The names of the centers (e.g. Shanghai University Medical Imaging Diagnostic Centre) can remain as is.

Response:
Revised as suggested (Page 26).

15. Reviewer Comment:

Please stay consistent with 'n=' or 'xxx CT & xxx PET images'. In the figure, the number of patients changes abruptly with the number of images, without justification.

Response:
We added a footnote to Fig. 1(Page 29):

"n representing the number of patients. "

16. Reviewer Comment:
It is not entirely clear to me what the concentrations (e.g. 15 g/mL or 7 g/mL) exactly are. Is the the upper limit of the window level? Please clarify.

Response:
As the footnote in Fig. 3 we have clarified as:

"For head images, we set an initial tracer concentration of 15 g/mL, and for other images, the concentration was set to 7 g/mL, allowing physicians to make slight adjustments according to their diagnostic preferences."

17. Reviewer Comment:
Please add space between Fig. and 6.

Response:
Revised as suggested (Page 34, Line 606).

18. Reviewer Comment:
(Small suggestion) Page 36, Figure 6: The boxplots are very small and illegible. Please consider placing Fig. 6b below 6a and increase the size of the figures.

Response:
Thank you for your valuable feedback. We have revised Fig. 6 to improve its clarity and legibility. The updated figure is now included in the revised manuscript.

General remarks:

Reviewer Comment:

a. In the discussion you mention the limitations that the IQA system is only trained on limited data from four centers and further research is necessary to assess generalizability. However, what I also would find interesting to see is the inter-institution comparison. How did the software manage to clasify the data of each center separately? This may provide some insight into how well it will be generalizable.

Response to Reviewer’s Comment:

Thank you for your insightful question.

This study focuses on developing a slice-based IQA system, trained on data from four medical centers. Although training on a limited dataset may present a limitation, the expert-driven slice-based approach remains a key strength. To enhance generalizability, we are actively incorporating data from additional centers.

The software manages to classify the data from each center separately by assigning a unique identifier to each patient's image. This identifier allows the software to trace back to the patient's basic information via tags, ensuring that data from different centers can be distinguished and appropriately classified. Although this study does not classify data by institution, future work will include center-based validation to assess adaptability across different clinical settings, providing valuable insights into its robustness and generalizability.

We appreciate your feedback and will address this in future research.

Reviewer Comment:

b. In this work, no mention is made about the homogeneity of the datasets. If you have inhomogeneous data, how does this IQA system perform. For example, data from long axial field of view PET/CT systems, has a variable sensitivity along the axial field of view. Likewise, for continuous bed motion with a longer scan duration or slower bed speed for only one region, the number of collected counts and thus the image quality varies along the axial dimension. How does the software classify the images in such situations?

Response to Reviewer’s Comment:

Thank you for your thoughtful question regarding dataset homogeneity and its impact on our IQA system’s performance.

Our study includes only whole-body [18F]FDG PET/CT scans (head to mid-thigh) (see Page 6, Line 117–118) and therefore excludes regional scans with extended or reduced scan durations to ensure data consistency. A key innovation of our approach is developing an IQA system that replicates expert subjective perception, aligning with clinical diagnostic standards while adapting to variations in scanner types and acquisition parameters.

While we acknowledge the complexity of factors affecting image quality, our approach simplifies this challenge by using algorithms to simulate expert evaluations as the reference standards. As mentioned earlier, this study lays the foundation for future validation on heterogeneous datasets to further assess the applicability of expert-perceptive IQA in real-world clinical practice. We appreciate your valuable feedback and hope this explanation clarifies our approach.

We believe these revisions adequately address your comments and further enhance the quality of the manuscript. Thank you once again for your invaluable feedback and continued support in refining our work.

Should you require any additional details or clarification, please do not hesitate to contact me at 15314282327@163.com.

Kind regards,
Jiahe Tian, MD
Department of Nuclear Medicine
Chinese PLA General Hospital