Supplemental Table 1. 3x3 agreement under laboratory conditions using proposed thresholds.

|  |  | Spectrophotometer (reference) |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Normal | Intermediate | Deficient | Sum |
| Wondfo G6PD/Hb Test | Normal | 112 | 0 | 0 | 112 |
|  | Intermediate | 93 | 11 | 0 | 104 |
|  | Deficient | 2 | 12 | 34 | 48 |
|  | Sum | 207 | 23 | 34 | 264 |

Supplemental Table 2. 3x3 agreement under simulated field conditions using proposed thresholds.

|  |  | Spectrophotometer (reference) |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Normal | Intermediate | Deficient | Sum |
| Wondfo G6PD/Hb Test | Normal | 106 | 0 | 0 | 106 |
|  | Intermediate | 97 | 12 | 0 | 109 |
|  | Deficient | 4 | 11 | 33 | 48 |
|  | Sum | 207 | 23 | 33 | 263 |

Supplemental Table 3. 3x3 agreement for fresh capillary specimens using proposed thresholds.

|  |  | Spectrophotometer (reference) |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Normal | Intermediate | Deficient | Sum |
| Wondfo G6PD/Hb Test (Capillary) | Normal | 195 | 0 | 0 | 195 |
|  | Intermediate | 3 | 6 | 0 | 9 |
|  | Deficient | 1 | 0 | 20 | 21 |
|  | Sum | 199 | 6 | 20 | 225 |

Supplemental Table 4. 3x3 agreement for fresh venous specimens using proposed thresholds.

|  |  | Spectrophotometer (reference) |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Normal | Intermediate | Deficient | Sum |
| Wondfo G6PD/Hb Test (Venous) | Normal | 195 | 1 | 0 | 196 |
|  | Intermediate | 3 | 5 | 0 | 8 |
|  | Deficient | 1 | 0 | 20 | 21 |
|  | Sum | 199 | 6 | 20 | 225 |