Supplemental Table 1. 3x3 agreement under laboratory conditions using proposed thresholds.
	
	Spectrophotometer (reference)

	
	Normal
	Intermediate
	Deficient
	Sum

	Wondfo G6PD/Hb Test
	Normal
	112
	0
	0
	112

	
	Intermediate
	93
	11
	0
	104

	
	Deficient
	2
	12
	34
	48

	
	Sum
	207
	23
	34
	264



Supplemental Table 2. 3x3 agreement under simulated field conditions using proposed thresholds.
	
	Spectrophotometer (reference)

	
	Normal
	Intermediate
	Deficient
	Sum

	Wondfo G6PD/Hb Test
	Normal
	106
	0
	0
	106

	
	Intermediate
	97
	12
	0
	109

	
	Deficient
	4
	11
	33
	48

	
	Sum
	207
	23
	33
	263



Supplemental Table 3. 3x3 agreement for fresh capillary specimens using proposed thresholds.
	
	Spectrophotometer (reference)

	
	Normal
	Intermediate
	Deficient
	Sum

	Wondfo G6PD/Hb Test (Capillary)
	Normal
	195
	0
	0
	195

	
	Intermediate
	3
	6
	0
	9

	
	Deficient
	1
	0
	20
	21

	
	Sum
	199
	6
	20
	225



Supplemental Table 4. 3x3 agreement for fresh venous specimens using proposed thresholds.
	
	Spectrophotometer (reference)

	
	Normal
	Intermediate
	Deficient
	Sum

	Wondfo G6PD/Hb Test (Venous)
	Normal
	195
	1
	0
	196

	
	Intermediate
	3
	5
	0
	8

	
	Deficient
	1
	0
	20
	21

	
	Sum
	199
	6
	20
	225



