Tables

Table 1 The label and specification of 10 studied ultraviolet C germicidal devices.

| UVC light source | Application | Label | UVC wave range (nm) | UV light power (W) |
| --- | --- | --- | --- | --- |
| Low-pressure mercury lamps | Handheld UVGI surface disinfection equipment | M1 | N/A | 2.5 |
|  |  | M2 | 253.7 | 2 |
|  | UVGI disinfection chamber | M3 | 253.7 | 24 |
|  |  | M4 | 253.7 | 8 |
|  | Movable UVGI air and surface purifier | M5 | N/A | 38 |
|  |  | M6 | 253.7 | 24 |
|  |  | M7 | 253.7 | 8 |
| Light-emitting diodes | Handheld UVGI surface disinfection equipment | L1 | 265-280 | 4.5 |
|  | UVGI disinfection chamber | L2 | 265 | 8 |
|  |  | L3 | 270-285 | 5-8 |

UVC, ultraviolet C; UVGI, ultraviolet germicidal irradiation; N/A, not applicable

Table 2 Ultraviolet C dosage and recommended settings of the 3 studied handheld UVGI surface disinfection equipment.

|  | Handheld UVGI surface disinfection equipment |  |  |
| --- | --- | --- | --- |
|  | Low-pressure mercury lamps |  | Light-emitting diodes |
| Label | M1 | M2 | L1 |
| Power (W) | 2.5 | 2 | 4.5 |
| Short distance | 2 cm | 2 cm | 2 cm |
| Cycle duration | 15/30 min | 30 sec | 2 sec |
| UVC dosage (mJ/cm2) | 260.60/436.54 | 4.52 | 0.29 |
| Long distance | 1.4 m | 5 cm | 3 cm |
| Cycle duration | 15/30 min | 30 sec | 2 sec |
| UVC dosage (mJ/cm2) | Undetectable | 2.46 | 0.22 |
| Recommend settings |  |  |  |
| Manufacturer |  |  |  |
| Distance | No more than 1.4 m | 3-5 cm | 1-3 cm |
| Cycle duration | 15/30 min | 30 sec | 2 sec |
| Author |  |  |  |
| Distance | No more than 10 cm | 3-5 cm | 1-3 cm |
| Cycle duration | 2 min | 45 sec | 35 sec |

UVC, ultraviolet C; UVGI, ultraviolet germicidal irradiation

Table 3 Ultraviolet C dosage and recommended settings of the 4 studied UVGI disinfection chambers.

|  | UVGI disinfection chambers |  |  |  |
| --- | --- | --- | --- | --- |
|  | Low-pressure mercury lamps |  | Light-emitting diodes |  |
| Label | M3 | M4 | L2 | L3 |
| Power (W) | 24 | 8 | 8 | 5-8 |
| Short distance | 2 cm | 2 cm | 2 cm | 2 cm |
| Cycle duration | 35 sec | 11/16/21 min | 5 min | 3 min |
| UVC dosage (mJ/cm2) | 86.14 | 790.22/1115.19/ 1423.00 | 19.40 | 7.58 |
| Long distance | 10 cm | 10 cm | 10 cm | 10 cm |
| Cycle duration | 35 sec | 11/16/21 min | 5 min | 3 min |
| UVC dosage (mJ/cm2) | 22.62 | 411.14/587.35/ 755.75 | 6.80 | 2.05 |
| Recommend setting |  |  |  |  |
| Manufacturer |  |  |  |  |
| Distance | In chamber | In chamber | In chamber | In chamber |
| Cycle duration | 35 sec | 11/16/21 min | 5 min | 3 min |
| Author |  |  |  |  |
| Distance | In chamber | In chamber | In chamber | In chamber |
| Cycle duration | 6 sec | 8 sec | 3 min | 6 min |

UVC, ultraviolet C; UVGI, ultraviolet germicidal irradiation

Table 4 Ultraviolet C dosage and recommended settings of the 3 studied movable UVGI air and surface purifiers.

|  | Movable UVGI air and surface purifiers |  |  |
| --- | --- | --- | --- |
|  | Low-pressure mercury lamps |  |  |
| Label | M5 | M6 | M7 |
| Power (W) | 38 | 24 | 8 |
| Short distance | 2 cm | 2 cm | 2 cm |
| Cycle duration | 15 min | 15 min | 15 min |
| UVC dosage (mJ/cm2) | 9622.97 | 9137.52 | 2794.49 |
| Long distance | 2 m | 2 m | 2 m |
| Cycle duration | 15 min | 15 min | 15 min |
| UVC dosage (mJ/cm2) | 36.55 | 21.61 | 5.37 |
| Recommend setting |  |  |  |
| Manufacturer |  |  |  |
| Distance | 2 m | 2 m | N/A |
| Cycle duration | 15 min | 15 min | N/A |
| Author |  |  |  |
| Distance | 2 m | 2 m | 2 m |
| Cycle duration | 2 min | 3 min | 11 min |

UVC, ultraviolet C; UVGI, ultraviolet germicidal irradiation; N/A, not applicable