Table 1. Participant demographics and health status.

|  | Total N=98 | PID cases N=49 | Control N=49 |
| --- | --- | --- | --- |
| Age in years Mean (SD) Range | 5.6 (4.1) [0.4, 15] | 5.5 (4.2) [0.5, 14.3] | 5.7 (4.1) [0.4, 15] |
| Gender, number (%) Female Male | 38 (39%) 60 (61%) | 19 (39%) 30 (61%) | 19 (39%) 30 (61%) |
| Breastfeeding number (%) Mean age in years (SD) Range | 16 (16%) 1.1(0.7) [0.4, 3.1] | 8 (16%) 0.9 (0.2) [0.6, 1.2] | 8 (16%) 1.3 (0.9) [0.4, 3.1] |
| Vaccinated against polio, number (%) | 95 (97%) | 46 (94%) | 49 (100%) |
| Polio vaccine type, first and/or second doses, % Oral Injection | 14 (15%) 84 (88%) | 9 (19.6%) 38 (82.6%) | 5 (10.2%) 46 (93.9%) |
| Jeffrey Modell Warning Signs (JM signs) | N=98 | N=49 | N=49 |
| Children with at least 1 JM sign | 54 (55%) | 49 (100%) | 5 (10%) |
| Children with at least 2 JM signs | 49 (50%) | 49 (100%) | 0 (0%) |
| PID diagnosis confirmed |  |  |  |
| Hypogammaglobulinemia: AG, CID, SCID | 11 (11%) | 11 (22%) | 0 (0%) |
| CVID | 9 (9%) | 9 (18%) | 0 (0%) |
| IgA deficiency | 1 (1%) | 1 (2%) | 0 (0%) |
| Suspected PID (Immune dysregulation, HIgM) | 27 (28%) | 27 (55%) | 0 (0%) |
| HIV | 1 (1%) | 1 (2%) | 0 (0%) |
| Healthy | 49 (50%) | 0 (0%) | 49 (100%) |

Table 2. Summary of reference and investigational test results amongst suspected PID, PID confirmed under IVIG treatment, and controls. Results for laboratory reference IgG testing category and IgG Total Test results for capillary and for serum are reported for IVIG-naive suspected PID cases, confirmed PID cases with known IVIG treatment, study Controls of healthy age-matched children, and all enrollees combined. Reference IgG is stratified into categories of IgG ranges <3 g/L, 3≤IgG<5 g/L, IgG ≥5 g/L, and described by median and interquartile values of IgG levels.

|  |  | Cases |  |  | Controls | Combined Cases and Controls |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Suspected PID No known IVIG |  | PID Receiving IVIG |  |  |
|  | PID diagnosis group ► | PID+ |  | PID+ | PID- | N/A |
| Test▼ | IgG category▼ | n=25* |  | n=24 | n=49 | n = 98 |
| Laboratory Reference IgG testing | Serum sample |  |  |  |  |  |
|  | IgG < 3 g/L | 0 | 1 |  | 0 | 1 |
|  | 3 g/l ≤IgG < 5 g/L | 0 | 2 |  | 2 | 4 |
|  | IgG≥5 g/L | 25 | 21 |  | 47 | 93 |
|  | Median IgG (Q1-Q3), g/L | 11.3 (8.6 -13.6) | 10.4 (7.0-14.8) |  | 10.9 (9.5 - 13.1) | 11.0 (8.8-13.3) |
| IgG Total Test | Capillary blood |  |  |  |  |  |
|  | IgG Total Test (1 or 2 lines) | 14 | 15 |  | 21 | 50 |
|  | IgG Total Test (0 lines) line | 11 | 9 |  | 28 | 48 |
|  | Serum |  |  |  |  |  |
|  | IgG Total Test (1 or 2 lines) | 1 | 4 |  | 1 | 6 |
|  | IgG Total Test (0 lines) | 24 | 10 |  | 48 | 91 |

*One suspected PID was later diagnosed as HIV-positive

Table 3. Analytical agreement between the reference assay (immunology panel) and the IgG Total Test results using capillary blood and serum samples. The results for the IgG Total Test are shown for capillary and serum samples, stratified by the IgG level of <3 g/L, 3 g/l ≤IgG < 5 g/L (“3-5 g/L”), and > 5g/L.  Percent correct results were calculated using highlighted cells corresponding to number of results divided by the total number in the category.

|  |  | Cases |  |  | Controls |  | Cases & Controls |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Reference IgG, g/L ► | <3* | 3-5* | >5 | 3-5* | >5 | <3 | 3-5 | >5 |
| Capillary | Test lines = result ▼ | n=1 | n=2 | n=46 | n=2 | n=47 | n=1 | n=4 | n=93 |
|  | 2 = very low IgG | 1 | 2 | 21 | 2 | 14 | 1 | 4 | 35 |
|  | 1 = moderately low IgG | 0 | 0 | 5 | 0 | 5 | 0 | 0 | 10 |
|  | 0 = normal IgG | 0 | 0 | 20 | 0 | 28 | 0 | 0 | 48 |
|  | % Correct | 100% | 0% | 43% | 0% | 60% | 100% | 0% | 52% |
| Serum | Serum | n=1 | n=1** | n=46 | n=2 | n=47 | n=1 | n=3** | n=93 |
|  | 2 = very low IgG | 1 | 1 | 2 | 0 | 1 | 1 | 1 | 3 |
|  | 1 = moderately low IgG | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 |
|  | 0 = normal IgG | 0 | 0 | 43 | 2 | 46 | 0 | 2 | 89 |
|  | % Correct | 100% | 0% | 93% | 0% | 98% | 100% | 0% | 96% |

*One case sample < 3 g/L was below the lower limit of quantification (LLOQ), with a result of <0.3 g/L.  It was analyzed as having a value of 0.3 g/L.  Two samples (one case, one control) from 3-5 g/L were below LLOQ, with results listed as <3.0 g/L.  They were analyzed as having a value of 3 g/L.

**One serum sample was not available to test.