Supplement 2.	Definitions Applied during the Usability Testing for Observed Outcomes for each Predefined Task

Supplement 3.	Specifications of Potential Use Errors That Can be Avoided through Safe Devise Design

Supplement 4. 	Frequency of Observed Outcome and Result, Analysis per Task Description

Supplement 5. 	The System Usability Scale (SUS) to Assess Participants’ Levels of Agreement with the Overall Usability of the System. Test Participants N=20 (Results per Question).

Supplement 6. 	The System Usability Scale (SUS) to Assess Participants’ Level of Agreement with the Overall Usability of the System (Results per Test Participant).

Supplement 7. 	After-Scenario Questionnaire (ASQ) to Assess the Participants’ Level of Satisfaction with the NILS model. Test participants N=20 (Results per Question).

Supplement 2. Definitions Applied during the Usability Testing for Observed Outcomes for each Predefined Task

| Term | Definition |
| --- | --- |
| Correct use | The user completes the task correctly. |
| Close call | Close calls are instances in which a user has difficulty or makes a use error that could result in harm, but the user takes an action to “recover” and prevents the harm from occurring. Close calls should be recorded when they are observed and discussed with the test participants after they have completed all the use scenarios. In addition, repeated attempts to complete a task and apparent confusion could indicate potential use error and therefore should also be collected as observational data and discussed during the interviews with test participants. |
| Use error | User action or lack of action that was different from that expected by the manufacturer and caused a result that: (1) was different from the result expected by the user and (2) was not caused solely by device failure and (3) did or could result in harm. |
| Use difficulty | Use difficulty is when a user seems to struggle to perform a task. This may be observed in different ways: the user may pause to read the manual for a longer time, or a task may take longer time than expected (longer than other tasks) to complete. It is also possible that the participant verbally expresses that a task is difficult. Use difficulty may lead to use errors. |

Supplement 3. Specifications of Potential Use Errors That Can be Avoided through Safe Devise Design

| Hazard | Hazardous situation | Cause |
| --- | --- | --- |
| Complex user interface | False estimated indication of benign SLN* | User interface is complex and allows misinterpretation |
|  | False estimated indication of benign SLN | User is confused when the user interface reports the probability of benign sentinel lymph node status instead of risk/probability for metastatic SLN status. This refer to the probability of "no-disease" which is not commonly used in risk prediction models. Because of this the probability and threshold can be misinterpreted, above threshold is in fact representing "no-disease" state. |
|  | False estimated indication of metastatic SLN | The interpreting physician does not understand the graph background data (histogram) and do not know how to interpret the results. |
|  | False estimated indication of metastatic SLN | Colours are used in a way that is not conventional and intuitive. The physician is confused how to interpret the results if signal colours are used inconsistently. |
|  | False estimated indication of metastatic SLN | Physician is unsure how to interpret data that is very close to the threshold or is unsure if the results can be used in order to omit SLNB |
| Use error | False estimated indication of benign SLN | The physician using NILS selects "Other/Mixed" as type in histopathological subtype with the intention of selecting "Unknown" for imputation. |
| Information provided by the manufacturer | False estimated indication of benign SLN | Incomplete instructions for use |
|  | False estimated indication of benign SLN | Inadequate description of performance characteristics |
|  | False estimated indication of benign SLN | Inadequate specification of intended use |
|  | False estimated indication of benign SLN | Inadequate disclosure of limitations |
|  | False estimated indication of benign SLN | Information provided by the manufacturer is not readily available and therefore omitted/missed at time point of NILS calculator |

* False estimated indication of benign sentinel lymph node is the most important harm to avoid.

The risk management procedure is based on the guidance provided in ISO 14971 (International Electrotechnical Commission IE. IEC TR 62366-2:2016 Medical devices - Part 2: Guidance on the application of usability engineering to medical devices (Edition 1.0) 2016. https://www.iso.org/standard/69126.html)

Supplement 4. Frequency of Observed Outcome and Result, Analysis per Task Description

|  | Task description | Potential harm | Observed outcome, N (%) |  |  |  | Result, N (%) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Correct use | Use error | Use difficulty | Close call | Pass | Fail |
| 1. | Reset the calculator | False estimated indication of benign SLN | 68 (85) | 12 (15) | 0 (0) | 0 (0) | 68 (85) | 12 (15) |
| 2. | Enter clinical data | False estimated indication of benign SLN | 83 (83) | 1 (1) | 16 (16) | 0 (0) | 99 (99) | 1 (1) |
| 3. | Enter mammography data | False estimated indication of benign SLN | 77 (96) | 1 (1) | 2 (3) | 0 (0) | 79 (99) | 1 (1) |
| 4. | Enter core biopsy data | False estimated indication of benign SLN | 73 (91) | 7 (9) | 0 (0) | 0 (0) | 73 (91) | 7 (9) |
| 5. | Perform calculation | - | 71 (89) | 9 (11) | 0 (0) | 0 (0) | 71 (89*) | 9 (11) |
| 6. | Fill out the questionnaire with the appropriate clinical pathway selected considering the results of NILS in combination with all other available information of the case. | False estimated indication of benign SLN | 94 (94) | 4 (4) | 1 (1) | 1 (1) | 96 (96) | 4 (4) |

Cell color indications: Green indicates a task result categorized as "pass" (meeting or exceeding the pre-defined acceptance criteria of ≥ 90%); Red indicates a task result categorized as "fail" (falling below the pre-defined acceptance criteria of ≥ 90%).

*The “Perform calculation” were below the acceptance criteria of ≥ 90% for the first of the four cases when analyzed per task. However, the summary of all cases resulted in a percentage just below the threshold. Altogether, reflecting a learning curve.

Supplement 5. The System Usability Scale (SUS) to Assess Participants’ Levels of Agreement with the Overall Usability of the System. Test Participants N=20 (Results per Question)

Please note that odd-numbered items are positive and even-numbered items are negative.

|  |  |  | Likert scale 1-5: 1, “Strongly Disagree”; 5 “Strongly Agree“ | SUS score |
| --- | --- | --- | --- | --- |
|  |  | Theme from SUS | Average point mean (SD) median (range) | Average score mean median |
| 1 | I think that I would like to use this system frequently | Readiness to use | 4.15 (0.75) 4.00 (3.00-5.00) | 3.15 3.00 |
| 2 | I found the system unnecessarily complex | Tool simplicity | 1.25 (0.44) 1.00 (1.00-2.00) | 3.75 4.00 |
| 3 | I thought the system was easy to use | Ease of use | 4.55 (0.61) 5.00 (3.00-5.00) | 3.55 4.00 |
| 4 | I think that I would need the support of a technical person to be able to use this system | Need for support to use tool | 1.15 (0.37) 1.00 (1.00-2.00) | 3.85 4.00 |
| 5 | I found the various functions in this system were well integrated | Understanding the input parameters | 4.40 (0.60) 4.00 (3.00-5.00) | 3.40 3.00 |
| 6 | I thought there was too much inconsistency in this system | Clarity of the tool | 1.25 (0.55) 1.00 (1.00-3.00) | 3.75 4.00 |
| 7 | I would imagine that most people would learn to use this system very quickly | Need for technical details | 4.70 (0.47) 5.00 (4.00-5.00) | 3.70 4.00 |
| 8 | I found the system very awkward to use | Easier-to-use tool | 1.20 (0.41) 1.00 (1.00-2.00) | 3.80 4.00 |
| 9 | I felt very confident using the system | Usability confidence | 4.05 (0.61) 4.00 (3.00-5.00) | 3.05 3.00 |
| 10 | I needed to learn a lot of things before I could get going with this system | Easy to use for everyone | 1.20 (0.70) 1.00 (1.00-4.00) | 3.80 4.00 |
|  | SUS score, mean | (3.15+3.75+3.55+3.85+3.40+3.75+3.70+3.80+3.05+3.80) x 2.5 Σ score: 89.5* |  |  |
|  | SUS score, median | (3+4+4+4+3+4+4+4+3+4) x 2.5 Σ score: 92.5 |  |  |

*Categorized as “Excellent”, according to the commonly used Adjective Rating Scale (A. Bangor et al, 2009). The Adjective Rating Scale: “Best imaginable” ≥ 90.9; “Excellent” ≥ 85.5; “Good” ≥ 71.4; “OK/Fair” ≥ 50.9; “Poor” ≥ 35.7; “Awful” ≥ 20.3; and “Worst imaginable” ≥ 12.5.

Supplement 6. The System Usability Scale (SUS) to Assess Participants’ Level of Agreement with the Overall Usability of the System (Results per Test Participant)

- The Table Shows Individual SUS Score for Each Test Participant

- Box Plot, Mean SUS Score per Test Participant = 89.4

| Test participant | Question |  |  |  |  |  |  |  |  |  | SUS score |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Σ |
| 1 | 2 | 4 | 3 | 4 | 3 | 4 | 4 | 3 | 3 | 4 | 85.0 |
| 2 | 3 | 4 | 2 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 75.0 |
| 3 | 3 | 4 | 4 | 4 | 3 | 4 | 4 | 4 | 3 | 4 | 92.5 |
| 4 | 3 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | 4 | 95.0 |
| 5 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | 4 | 97.5 |
| 6 | 4 | 4 | 3 | 4 | 4 | 4 | 4 | 3 | 4 | 4 | 95.0 |
| 7 | 3 | 4 | 3 | 4 | 2 | 3 | 3 | 3 | 2 | 1 | 70.0 |
| 8 | 3 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | 4 | 95.0 |
| 9 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 100.0 |
| 10 | 3 | 3 | 3 | 4 | 3 | 4 | 3 | 4 | 2 | 4 | 82.5 |
| 11 | 3 | 3 | 4 | 4 | 4 | 2 | 3 | 4 | 3 | 4 | 85.0 |
| 12 | 2 | 4 | 4 | 4 | 3 | 3 | 4 | 4 | 4 | 4 | 90.0 |
| 13 | 2 | 4 | 4 | 4 | 4 | 4 | 3 | 4 | 3 | 4 | 90.0 |
| 14 | 4 | 3 | 4 | 3 | 3 | 4 | 4 | 4 | 3 | 4 | 90.0 |
| 15 | 2 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 2 | 4 | 90.0 |
| 16 | 4 | 3 | 4 | 3 | 3 | 4 | 4 | 4 | 3 | 4 | 90.0 |
| 17 | 4 | 4 | 3 | 4 | 3 | 4 | 4 | 4 | 3 | 3 | 90.0 |
| 18 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 100.0 |
| 19 | 3 | 3 | 3 | 4 | 3 | 4 | 4 | 4 | 3 | 4 | 87.5 |
| 20 | 3 | 4 | 3 | 4 | 3 | 4 | 3 | 4 | 3 | 4 | 87.5 |
| SUS score per test participants mean |  |  |  |  |  |  |  |  |  |  | Σ score: 89.4* |
| SUS score per test participant median |  |  |  |  |  |  |  |  |  |  | Σ score: 90.0 |

*Categorized as “Excellent”, according to the commonly used Adjective Rating Scale (A. Bangor et al, 2009). The Adjective Rating Scale: “Best imaginable” ≥ 90.9; “Excellent” ≥ 85.5; “Good” ≥ 71.4; “OK/Fair” ≥ 50.9; “Poor” ≥ 35.7; “Awful” ≥ 20.3; and “Worst imaginable” ≥ 12.5.

b)

Supplement 7. After-Scenario Questionnaire (ASQ) to Assess the Participants’ Level of Satisfaction with the NILS model. Test participants N=20 (Results per Question).

|  |  | Likert scale 1-7: 1, “Strongly Disagree”; 7, “Strongly Agree“ |
| --- | --- | --- |
|  |  | Average point mean (SD) median (range) |
| 1 | Overall, I am satisfied with the ease of completing the tasks in this scenario. | 6.20 (0.70) 6.00 (5.00-7.00) |
| 2 | Overall, I am satisfied with the amount of time it took to complete the tasks in this scenario. | 6.45 (0.61) 6.50 (5.00-7.00) |
| 3 | Overall, I am satisfied with the support information (on-line help, documentation) when completing the tasks. | 6.25 (0.97) 6.50 (4.00-7.00) |
|  | Total | mean: 6.30 median: 6.50 |