Additional file 1 detailed methods and results of validation

Included to the Jet-iStar 800 is test device, detection buffer and calibration curve. Test device consists of HBP monoclonal antibody, goat anti-rabbit IgG and a nitrocellulose membrane found on the reaction plate. The detection buffer is pre-aliquoted in tubes containing 0.15mL fluorescently labeled HBP monoclonal antibody, fluorescently labeled rabbit IgG and phosphate buffer (pH 7.4). For calibrating the kit, only the provided calibration curve found as a QR code is necessary.

Before analyzing the plasma samples all reagents should be equilibrated to room temperature. Following, 50L of homogenized plasma was pipetted into the detection buffer tube and thoroughly mixed by inverting. From the detection buffer tube 100L of mixture was added vertically to the well of the reaction plate. Mixture was set for incubation for 18 min. During incubation the mixture diffuses into the nitrocellulose membrane by capillary action and is captured by the paired antibody immobilized by the membrane. Prior to inserting the reaction plate into the machine, the QR code should be scanned and the project- and product basic information should be confirmed. On the screen of the Jet-iStar 800 the sample type is selected, and the reading starts by clicking on the Quick test button. The intensity of the fluorescent antibody reflects the amount of antigen captured and the HBP concentration found in the plasma sample is determined by immunoassay analysis. The Jet-iStar 800 analyzer instant report detected concentration for each sample in the range between 5.90ng/mL- 300.00 ng/mL HBP. Multiple drugs have been tested for interference with the Jet-iStar 800, many of which show no obvious effect on the test results at tested concentrations. However, some interference was found when analyzing hemoglobin (4.0mg/mL), Triglyceride (8.5mg/mL) and Bilirubin (0.2mg/mL).

Additional table 1. Validation of performance for Jet-iStar 800 point of care test for plasma HBP

| Test item | Tested value | Target value according to the manufacturer |
| --- | --- | --- |
| Accuracy | R= 100.75% | Recovery rate (R) should be within the range 85-115% |
| Lower detection limit | LOD <5.90ng/mL | Lower limit of detection (LOD) should be <5.90ng/mL |
| Linearity | r= 0.99 | Correlation coefficient (r)>0.990 |
| Precision | CV= 4.36% | Coefficient of variation (CV) ≤10% |