Table S1. HPLC-MS/MS and UPLC-MS/MS operating conditions

| Parameters | Conditions |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| LC type | HPLC |  |  | UPLC |  |  |
| Injection volume (μL) | 10 |  |  | 5 |  |  |
| Column temperature (℃) | 40 |  |  | 40 |  |  |
| Flow rate (mL/min) | 0.2 |  |  | 0.2 |  |  |
| Column type | OSAKA SODA capcell core C18 column (2.1 × μ 150 mm, 2.7 μm) |  |  | Waters BEH C18 column (2.1 × μ 100 mm, 1.7 μm) |  |  |
| Gradient | Time (min) | A (%) | B (%) | Time (min) | A (%) | B (%) |
|  | 0.0 | 95 | 5 | 0.5 | 95 | 5 |
|  | 1.0 | 95 | 5 | 9.5 | 50 | 50 |
|  | 3.0 | 40 | 60 | 13.5 | 10 | 90 |
|  | 13.0 | 0 | 100 | 17.0 | 2 | 98 |
|  | 18. | 0 | 100 | 17.1 | 95 | 5 |
|  | 18.1 | 95 | 5 | 20.0 | 95 | 5 |
|  | 27.0 | 95 | 5 |  |  |  |
| MSMS |  |  |  |  |  |  |
| Polarity | Positive mode |  |  | Positive mode |  |  |
| Curtain Gas | 35 |  |  | 20 |  |  |
| Collision Gas | 10 |  |  | 7 |  |  |
| Ionspray voltage (V) | 5500 |  |  | 5500 |  |  |
| Temperature (℃) | 500 |  |  | 450 |  |  |
| Ion source gas 1 | 50 |  |  | 50 |  |  |
| Ion source gas 2 | 50 |  |  | 60 |  |  |

Table S2. Multi-reaction monitoring (MRM) conditions for analyzing 71 pesticides using HPLC-MS/MS

| Pesticides | Rt (min) | Quantification ion (m/z) | CE | Confirmation ion (m/z) | CE | DP |
| --- | --- | --- | --- | --- | --- | --- |
| Acetamiprid | 7.91 | 223.0>126.0 | 31 | 223.0 > 90.0 | 50 | 41 |
| Acibenzolar-S-methyl | 12.29 | 210.9>136.1 | 48 | 210.9 > 140.0 | 38 | 48 |
| Ametoctradin | 14.03 | 276.2>149.1 | 55 | 276.2 > 176.1 | 55 | 71 |
| Amisulbrom | 14.65 | 465.9>227 | 30 | 465.9> 148.1 | 65 | 38 |
| Azoxystrobin | 11.06 | 404.1>372.2 | 21 | 404.1> 344.2 | 35 | 31 |
| Bixafen | 12.91 | 414.0>394.0 | 23 | 414.0> 266.0 | 35 | 43 |
| Bromacil | 9.48 | 261.0>205.0 | 24 | 261.0> 187.9 | 40 | 29 |
| Carpropamide | 13.34 | 334.0>139.0 | 30 | 334.0> 196.0 | 20 | 39 |
| Chlorantraniliprole | 10.83 | 483.9>452.9 | 25 | 483.9> 285.9 | 35 | 30 |
| Chlorfluazuron | 16.11 | 539.9>383.0 | 31 | 539.9> 347.0 | 50 | 55 |
| Chloridazon | 8.09 | 222.0>104.0 | 35 | 222.0> 92.0 | 39 | 56 |
| Chlorobenzuron | 13.31 | 309.0>156.0 | 35 | 309.0> 139.0 | 49 | 35 |
| Chlorotoluron | 10.23 | 213.0>72.0 | 37 | 213.0> 46.0 | 30 | 38 |
| Chloroxuron | 12.2 | 291.0>72.0 | 49 | 291.0> 46.0 | 37 | 40 |
| Chromafenozide | 12.24 | 395.2>175.1 | 35 | 395.2> 339.2 | 14 | 21 |
| Clomeprop | 15.03 | 324.0>203.0 | 25 | 324.0> 120.1 | 33 | 43 |
| Clothianidin | 7.72 | 250.0>132.0 | 27 | 250.0> 169.0 | 21 | 30 |
| Cyantraniliprole | 9.7 | 475.0>285.9 | 35 | 475.0> 112.1 | 70 | 40 |
| Cyclaniliprole | 12.77 | 599.9>284.0 | 39 | 599.9> 177.0 | 70 | 37 |
| Cycloate | 14.38 | 216.1>154.1 | 20 | 216.1> 134.1 | 22 | 34 |
| Cyflumetofen | 14.48 | 465.1>173.1 | 33 | 465.1> 249.0 | 20 | 42 |
| Cymoxanil | 8.28 | 199.0>128.0 | 15 | 199.0> 83.0 | 35 | 22 |
| Diflubenzuron | 12.92 | 311.0>141.0 | 55 | 311.0> 158.0 | 25 | 38 |
| Diuron | 10.04 | 233.0>72.0 | 37 | 233.0> 46.1 | 34 | 41 |
| Famoxadone | 13.45 | 392.1>331.1 | 15 | 392.1> 238.1 | 30 | 24 |
| Fenhexamid | 12.28 | 302.0>97.1 | 37 | 302.0> 55.0 | 69 | 57 |
| Fenpyroximate | 16.11 | 422.2>366.2 | 25 | 422.2> 107.0 | 70 | 40 |
| Fludioxonil | 11.61 | 266.0>229.0 | 22 | 266.0> 158.0 | 50 | 21 |
| Fluometuron | 10.03 | 233.0>46.0 | 33 | 233.0> 160.0 | 37 | 41 |
| Flupoxam | 12.74 | 461.0>444.1 | 32 | 461.0> 123.1 | 69 | 70 |
| Fluxametamide | 15.34 | 474.0>399.9 | 28 | 474.0> 159.9 | 55 | 40 |
| Imidacloprid | 7.62 | 256.0>209.1 | 26 | 256.0> 175.0 | 30 | 35 |
| Isoproturon | 10.49 | 207.1>72.0 | 37 | 207.1> 165.0 | 22 | 40 |
| Isoxaben | 11.59 | 333.1>165.1 | 35 | 333.1> 107.0 | 75 | 41 |
| Lenacil | 10.47 | 235.1>153.1 | 25 | 235.1> 136.0 | 50 | 33 |
| Linuron | 11.58 | 249.0>160.0 | 29 | 249.0>182.1 | 25 | 39 |
| Mandestrobin | 13.4 | 314.1>192.0 | 17 | 314.1>132.0 | 36 | 25 |
| Mandipropamid | 11.47 | 412.1>328.0 | 22 | 412.1>125.0 | 60 | 42 |
| Metaflumizone_E | 14.74 | 507.1>287.0 | 35 | 507.1>178.0 | 35 | 60 |
| Metaflumizone_Z | 14.74 | 507.1>178.0 | 40 | 507.1>287.0 | 39 | 56 |
| Methabenzthiazuron | 10.51 | 222.0>165.0 | 26 | 222.0>150.0 | 50 | 32 |
| Methoxyfenozide | 11.83 | 369.2>149.1 | 27 | 369.2>313.0 | 15 | 20 |
| Metominostrobin | 10.68 | 285.1>238.0 | 19 | 285.1>194.1 | 29 | 30 |
| Neburon | 13.08 | 275.0>88.1 | 26 | 275.0>114.1 | 24 | 45 |
| Nitenpyram | 6.89 | 271.0>225.1 | 20 | 271.0>126.0 | 37 | 32 |
| Norflurazon | 10.65 | 304.0>140.1 | 59 | 304.0>145.0 | 69 | 60 |
| Orysastrobin | 11.74 | 392.1>205.0 | 21 | 392.1>116 | 40 | 38 |
| Pebulate | 14.19 | 204.1>128.1 | 19 | 204.1>72.0 | 24 | 31 |
| Pencycuron | 13.93 | 329.1>125.0 | 35 | 329.1>218.1 | 24 | 43 |
| Phosfolan | 8.52 | 256.0>140.0 | 35 | 256.0>228.0 | 20 | 39 |
| Picarbutrazox | 12.69 | 410.1>310.0 | 20 | 410.1>107.1 | 40 | 19 |
| Picolinafen | 15.1 | 377.0>238.0 | 42 | 377.0>359.0 | 30 | 50 |
| Proquinazid | 16.64 | 373.0>330.9 | 24 | 373.0>288.9 | 35 | 40 |
| Pyraclonil | 9.88 | 315.1>169.1 | 40 | 315.1>241.0 | 31 | 48 |
| Pyraclostrobin | 13.7 | 388.1>194.1 | 20 | 388.1>163.1 | 31 | 30 |
| Pyrazolate | 13.93 | 439.0>91.0 | 70 | 439.0>172.9 | 30 | 50 |
| Pyridaben | 16.45 | 365.1>309.1 | 20 | 365.1>147.1 | 35 | 38 |
| Pyridate_(a) | 16.97 | 379.1>207.0 | 30 | 379.1>351.1 | 19 | 39 |
| Pyriproxyfen_(a) | 15.68 | 322.1>96.0 | 26 | 322.1>185.1 | 33 | 40 |
| Saflufenacil | 10.91 | 501.0>349.0 | 39 | 501.0>198.0 | 60 | 55 |
| Sulfoxaflor | 8.12 | 278.0>174.1 | 23 | 278.0>154.0 | 39 | 32 |
| Tebufenozide | 12.79 | 353.2>297.2 | 15 | 353.2>133.1 | 30 | 19 |
| Tebuthiuron | 9.53 | 229.1>172.1 | 27 | 229.1>116.0 | 39 | 40 |
| Thiacloprid | 8.19 | 253.0>126.0 | 35 | 253.0>90.0 | 55 | 42 |
| Thiamethoxam | 7.23 | 292.0>211.1 | 20 | 292.0>181.0 | 33 | 30 |
| Thidiazuron | 9.39 | 221.0>102.0 | 25 | 221.0>128.0 | 25 | 40 |
| Tolfenpyrad | 15.12 | 384.1>197.0 | 40 | 384.1>154.0 | 60 | 60 |
| Triflumuron | 13.69 | 359.0>156.0 | 25 | 359.0>139.0 | 50 | 40 |
| Vernolate | 14.19 | 204.1>128.0 | 17 | 204.1>86.0 | 23 | 30 |
| BTS 44595 | 13.6 | 325.0>282.1 | 26 | 325.0>129.1 | 26 | 42 |
| BTS 44596 | 13.51 | 353.0>308.1 | 20 | 353.0>70.1 | 40 | 39 |

Table S3. Multi-reaction monitoring (MRM) conditions for analyzing 71 pesticides using UPLC-MS/MS

| Pesticides | Rt (min) | Quantification ion (m/z) | CE | Confirmation ion | CE | DP | CXP |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Acetamiprid | 10.75 | 223.0>126.0 | 26 | 223.0>90.0 | 40 | 45 | 10 |
| Acibenzolar-S-methyl | 15.2 | 210.9>136.1 | 39 | 210.9>140.0 | 31 | 15 | 4 |
| Ametoctradin | 16.33 | 276.2>149.1 | 47 | 276.2>176.1 | 60 | 50 | 10 |
| Amisulbrom | 16.35 | 465.9>227.0 | 29 | 465.9>148.1 | 68 | 55 | 14 |
| Azoxystrobin | 14.77 | 404.1>372.2 | 20 | 404.1>344.2 | 34 | 36 | 6 |
| Bixafen | 15.6 | 414.0>394.0 | 20 | 414.0>266.0 | 31 | 50 | 16 |
| Bromacil | 13.33 | 261.0>205.0 | 17 | >261.0>187.9 | 36 | 30 | 6 |
| Carpropamide | 15.89 | 334.0>139.0 | 22 | 334.0>196.0 | 17 | 50 | 14 |
| Chlorantraniliprole | 14.56 | 483.9>452.9 | 24 | 483.9>285.9 | 20 | 45 | 18 |
| Chlorfluazuron | 17.14 | 539.9>383.0 | 34 | 539.9>347.0 | 55 | 70 | 14 |
| Chloridazon | 10.81 | 222.0>104.0 | 29 | 222.0>92.0 | 33 | 55 | 6 |
| Chlorobenzuron | 15.75 | 309.0>156.0 | 19 | 309.0>139.0 | 40 | 35 | 14 |
| Chlorotoluron | 14.1 | 213.0>72.0 | 35 | 213.0>46.0 | 30 | 48 | 4 |
| Chloroxuron | 15.31 | 291.0>72.0 | 45 | 291.0>46.0 | 45 | 50 | 4 |
| Chromafenozide | 15.36 | 395.2>175.1 | 22 | 395.2>339.2 | 10 | 30 | 12 |
| Clomeprop | 16.62 | 324.0>203.0 | 21 | 324.0>120.1 | 21 | 50 | 6 |
| Clothianidin | 9.94 | 250.0>132.0 | 25 | 250.0>169.0 | 28 | 15 | 10 |
| Cyantraniliprole | 13.76 | 475.0>285.9 | 18 | 475.0>112.1 | 80 | 40 | 12 |
| Cyclaniliprole | 15.55 | 599.9>284.0 | 21 | 599.9>177.0 | 65 | 45 | 14 |
| Cycloate | 16.39 | 216.1>154.1 | 17 | 216.1>134.1 | 22 | 35 | 14 |
| Cyflumetofen | 16.28 | 465.1>173.1 | 30 | 465.1>249.0 | 19 | 21 | 10 |
| Cymoxanil | 11.18 | 199.0>128.0 | 15 | 199.0>83.0 | 37 | 51 | 10 |
| Diflubenzuron | 15.57 | 311.0>141.0 | 50 | 311.0>158.0 | 22 | 45 | 10 |
| Diuron | 13.97 | 233.0>72.0 | 35 | 233.0>46.1 | 35 | 45 | 4 |
| Famoxadone | 15.81 | 392.1>331.1 | 14 | 392.1>238.1 | 24 | 15 | 12 |
| Fenhexamid | 15.35 | 302.0>97.1 | 30 | 302.0>55.0 | 65 | 52 | 10 |
| Fenpyroximate | 17.31 | 422.2>366.2 | 25 | 422.2>107.0 | 75 | 30 | 14 |
| Fludioxonil | 15.05 | 266.0>229.0 | 19 | 266.0>158.0 | 45 | 20 | 4 |
| Fluometuron | 13.96 | 233.0>46.0 | 32 | 233.0>160.0 | 36 | 50 | 12 |
| Flupoxam | 15.57 | 461.0>444.1 | 37 | 461.0>123.1 | 75 | 80 | 14 |
| Fluxametamide | 16.71 | 474.0>399.9 | 26 | 474.0>159.9 | 53 | 35 | 16 |
| Imidacloprid | 9.86 | 256.0>209.1 | 21 | 256.0>175.0 | 29 | 45 | 6 |
| Isoproturon | 14.37 | 207.1>72.0 | 33 | 207.1>165.0 | 21 | 50 | 4 |
| Isoxaben | 15.05 | 333.1>165.1 | 25 | 333.1>107.0 | 75 | 35 | 10 |
| Lenacil | 14.38 | 235.1>153.1 | 21 | 235.1>136.0 | 43 | 31 | 10 |
| Linuron | 14.92 | 249.0>160.0 | 25 | 249.0>182.1 | 22 | 35 | 14 |
| Mandestrobin | 15.86 | 314.1>192.0 | 15 | 314.1>132.0 | 35 | 28 | 12 |
| Mandipropamid | 14.95 | 412.1>328.0 | 21 | 412.1>125.0 | 55 | 60 | 10 |
| Metaflumizone_E | 16.36 | 507.1>287.0 | 34 | 507.1>178.0 | 31 | 80 | 12 |
| Metaflumizone_Z | 16.36 | 507.1>178.0 | 31 | 507.1>287.0 | 33 | 70 | 12 |
| Methabenzthiazuron | 14.32 | 222.0>165.0 | 22 | 222.0>150.0 | 43 | 39 | 10 |
| Methoxyfenozide | 15.14 | 369.2>149.1 | 21 | 369.2>313.0 | 10 | 20 | 10 |
| Metominostrobin | 14.49 | 285.1>238.0 | 15 | 285.1>194.1 | 26 | 45 | 8 |
| Neburon | 15.72 | 275.0>88.1 | 20 | 275.0>114.1 | 20 | 35 | 6 |
| Nitenpyram | 7.56 | 271.0>225.1 | 16 | 271.0>126.0 | 35 | 35 | 10 |
| Norflurazon | 14.49 | 304.0>140.1 | 47 | 304.0>145.0 | 65 | 50 | 10 |
| Orysastrobin | 15.1 | 392.1>205.0 | 21 | 392.1>116.0 | 40 | 39 | 8 |
| Pebulate | 16.28 | 204.1>128.1 | 15 | 204.1>72.0 | 18 | 41 | 12 |
| Pencycuron | 16.12 | 329.1>125.0 | 26 | 329.1>218.1 | 23 | 65 | 10 |
| Phosfolan | 12.14 | 256.0>140.0 | 37 | 256.0>228.0 | 17 | 35 | 4 |
| Picarbutrazox | 15.49 | 410.1>310.0 | 19 | 410.1>107.1 | 31 | 30 | 10 |
| Picolinafen | 16.61 | 377.0>238.0 | 35 | 377.0>359.0 | 29 | 65 | 8 |
| Proquinazid | 17.63 | 373.0>330.9 | 20 | 373.0>288.9 | 33 | 28 | 12 |
| Pyraclonil | 14.02 | 315.1>169.1 | 40 | 315.1>241.0 | 32 | 55 | 10 |
| Pyraclostrobin | 15.97 | 388.1>194.1 | 17 | 388.1>163.1 | 33 | 48 | 14 |
| Pyrazolate | 16.08 | 439.0>91.0 | 64 | 439.0>172.9 | 30 | 60 | 6 |
| Pyridaben | 17.59 | 365.1>309.1 | 18 | 365.1>147.1 | 31 | 25 | 16 |
| Pyridate_(a) | 17.95 | 379.1>207.0 | 25 | 3779.1>351.1 | 15 | 20 | 6 |
| Pyriproxyfen_(a) | 16.82 | 322.1>96.0 | 20 | 322.1>185.1 | 32 | 31 | 10 |
| Saflufenacil | 14.59 | 501.0>349.0 | 37 | 501.0>198.0 | 60 | 76 | 6 |
| Sulfoxaflor | 11.11 | 278.0>174.1 | 15 | 278.0>154.0 | 40 | 50 | 12 |
| Tebufenozide | 15.62 | 353.2>297.2 | 12 | 353.2>133.1 | 37 | 20 | 8 |
| Tebuthiuron | 13.56 | 229.1>172.1 | 25 | 229.1>116.0 | 37 | 50 | 12 |
| Thiacloprid | 11.63 | 253.0>126.0 | 27 | 253.0>90.0 | 51 | 45 | 12 |
| Thiamethoxam | 8.49 | 292.0>211.1 | 17 | 292.0>181.0 | 30 | 42 | 12 |
| Thidiazuron | 13.23 | 221.0>102.0 | 20 | 221.0>128.0 | 21 | 48 | 8 |
| Tolfenpyrad | 16.63 | 384.1>197.0 | 34 | 384.1>154.0 | 55 | 95 | 12 |
| Triflumuron | 15.9 | 359.0>156.0 | 25 | 359.0>139.0 | 55 | 45 | 12 |
| Vernolate | 16.27 | 204.1>128.0 | 15 | 204.1>86.0 | 19 | 25 | 8 |
| BTS 44595 | 16.01 | 325.0>282.1 | 22 | 325.0>129.1 | 22 | 48 | 10 |
| BTS 44596 | 15.94 | 353.0>308.1 | 20 | 353.0>70.1 | 47 | 45 | 16 |