## Tables and figures

Table 1. Checklist - key performance metrics, study design elements, calibration specifications, reporting standards, and safety measures for CMG devices in Latin America

| Criteria |  | Yes | No | Comment |
| --- | --- | --- | --- | --- |
| 1. Accuracy criteria |  |  |  |  |
| 1 | Precision and accuracy are reported using the following metrics: Mean Absolute Relative Difference (MARD), consensus error grid, and percent of results within a defined range vs reference |  |  |  |
| 1.1. | MARD: The device demonstrates a Mean Absolute Relative Difference (MARD) < 10% when compared with venous blood glucose |  |  |  |
| 1.1.1. | MARD is reported throughout the sensor’s operational lifespan, considering different glucose thresholds such as hypoglycemia, hyperglycemia, and Time in Range (TIR) |  |  |  |
| 1.2. | Consensus error grid: Report results in Zones A and B, in which have no effect on the clinical action or the action would have little effect on the clinical outcome |  |  |  |
| 1.3. | Agreement rate: Agreement thresholds for iCGM accuracy: Each item below specifies the minimum percentage of acceptable agreement relative to blood glucose values across defined glycemic ranges(5,14): iCGM special control performance requirement (all with 95% LCL*): - A: Within 15 mg/dL, <70 mg/dL: >85% - B: Within 15%, 70–180 mg/dL: >70% - C: Within 15%, >180 mg/dL: >80% - D: Within 40 mg/dL, <70 mg/dL: >98% - E: Within 40%, 70–180 mg/dL: >99% - F: Within 40%, >180 mg/dL: >99% - G: Within 20%: >87% - H: CGM <70 mg/dL & Ref** >180 mg/dL: 0 - I: CGM >180 mg/dL & Ref <70 mg/dL: 0 - J: CGM ROC*** >1 mg/dL/min & Ref ROC <–2 mg/dL/min: <1% - K: CGM ROC <–1 mg/dL/min & Ref ROC >2 mg/dL/min: <1% *LCL: lower confidence limit, **Ref: reference glucose value, ***ROC: rate of change (glucose rate of change in mg/dL per minute) |  |  |  |
| 2. Design of performance studies |  |  |  |  |
| 2.1. | Population Representation: Performance is validated in diverse populations, including pediatric, adult, older adults, and pregnant individuals*. |  |  |  |
| 2.2. | Anatomical Sites: Performance is reported for all intended anatomical application sites, including paired readings. |  |  |  |
| 2.3. | Study Setting: Performance studies clearly describe whether the study was conducted in-clinic, free-living, or hybrid setting. |  |  |  |
| 2.4. | Comparator measurement: Sampling compartment (capillary, venous, or arterialized-venous) is clearly defined. |  |  |  |
| 2.5. | Reporting: Performance studies are published in indexed, peer-reviewed journals. |  |  |  |
| 2.6. | RWE: If approved, performance is reported for everyday use, including high-stress environments (e.g., hospitals). |  |  |  |
| 3. Calibration |  |  |  |  |
| 3.1. | Factory-calibrated: The device is factory-calibrated and calibration remains accurate throughout the sensor’s lifespan. |  |  |  |
| 3.2. | Calibration method: If the device requires manual calibration, the calibration method is explicitly described and precise instructions for healthcare professionals and patients on how to perform manual calibration is provided. |  |  |  |
| 4. Report generation and data sharing |  |  |  |  |
| 4.1. | Standardized reports: The report adheres to the recommendations by the 2019 international ATTD consensus (Battelino et al. 2019 [14]). |  |  |  |
| 4.2. | Interoperability: The device integrates seamlessly with smartphones, insulin pumps, and other diabetes management devices. |  |  |  |
| 4.3. | Privacy compliance: The device adheres to local data privacy regulations. |  |  |  |
| 5. User support, post-sale services, and patient safety |  |  |  |  |
| 5.1. | Technical support: Manufacturer provides omnichannel customer support in every market where the device is available. |  |  |  |
| 5.2. | Warranty and repairs: Manufacturer ensure robust post-sale policies, including user-friendly educational resources, warranties, and device replacement protocols. |  |  |  |
| 5.3. | Biocompatibility: The device complies with standards and explicitly reports materials, potential allergens, and interactions with medications or other devices. |  |  |  |
| 5.4. | Customizable alerts: The device allows users to customize alarm thresholds and settings based on individual needs. |  |  |  |
| 5.4.1. | Alerts metrics: True alert, false alert, missed alert, and correct alert rates are provided. |  |  |  |
| 5.5. | Signal loss: The device includes alerts for signal loss. |  |  |  |

*The population representation is for the intended approval.