On-Line Electronic Supplement

Intracycle Power and Ventilation Mode as Potential Contributors to

Ventilator-Induced Lung Injury

Section E1:  Mathematical Model Development

A mathematical model was used to simulate ventilation into and out of a simplified one-compartment respiratory system. Total duration of each breath (ttot) was subdivided into an inspiratory phase (0≤t≤ti) and an expiratory phase (ti≤t≤ttot). We assumed ‘lumped’ but clinically-measurable mechanical characteristics for resistance (R) and compliance (C).  In mathematical terms, these assumptions translate into the equations:

Presistive=R∙flow=RdVdt

Pelastic=E∙volume=1CVt.

Here, the inflation volume  V(t) delivered by the ventilator is defined as zero at the beginning and the ending of the breath.

The “equations of motion” build upon this base by assuming that the applied pressure at the airway opening is balanced by pressures developing in response to resistance and elastance (inverse of compliance) as well as the residual alveolar pressure value at end-expiration, defined as Presidual=Pex=total PEEP, the sum of PEEP and autoPEEP. Thus:

Presistive+Pelastic+Presidual=Papplied.

These assumptions lead to the following mathematical model comprised of a coupled pair of initial-value problems for the unknown volume functions for inflation and expiration Vit and Vet, respectively.

Inspiration (i):
RdVidt+ViC+Pex=Paw , 0≤t≤ti

Vi0=0

*Expiration (e):

RdVedt+VeC+Pex=PEEP, ti≤t≤ttot

Veti=Vi(ti)=VT.

*(During expiration flow is assumed to be a negative quantity.)

Each differential equation in the mathematical model can be solved for Vit and Ve(t).  For example, in the case of pressure-controlled ventilation:

Vit=CPex1-e-tCR+CPset1-e-tCR.

Our simplified mathematical model emulates ventilation-related variables that result from different ventilatory patterns (flow waveforms). In addition to the sinusoidal flow pattern of normal spontaneous breathing (SF), we studied model predictions for three popular modes of ventilation: (1) constant inflation pressure ventilation (Pset); (2) constant flow ventilation (CF); and (3) linearly decelerating  flow ventilation (DF).

Section E2:  Intracycle Power Definitions

The total intracycle total power can be considered as the sum of its elastic PEEP and resistive components.

We define the IntraCycle Driving Power (ICPD) as the product of tidal elastic lung pressure above total PEEP and flow (Q).  That is,

ICPDt=1CVit∙Qit=1CVitdVidtt.

Similarly,  we define the IntraCycle Elastic Power (ICPE), a product of the values at inspiratory time t for dynamic alveolar pressure (driving pressure plus PEEP) and flow, as:

ICPEt=1CVit+PexQit=1CVit+PexdVidt(t)

and the IntraCycle Total Power (ICPT) as the sum of the values at inspiratory time t for flow resistive power and ICPE :

ICPTt=RQit+1CVit+PexQit=RdVidtt+1CVit+PexdVidtt.

Section E3:  Comparison of Mathematical Model Predictions and Behaviors of the Mechanical (Physical) Simulation for Three Prototypical Scenarios

The predictions of the mathematical model were compared with experiments performed on a servo-actuated lung simulator (Ingmar ASL-5000) driven by a commercial ventilator.  Ventilator settings: tidal volume, inspiratory time, frequency, and PEEP and mode of ventilation (Pset, CF, or DF), as well as properties for patient lung mechanics (compliance and resistance), were chosen to resemble ARDS, airway obstruction (COPD), and normal lung scenarios (Table E1).  Measurements of time, volume, flow, airway pressure in and out of the mechanical lung were recorded and compared against the mathematical model.  In doing this, our goal was to

Determine if the mathematical model produces accurate predictions for volume and flow into the mechanical lung during inspiration, as these two quantities are used to compute values for the intracycle power functions,  e.g., ICPTt=RQit+VitC+Pex , 0≤t≤ti.

The ventilator-lung simulator experiments were performed for a wide variety of ventilator settings and compliance-resistance combinations using 3 modes of ventilation.  For example, with ventilator settings of ti=1.0 s, ttot=3.75 s, VT=0.35 L, frequency = 16/min PEEP=8.0 cmH2O and patient parameters of C=0.04 L/cmH2O, R=8.0cmH2OL/s with CF ventilation, we have the following comparison between the mechanical lung outputs (red dots) and mathematical model predictions (blue curve) for the volumes (L.) generated during inspiration and expiration:

The mathematical model predicted the total end-expiratory pressure (including auto-PEEP) as Pex=8.023 cmH2O. The maximum separation ‘error’ between the simulation and model was 0.032 L and the minimum error was 5.511 x 10-16 L  As one can readily observe, the maximum difference between the model and simulator volumes (liters) occurs during expiration. For CF and DF ventilation experiments, maximum errors were observed generally during expiration.  Using DF ventilation, another experiment produced these results:

The maximum separation here is 0.039 L.,  again occurring during the expiratory portion of the breath.

The most noteworthy differences occurred using Pset ventilation.  This can be seen in the following example of Pset ventilation.

The maximum error is 0.042 L and occurs during inflation.  The central problem with Pset ventilation experiments lies with the airway pressure actually being delivered by the ventilator during inspiration.  In constant pressure ventilation (Pset), the mathematical model assumes that the airway pressure is constant throughout inspiration.  If one examines the airway pressure developed by the ventilator during inspiration, however, it only becomes constant some brief time after inflation onset.  This can be seen in the following tracing of airway pressure (cmH2O), taken from the above example.

One can see about 10-15% of the inspiratory time is spent building the airway pressure to its pre-set value.  By design, the ventilator tends to generate lower than maximal pressures at the beginning of inspiration and therefore delivers lower flows (and hence, smaller volumes).  This explains why the red dots of the simulator volumes undershoot the mathematical model’s predicted volume curve.  The tested simulator settings and virtual patient parameters are shown in Table E2. In summary, if the mathematical model is provided accurate information about its inputs, then its predictions are very consistent with those experienced by the mechanical ventilator-simulator system.

Table E2  Conditions tested using the ventilator-simulator system

| Exp # | Mode | Flow Waveform | PEEP (cm H2O) | f Rate (bp) | VT (ml) |  | TI (sec) |  | TR (sec) | Flow Ramp (%) | C (L/cm H2O) | R (cm H2O/L/s) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | VC-CMV | Constant | 16 | 30 |  | 300 |  | 0.8 | 0.0 | 100 | 0.015 | 10 |
| 2 | VC-CMV | Constant | 10 | 30 |  | 300 |  | 0.8 | 0.0 | 100 | 0.015 | 15 |
| 3 | VC-CMV | Constant | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 100 | 0.040 | 8 |
| 4 | VC-CMV | Constant | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 100 | 0.060 | 8 |
| 5 | VC-CMV | Constant | 8 | 14 |  | 500 |  | 1.0 | 0.0 | 100 | 0.080 | 10 |
| 6 | VC-CMV | Constant | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 100 | 0.100 | 20 |
| 7 | VC-CMV | Constant | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 100 | 0.100 | 30 |
| 8 | VC-CMV | Descending Ramp 0% | 16 | 30 |  | 300 |  | 0.8 | 0.0 | 0 | 0.015 | 10 |
| 9 | VC-CMV | Descending Ramp 0% | 10 | 30 |  | 300 |  | 0.8 | 0.0 | 0 | 0.015 | 15 |
| 10 | VC-CMV | Descending Ramp 0% | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 0 | 0.040 | 8 |
| 11 | VC-CMV | Descending Ramp 0% | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 0 | 0.060 | 8 |
| 12 | VC-CMV | Descending Ramp 0% | 8 | 14 |  | 500 |  | 1.0 | 0.0 | 0 | 0.080 | 10 |
| 13 | VC-CMV | Descending Ramp 0% | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 0 | 0.100 | 20 |
| 14 | VC-CMV | Descending Ramp 0% | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 0 | 0.100 | 30 |
| 15 | VC-CMV | Descending Ramp 30% | 16 | 30 |  | 300 |  | 0.8 | 0.0 | 30 | 0.015 | 10 |
| 16 | VC-CMV | Descending Ramp 30% | 10 | 30 |  | 300 |  | 0.8 | 0.0 | 30 | 0.015 | 15 |
| 17 | VC-CMV | Descending Ramp 30% | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 30 | 0.040 | 8 |
| 18 | VC-CMV | Descending Ramp 30% | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 30 | 0.060 | 8 |
| 19 | VC-CMV | Descending Ramp 30% | 8 | 14 |  | 500 |  | 1.0 | 0.0 | 30 | 0.080 | 10 |
| 20 | VC-CMV | Descending Ramp 30% | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 30 | 0.100 | 20 |
| 21 | VC-CMV | Descending Ramp 30% | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 30 | 0.100 | 30 |
| 22 | VC-CMV | Descending Ramp 50% | 16 | 30 |  | 300 |  | 0.8 | 0.0 | 50 | 0.015 | 10 |
| 23 | VC-CMV | Descending Ramp 50% | 10 | 30 |  | 300 |  | 0.8 | 0.0 | 50 | 0.015 | 15 |
| 24 | VC-CMV | Descending Ramp 50% | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 50 | 0.040 | 8 |
| 25 | VC-CMV | Descending Ramp 50% | 8 | 16 |  | 400 |  | 1.0 | 0.0 | 50 | 0.060 | 8 |
| 26 | VC-CMV | Descending Ramp 50% | 8 | 14 |  | 500 |  | 1.0 | 0.0 | 50 | 0.080 | 10 |
| 27 | VC-CMV | Descending Ramp 50% | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 50 | 0.100 | 20 |
| 28 | VC-CMV | Descending Ramp 50% | 5 | 12 |  | 500 |  | 1.2 | 0.0 | 50 | 0.100 | 30 |
| 29 | PC-CMV | Exponential Decay PR0 | 16 | 30 |  | 300 |  | 0.8 | 0.0 | NA | 0.015 | 10 |
| 30 | PC-CMV | Exponential Decay PR0 | 10 | 30 |  | 300 |  | 0.8 | 0.0 | NA | 0.015 | 15 |
| 31 | PC-CMV | Exponential Decay PR0 | 8 | 16 |  | 400 |  | 1.0 | 0.0 | NA | 0.040 | 8 |
| 32 | PC-CMV | Exponential Decay PR0 | 8 | 16 |  | 400 |  | 1.0 | 0.0 | NA | 0.060 | 8 |
| 33 | PC-CMV | Exponential Decay PR0 | 8 | 14 |  | 500 |  | 1.0 | 0.0 | NA | 0.080 | 10 |
| 34 | PC-CMV | Exponential Decay PR0 | 5 | 12 |  | 500 |  | 1.2 | 0.0 | NA | 0.100 | 20 |
| 35 | PC-CMV | Exponential Decay PR0 | 5 | 12 |  | 500 |  | 1.2 | 0.0 | NA | 0.100 | 30 |
| 36 | PC-CMV | Exponential Decay PR0.4 | 16 | 30 |  | 300 |  | 0.8 | 0.4 | NA | 0.015 | 10 |
| 37 | PC-CMV | Exponential Decay PR0.4 | 10 | 30 |  | 300 |  | 0.8 | 0.4 | NA | 0.015 | 15 |
| 38 | PC-CMV | Exponential Decay PR0.4 | 8 | 16 |  | 400 |  | 1.0 | 0.4 | NA | 0.040 | 8 |
| 39 | PC-CMV | Exponential Decay PR0.4 | 8 | 16 |  | 400 |  | 1.0 | 0.4 | NA | 0.060 | 8 |
| 40 | PC-CMV | Exponential Decay PR0.4 | 8 | 14 |  | 500 |  | 1.0 | 0.4 | NA | 0.080 | 10 |
| 41 | PC-CMV | Exponential Decay PR0.4 | 5 | 12 |  | 500 |  | 1.2 | 0.4 | NA | 0.100 | 20 |
| 42 | PC-CMV | Exponential Decay PR0.4 | 5 | 12 |  | 500 |  | 1.2 | 0.4 | NA | 0.100 | 30 |

Table E2: Ventilating settings, physiological parameters, and ventilation modes used in 42 mechanical simulator experiments.  For decelerating flow (Descending Ramp), the end-inflation flow rate was either 0%, 30% or 50% of the peak initial flow rate. Pressure Control (Pset, PC-CMV) produced an exponential decay and was tested with two approaches to the targeted value: PRO and PRO.4. These corresponded to rapid and slow ‘attack’ rates to the targeted pressure, respectively.

Section E4  Tested Ranges for Normal, Obstructed (COPD) and ARDS ‘Virtual’ Patients

Table E1        Virtual Patients : Sample Generation and Parameter Ranges

# 5,000 Normal    Lung Samples

| Statistic/Parameter | Ti | Ttot | Vt | C | R | PEEP |
| --- | --- | --- | --- | --- | --- | --- |
| Mean | 0.898 | 2.241 | 0.524 | 0.100 | 9.984 | 5.042 |
| SD | 0.172 | 0.527 | 0.131 | 0.011 | 1.152 | 2.884 |
| Max | 1.200 | 3.390 | 0.750 | 0.120 | 11.999 | 9.992 |
| Min | 0.600 | 1.102 | 0.300 | 0.080 | 8.000 | 0.006 |

# 5,000 Severe Obstruction Samples

| Statistic/Parameter | Ti | Ttot | Vt | C | R | PEEP |
| --- | --- | --- | --- | --- | --- | --- |
| Mean | 1.199 | 2.397 | 0.499 | 0.085 | 26.082 | 5.018 |
| SD | 0.232 | 0.326 | 0.114 | 0.020 | 8.144 | 2.872 |
| Max | 1.600 | 3.185 | 0.700 | 0.120 | 39.997 | 9.999 |
| Min | 0.800 | 1.616 | 0.300 | 0.050 | 12.001 | 0.002 |

# 5,000 ARDS Samples

| Statistic/Parameter | Ti | Ttot | Vt | C | R | PEEP |
| --- | --- | --- | --- | --- | --- | --- |
| Mean | 1.105 | 2.333 | 0.368 | 0.034 | 9.970 | 11.228 |
| SD | 0.197 | 0.475 | 0.071 | 0.014 | 1.202 | 3.850 |
| Max | 1.389 | 3.281 | 0.499 | 0.060 | 11.974 | 19.764 |
| Min | 0.703 | 1.342 | 0.250 | 0.010 | 8.038 | 5.222 |

Units:  Ti (s); Ttot (s); Vt (L); C (L/cmH2O); R (cmH2O/L/sec); PEEP (cmH2O)

Section E5  Computing Intracycle Power on a Virtual Population of ARDS Patients

In order to detail the behavior of driving, elastic, and total power for ARDS patients, a virtual population of 5,000 samples was created by randomly choosing inspiratory times, frequency, tidal volume, applied PEEP, compliance, and resistance values within certain prescribed ranges that are common to ARDS patients of varying severity. In our computations, we chose the following restrictions on these parameters: 0.7≤ti≤1.4 (s), 0.25≤VT≤0.5 (L), 5.0≤PEEP≤20.0 (cmH2O), 0.01≤C≤0.06 (LcmH2O), and 8.0≤R≤12.0 (cmH2OL/s).  Lastly, ttot=ti+te (s) with 0.5≤te≤2.0.  Upon building the 5,000 ‘virtual ARDS patient’ samples, the means and standard deviations of these parameters using those restrictions were: ti=1.105±0.197,  ttot=2.333±0.475, VT=0.368±0.071, PEEP=11.228±3.850, C=0.034±0.014, and R=9.970±1.202. (Table E1)

We examined the intracycle power functions (ICPD(t), ICPE(t), ICPT(t)) on this virtual ARDS population for four modes of ventilation (Pset, CF, DF, and SF).  Before we describe the details, we note that it can be proven mathematically for our model that

0tiICPDtdt=12CVT2

and

0tiICPEtdt=12CVT2+PexVT

regardless which mode of ventilation is used. In other words, the areas under the intracycle driving power and intracycle elastic power functions are determined by only the tidal volume, compliance and end-expiratory pressure.   However, this independence for the modes of ventilation does not hold for ICPT(t).

Section E6    Intracycle power above Alveolar Pressure Threshold for Damage

In the following, we examine the areas under the intracycle power curves, not on the entire inspiratory interval, but rather on the proportion of the inspiratory interval when the elastic pressure, Pelastict=VitC+Pex, exceeds a specified threshold for ‘safe’ elastic pressure.  Since the elastic pressure is an increasing function in time, there is an inspiratory time, tep, after which the elastic pressure (ep) will be above its threshold over the interval: tep≤t≤ti.  Using this threshold pressure, we can calculate the area under the intracycle power curve (A, equivalent to what may be potentially damaging elastic energy) for all four modes of ventilation.

A=teptiICPtdt

The setup for energy and power is shown in the following generic figure, whose conceptual principle applies

equally well to total power and to its driving and dynamic elastic components.

A

The area A below the power curve (ICP) is a measure of the energy after the elastic pressure has crossed its threshold (at the time t=tep).  The value of A varies from sample to sample, as does the threshold time (tep).  However, there is an interesting fact about two intracycle power functions,  which are the driving  (AD) and elastic (AE) components of intracycle power  (ICPD and ICPE) that meet the elastic pressure threshold criterion:

AD=teptiICPDtdt

AE=teptiICPEtdt

For a given sample, each of these area types is unaffected by changes in the flow profile used on that sample.  This leaves the intracycle total power area function (AT, the sum of resistive area above the pressure threshold and AE) as a possible discriminator between the modes of ventilation.  Therefore, we considered the area:

AT=teptiICPTtdt

AT was individually computed for 5,000 virtual patient samples using each mode of ventilation.  Using an elastic pressure threshold of 20 cmH2O, the following statistics for the means of AT were found:

- For Pset ventilation, AT=0.513 with a standard deviation of 0.538

- For CF ventilation, AT=0.469 with a standard deviation of 0.457

- For DF ventilation, AT=0.475 with a standard deviation of 0.471

- For SF ventilation, AT=0.485 with a standard deviation of 0.475.

We note that in these calculations of AT, if the elastic pressure threshold was not reached during inspiration, then we set AT=0 for all modes of ventilation for that sample. The mean values of AT for samples in which the elastic pressure threshold was not achieved, (no zero areas included) would undoubtedly be considerably higher. To determine if there are statistical differences between the modes of ventilation regarding total power, pairs of modes (e.g., CF versus SF) were tested with the Mann-Whitney Test on AT, yielding the following results for their p-values:

- Pset versus CF: p-value = 0.092

- Pset versus DF: p-value = 0.120

- Pset versus SF: p-value = 0.492

- CF versus DF: p-value = 0.855

- CF versus SF: p-value = 0.277

- DF versus SF: p-value = 0.376.

Thus, with samples having AT = 0 included, the level of statistical significance (p-value = 0.05) for the area comparisons was not achieved, even though the differences between Pset and CF or DF were very close to significance and may be achievable with a larger sample size.  (This was the case with a 10,000 virtual patient sampling we conducted:  Pset versus CF (p-value = 0.0092) and Pset versus DF (p-value = 0.020).

The total power vs. time area quantity (AT) is a measure of the energy applied to the lung once the elastic pressure threshold is met.  As such it is a macro measurement.  To further understand the dynamics of the power functions, the maximum instantaneous values of intracycle power in excess of the elastic pressure criterion, ICP(t), tep≤t≤ti, were examined for each mode of ventilation.  We introduce the notation: M=maxICPt:tep≤t≤ti for each type of intracycle power function, with notation of its driving (MD) and elastic (ME) components of the total (MT) values.  As in the case of areas, if the elastic pressure threshold of 20 cmH2O was not achieved during inspiration, then we set M=0.  Calculating M (watts) for each mode of ventilation and each type of intracycle power function, the following mean and standard deviation statistics were found:

- For Pset ventilation, MD=0.551  with a standard deviation 0.825

- For CF ventilation, MD=0.430  with a standard deviation   0.427

- For DF ventilation, MD=0.322  with a standard deviation  0.330

- For SF ventilation, MD=0.430  with a standard deviation   0.437

- For Pset ventilation, ME=1.465  with a standard deviation 1.664

- For CF ventilation, ME=0.814  with a standard deviation   0.641

- For DF ventilation, ME=0.820  with a standard deviation  0.669

- For SF ventilation, ME=0.966  with a standard deviation   0.772

- For Pset ventilation, MT=2.595  with a standard deviation  3.80

- For CF ventilation, MT=0.926  with a standard deviation  0.732

- For DF ventilation, MT=1.053  with a standard deviation  0.910

- For SF ventilation, MT=1.200  with a standard deviation  0.980.

As shown below, when checking for statistical differences between the modes of ventilation for MT and its subcomponents, we find (for ARDS alone) that all comparisons of modes of ventilation are significant (p-values < 10—5), except CF versus SF for MD (p-value = 0.537) and CF versus DF for ME (p-value = 0.224).  For MT, the largest p-value was 1.8 x 10-9.  The maximum value of the intracycle power function that exceeds threshold is therefore an excellent discriminator between modes of ventilation for the ARDS state.

Section E7:  Statistical Comparisons of Supra-threshold  Energy and Power for 4 Flow Waveforms in 3 Disease States.

We examined binary statistical comparisons between the ventilation modes for A, for M, and for their intracycle subcomponents that relate to driving and elastic power.

Statistical Analysis**

NORMAL Virtual Subjects

Driving Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ------ | 0.999 | 0.999 | 0.999 |
| Constant | 0.999 | ------ | 0.999 | 0.999 |
| Decelerating | 0.999 | 0.999 | ----- | 0.999 |
| Sinusoidal | 0.999 | 0.999 | 0.999 | ----- |

Elastic Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.999 | 0.999 | 0.999 |
| Constant | 0.999 | ----- | 0.999 | 0.999 |
| Decelerating | 0.999 | 0.999 | ----- | 0.999 |
| Sinusoidal | 0.999 | 0.999 | 0.999 | ----- |

Total Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.987 | 0.997 | 0.988 |
| Constant | 0.987 | ----- | 0.985 | 0.999 |
| Decelerating | 0.997 | 0.985 | ----- | 0.985 |
| Sinusoidal | 0.988 | 0.999 | 0.985 | ----- |

Driving Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.788 | 0.999 | 0.000 |
| Constant | 0.788 | ----- | 0.796 | 0.946 |
| Decelerating | 0.999 | 0.796 | ----- | 0.859 |
| Sinusoidal | 0.999 | 0.946 | 0.859 | ----- |

Elastic Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.858 | 0.964 | 0.000 |
| Constant | 0.858 | ----- | 0.933 | 0.936 |
| Decelerating | 0.964 | 0.933 | ----- | 0.903 |
| Sinusoidal | 0.000 | 0.936 | 0.903 | ----- |

Total Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.880 | 0.964 | 0.000 |
| Constant | 0.880 | ----- | 0.953 | 0.916 |
| Decelerating | 0.964 | 0.953 | ----- | 0.907 |
| Sinusoidal | 0.000 | 0.916 | 0.907 | ----- |

** p values by Mann-Whitney test

Statistical Analysis**

OBSTRUCTED Virtual Subjects

Driving Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ------ | 0.999 | 1.000 | 0.999 |
| Constant | 0.999 | ------ | 0.999 | 0.999 |
| Decelerating | 1.000 | 0.999 | ----- | 0.999 |
| Sinusoidal | 0.999 | 0.999 | 0.999 | ----- |

Elastic Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 1.000 | 1.000 | 1.000 |
| Constant | 1.000 | ----- | 1.000 | 1.000 |
| Decelerating | 1.000 | 1.000 | ----- | 1.000 |
| Sinusoidal | 1.000 | 1.000 | 1.000 | ----- |

Total Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.987 | 0.997 | 0.988 |
| Constant | 0.987 | ----- | 0.985 | 1.000 |
| Decelerating | 0.997 | 0.985 | ----- | 0.985 |
| Sinusoidal | 0.988 | 1.000 | 0.985 | ----- |

Driving Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.788 | 1.000 | 0.000 |
| Constant | 0.788 | ----- | 0.796 | 0.946 |
| Decelerating | 1.000 | 0.796 | ----- | 0.859 |
| Sinusoidal | 0.000 | 0.946 | 0.859 | ----- |

Elastic Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.858 | 0.964 | 0.000 |
| Constant | 0.858 | ----- | 0.933 | 0.936 |
| Decelerating | 0.964 | 0.933 | ----- | 0.903 |
| Sinusoidal | 0.000 | 0.936 | 0.903 | ----- |

Total Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.881 | 0.964 | 0.000 |
| Constant | 0.881 | ----- | 0.953 | 0.916 |
| Decelerating | 0.964 | 0.953 | ----- | 0.907 |
| Sinusoidal | 0.000 | 0.916 | 0.907 | ----- |

** p values by Mann-Whitney test

Statistical Analysis**

ARDS Virtual Subjects

Driving Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ------ | 1.000 | 1.000 | 1.000 |
| Constant | 1.000 | ------ | 1.000 | 1.000 |
| Decelerating | 1.000 | 1.000 | ----- | 1.000 |
| Sinusoidal | 1.000 | 1.000 | 1.000 | ----- |

Elastic Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 1.000 | 1.000 | 1.000 |
| Constant | 1.000 | ----- | 1.000 | 1.000 |
| Decelerating | 1.000 | 1.000 | ----- | 1.000 |
| Sinusoidal | 1.000 | 1.000 | 1.000 | ----- |

Total Power Areas

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.092 | 0.130 | 0.492 |
| Constant | 0.092 | ----- | 0.855 | 0.277 |
| Decelerating | 0.130 | 0.855 | ----- | 0.376 |
| Sinusoidal | 0.492 | 0.277 | 0.376 | ----- |

Driving Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 0.000 | 8.27x10-9 | 0.000 |
| Constant | 0.000 | ----- | 6.58x10-35 | 0.537 |
| Decelerating | 8.27x10-9 | 6.58x10-35 | ----- | 6.37x10-30 |
| Sinusoidal | 0.000 | 0.537 | 6.37x10-30 | ----- |

Elastic Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 1.97x10-53 | 1.09x10-50 | 0.000 |
| Constant | 1.97x10-53 | ----- | 0.224 | 4.55x10-29 |
| Decelerating | 1.09x10-50 | 0.224 | ----- | 3.98x10-20 |
| Sinusoidal | 0.000 | 4.55x10-29 | 3.98x10-20 | ----- |

Total Power Maximum

|  | Pset | Constant | Decelerating | Sinusoidal |
| --- | --- | --- | --- | --- |
| Pset | ----- | 3.41x10-81 | 6.55x10-57 | 0.000 |
| Constant | 3.41x10-81 | ----- | 1.87x10-9 | 1.22x10-49 |
| Decelerating | 6.55x10-57 | 1.87x10-9 | ----- | 1.39x10-13 |
| Sinusoidal | 0.000 | 1.22x10-49 | 1.39x10-13 | ----- |

** p values by Mann-Whitney test

Table E1            Virtual Patients: Parameters and Sampling Ranges

# 5,000 Normal  Lung

Units:  Ti (sec.); Ttot (sec.); Vt (L.); C (L/cmH2O); R (cmH2O/L/sec); PEEP (cmH2O)

| Statistic/Parameter | Ti | Ttot | Vt | C | R | PEEP |
| --- | --- | --- | --- | --- | --- | --- |
| Mean | 0.898286 | 2.24145 | 0.524453 | 0.100078 | 9.98358 | 5.04238 |
| SD | 0.172626 | 0.526716 | 0.130647 | 0.0113831 | 1.15246 | 2.88427 |
| Max | 1.19987 | 3.39309 | 0.749906 | 0.119989 | 11.9994 | 9.99176 |
| Min | 0.600059 | 1.10193 | 0.300301 | 0.0800088 | 8.00051 | 0.00554778 |

# 5,000 COPD

Units:  Ti (sec.); Ttot (sec.); Vt (L.); C (L/cmH2O); R (cmH2O/L/sec); PEEP (cmH2O)

| Statistic/Parameter | Ti | Ttot | Vt | C | R | PEEP |
| --- | --- | --- | --- | --- | --- | --- |
| Mean | 1.19906 | 2.39651 | 0.498688 | 0.0848497 | 26.0815 | 5.01772 |
| SD | 0.231903 | 0.326325 | 0.114425 | 0.0201595 | 8.14367 | 2.87205 |
| Max | 1.59982 | 3.18508 | 0.699951 | 0.119997 | 39.9972 | 9.99883 |
| Min | 0.800219 | 1.61206 | 0.300031 | 0.0500169 | 12.0012 | 0.00182662 |

# 5,000 ARDS

Units:  Ti (sec.); Ttot (sec.); Vt (L.); C (L/cmH2O); R (cmH2O/L/sec); PEEP (cmH2O)

| Statistic/Parameter | Ti | Ttot | Vt | C | R | PEEP |
| --- | --- | --- | --- | --- | --- | --- |
| Mean | 1.10537 | 2.33331 | 0.368201 | 0.0340092 | 9.97041 | 11.2281 |
| SD | 0.19703 | 0.474589 | 0.0713678 | 0.0136919 | 1.20209 | 3.84966 |
| Max | 1.3893 | 3.28107 | 0.498916 | 0.0599774 | 11.9748 | 19.7644 |
| Min | 0.7032 | 1.34288 | 0.250325 | 0.0100416 | 8.03833 | 5.22235 |

Supplement Figure Legend

Figure E1   Elastic and Total Intracycle power (cmH2O x L /sec) as functions of time for ARDS (a,b) and severe obstruction, exemplified by COPD (c,d). Note that for each flow profile (Pressure Control, Pset, Constant Flow (CF), Decelerating Flow (DF), Accelerating Flow (AF) and Sinusoidal Flow (SF), the numerical separations between Total Power and Elastic Power widen as a function of resistance and compliance, which are both higher for COPD than for ARDS.