## Text desciption
Untargeted Lipidomic and metabolomic analyses (plasma, cells, supernatant)
Human blood was collected in 9 ml K3+ EDTA tubes, mouse blood in 500 µl K+ EDTA tubes, centrifuged at 1300 g or 1500 g and plasma was transferred into 1.5 ml tubes and stored at -80°C until further use. Cells were collected by trypsinization, washed, pelleted, resuspended in 100 µl counted and adjusted to yield 1.5 x 10exp6 cells per pellet (experiment-1) or 2.5 x 10exp5 cells (experiment-2). Finally, the cells were pelleted at high speed and supernatant removed with 100 µl and subsequently 20 µl pipet. The "dry" pellet was frozen at -80°C until lipid/metabolite extraction. Fibroblasts were homogenized as explained below (targeted analyses). Lipidomic and metabolomic analysis were conducted applying the same procedure as previously described 49. Lipidomic analysis of the human plasma samples used a slightly different procedure 50. Further protocol details for extraction and analyses of different matrices (cells, tissue, plasma) are described in the supplementary methods (Excel file).
Briefly, a methyl-tert-butyl-ether (MTBE) and methanol-based liquid-liquid extraction was used to allow for the simultaneous analysis of polar metabolites and lipids from the same sample. For analysis of cell pellets, some experiments also included the targeted analyses of sphingoid bases and ceramides from the same samples. Tissue samples were homogenized prior to extraction and 20 µl of the resulting homogenate containing 1 mg of brain tissue were used for sample preparation. For the analysis of the human plasma lipidome 20 µl of plasma were used.
For chromatographic separation of lipids, a Zorbax RRHD Eclipse Plus C8 1.8 µm 50 x 2.1 mm ID column (Agilent, Waldbronn, Germany) with a pre-column of the same type was used. The mobile phases were (A) 0.1% formic acid and 10 mM ammonium formate and (B) 0.1% formic acid in acetonitrile:isopropanol (2:3, v/v).
Analyses of lipids and polar metabolites in cells or tissue were performed on an Orbitrap Exploris 480 with a Vanquish horizon UHPLC system (both Thermo Fisher Scientific, Dreieich, Germany). Data was acquired using Thermo Scientific XCalibur v4.4 (RRID:SCR_014593) and relative quantification was performed in Thermo Scientific TraceFinder 5.1 (RRID:SCR_023045). Full scan spectra were acquired from 180-1500 m/z (lipidomics), 70-700 m/z (metabolomics positive ion mode) or 59-590 m/z (metabolomics negative ion mode) at 120,000 mass resolution each for 0.6 sec, and data dependent MS/MS spectra at 15,000 mass resolution in between. Analysis of lipids in human plasma samples was performed on a Nexera X2-UHPLC system (Shimadzu Corporation, Kyoto, Japan) coupled to a Sciex TripleTOF 6600 operated by Analyst®TF Software (RRID:SCR_015785) v1.7.1 (both Sciex, Darmstadt, Germany). The TOF-MS1 scan covered a mass range of 100-1000 m/z at 30,000 mass resolution and six data dependent MS/MS spectra in mass range of 50-1000 m/z in high sensitivity mode. Relative quantification was performed in MultiQuant 3.02 software and peak identification in MasterView v1.1 software (both Sciex, Darmstadt, Germany).
For all lipid analyses, the area under the curve (AUC) was used for quantification. AUC were transformed to square root of the AUCs (sqrt AUC) for data analysis. For multivariate analyses sqrt AUCs were scaled to have a common mean and SD of 1 (autoscaling in MetaboAnalyst).
For metabolomic analysis, polar metabolites were separated on a SeQuant ZIC-HILIC, 3.5 µm, 100 mm × 2.1 mm I.D. column coupled to a guard column with the same chemistry (both Merck, Darmstadt, Germany) and a KrudKatcher inline filter (Phenomenex, Aschaffenburg, Germany). Using 0.1 % formic acid in water (solvent A) and 0.1 % formic acid in acetonitrile (solvent B), binary gradient elution was performed with a run time of 25 min. Relative quantification of metabolites was performed based on peak areas in extracted ion chromatograms, normalized by median-based probabilistic quotient normalization (PQN).
Targeted lipid analyses
Targeted sphingolipid analyses were done in plasma, cell and tissue extracts essentially as described previously 41 using liquid chromatography-electrospray ionization-tandem mass spectrometry (LC-ESI-MS/MS), according to procedures described in detail in 51. Cell pellets were homogenized by adding ethanol:water (1:3, v/v, tissue concentration 0.02 µg/ml) using a ball mill mixer (MM400, Retsch, Haan, Germany) with 4 zirconium oxide grinding balls (25 Hz for 2.5 min). Subsequently, 50 µl of the homogenate were extracted using a liquid-liquid-extraction method. For plasma samples, we used 10 μl for lipid extraction. Sphingolipids were separated using an Agilent 1260 Infinity II UHPLC System (RRID:SCR_019365) equipped with a Zorbax C18 Eclipse Plus UHPLC column (50 Å~ 2.1 mm, 1.8 μm, Agilent, Waldbronn, Germany). For mass spectrometry, we used a hybrid triple quadrupole-ion trap mass spectrometer QTRAP 5500 (Sciex, Darmstadt, Germany), equipped with a Turbo-V-source operating in positive ESI mode. Quality control samples of three different concentration levels (low, middle, high) were run as initial and final samples of each run. For all analytes, the concentrations of the calibration standards, quality controls and samples were evaluated by Analyst®TF Software (RRID:SCR_015785) v.1.6 and MultiQuant Software 3.02 (Sciex) using the internal standard method (isotope-dilution mass spectrometry). Calibration curves were calculated by linear or quadratic regression with 1/x weighting or 1/x2 weighting. Variations in accuracy of the calibration standards were less than 15 % over the range of calibration, except for the lower limit of quantification (LLOQ), where a variation in accuracy of 20 % was accepted. Lipid concentrations in tissue are expressed as ng/ml of plasma, or pg/2.5 x 10exp5 cells. For multivariate statistical analyses, lipid data were scaled to the average or were normalized to have a common mean and standard deviation of 1 (referred to as “autoscaling” in MetaboAnalyst).

## Extraction protocols
Extraction protocols
Extraction protocol - lipidomics+metabolomics+ sphingolipids from cells - cell pellet in a 1.5 mL tube + 150 µL of lipidomics IS solved in methanol + 40 µL of sphingolipid IS + 500 µL of MTBE + 140 µL of 50 mM ammonium formate solution - Vortexing for 1 min - Centrifugation for 5 min at 18,213 x g and 4 °C - Transfer the upper layer (organic layer) into a glass vial - Re-extract the lower layer (water layer) with 200 µL MTBE:methanol:water (10:3:2.5, v/v/v, upper layer) - Vortexing for 1 min - Centrifugation for 5 min at 18,213 x g and 4 °C - Combine upper layers and split into 3 x 200 µL - evaporate under a stream of nitrogen at 45 °C (10 min) for lipidomics, sphingoid bases and ceramide analysis - dry lower phase under a stream of nitrogen at 45 °C (20-30 min) for metabolomics analysis - Store dried samples at < -70 °C until analysis - Reconstitute dried organic layer (lipidomics) with 100 µL methanol - Reconstute dried organic layer (sphingoid bases) with 100 µL methanol + 5% formic acid - Reconsitute dried organic layer (ceramides) with 100 µL tetrahydrofuran + 0.2% formic acid + 10 mM ammonium formate (9:1,v/v) - Reconstitute dried aqueous layer (metabolomics) with 100 µL of acetonitril:water (50:50, v/v) | Extraction protocol - lipidomics from plasma (QTOF) - 20 µL of plasma in a 1.5 mL tube + 150 µL of internal standards solved in methanol + 500 µL of MTBE + 125 µL of 50 mM ammonium formate solution - Vortexing for 1 min - Centrifugation for 5 min at 18,213 x g and 4 °C - Transfer the upper layer (organic layer) into a glass vial - Re-extract the lower layer (water layer) with 200 µL MTBE:methanol:water (10:3:2.5, v/v/v, upper layer) - Vortexing for 1 min - Centrifugation for 5 min at 18,213 x g and 4 °C - Combine upper layers and split into 2x290 µL (for separate measurement in positive and negative ion mode) - evaporate under a stream of nitrogen at 45 °C (10 min) - Store dried samples at < -70 °C until analysis - Reconstitute dried organic layer with 120 µL methanol
Extraction protocol - lipidomics+metabolomics tissue - homogenize the tissue with ethanol 25% to a tissue concentration of 0.05 mg/µL - 20 µL of tissue homogenate in a 1.5 mL tube + 75 µL of internal standards solved in methanol + 250 µL of MTBE + 50 µL of 50 mM ammonium formate solution - Vortexing for 1 min - Centrifugation for 5 min at 18,213 x g and 4 °C - Transfer the upper layer (organic layer) into a glass vial - Re-extract the lower layer (water layer) with 100 µL MTBE:methanol:water (10:3:2.5, v/v/v, upper layer) - Vortexing for 1 min - Centrifugation for 5 min at 18,213 x g and 4 °C - Combine upper layers and evaporate under a stream of nitrogen at 45 °C (10 min) for lipidomics analysis - add 200 µl of acetonitrile then dry the lower phase under a stream of nitrogen at 45 °C (20-30 min) for metabolomics analysis - Store dried samples at < -70 °C until analysis - Reconstitute dried organic layer (lipidomics) with 100 µL methanol - Reconstitute dried aqueous layer (metabolomics) with 100 µL of acetonitril:water (50:50, v/v)
Sphingolipids
For calibrants and QCs 20 µL of working solution are mixed with 200 µL Z-buffer and processed as described, starting at step 3
1. Aliquot 10 μL plasma in a 1.5 mL polypropylene tube
2. Add 200 μL Z-Puffer (citric acid 30 mM, disodium hydrogen phosphate 40 mM),
3. Add 20 μL internal standard working solution (in CHCl3/MeOH 2:1 (v/v) +0,1%BHT)
4. Add 600 μL CHCl3/MeOH/conc. HCl (83:15:2, v/v/v)
5. Vortex for one minute, centrifuge 5 min at approx. 20,000 *g
6. Split the organic (lower!) layer to two new polypropylene tubes (1.5 mL)
7. Dry under a nitrogen stream at 45 °C
8. Prior analysis, redissolve the samples in
A) for analysis of sphingoid bases: 50 μL Methanol/formic acid (95:5, v/v)
B) for analysis of ceramides: 50 μL THF/0.2% formic acid and 10 mM ammonium formate (9:1, v/v)
For cell supernatants: Mix 200 µL of sample with 20 µL of internal standard working solution and continue from step 4

## Materials
Materials
Reference and internal standards see the respective method details
Material | Purity | Vendor | Method | Method
Acetone | LC-MS grade | CR | Sphingoid bases + ceramides
Acetonitrile | LC-MS grade | CR | Sphingoid bases + ceramides | Lipid screening
Ammonium formate | for mass spectrometry, ≥ 99.0 % | SA | Sphingoid bases + ceramides | Lipid screening
Butylated hydroxytoluene | ≥99% | SA | Sphingoid bases + ceramides
Chloroform | ≥ 99.8 % | CR | Sphingoid bases + ceramides
Citric acid | ≥ 99.5 % | CR | Sphingoid bases + ceramides
Disodium hydrogen phosphate | x2H2O, 99.9% | VWR | Sphingoid bases + ceramides
Ethanol | for GC, ≥ 99.8 % | SA | Sphingoid bases + ceramides
Formic acid | 98-100 % | AC | Sphingoid bases + ceramides | Lipid screening
Hydrochloric acid | 37% | VWR | Sphingoid bases + ceramides
Isopropanol | LC-MS grade | CR | Sphingoid bases + ceramides | Lipid screening
Methanol | LC-MS grade | CR | Sphingoid bases + ceramides | Lipid screening
Methyl-tert butyl ether | HPLC grade | CR |  | Lipid screening
Phosphate buffered saline |  |  | Sphingoid bases + ceramides | Lipid screening
Tetrahydrofuran | ≥ 99.9 % | CR | Sphingoid bases + ceramides
Water | LC-MS grade | CR | Sphingoid bases + ceramides | Lipid screening
AC = AppliChem, Darmstadt, Germany
Bi = Biomol GmbH, Hamburg, Germany
CR = Carl Roth, Karlsruhe, Germany
FS = fisher scientific, Schwerte, Germany
M = Merck, Darmstadt, Germany
SA = Sigma Aldrich, Munich, Germany
VWR = VWR International, Radnor, PA, USA

## Lipidomics_cells_tissue
Lipidomic analysis of tissue and cells extracts
LC-MS/MS Instrumentation and Analysis
Instrumentation
HPLC: | Thermo Vanquish Horizon
MS: | Thermo Orbitrap Exploris 480
Mode | H-ESI - positive and negative electrospray ionization
Column | Zorbax RRHD Eclipse Plus C8 2.1 x 50 mm, 1.8µm (Agilent)
Pre-column | Zorbax Eclipse Plus C8 2.1 mm, 1.8 µm (Agilent)
Analysis
Data Acquisition | Xcalibur 4.4
Data Evaluation | TraceFinder 5.1
Method | Scan + ddMS2
LC parameters
Solvents
A | Water + 0,1 % formic acid + 10 mM ammonium formate
B | Acetonitrile:isopropanol (2:3) + 0.1% formic acid
Gradient Elution
Time [min] | A [%] | B [%] | Flowrate [mL/min]
0 | 75 | 25 | 0.3
0.3 | 75 | 25 | 0.3
1.5 | 20 | 80 | 0.3
11 | 0 | 100 | 0.3
12 | 0 | 100 | 0.3
12.5 | 75 | 25 | 0.3
14 | 75 | 25 | 0.3
Injection Volume |  | Column Temperature
Positive Ionization Mode | 2 µL | 40 °C
Negative Ionization Mode | 5 µL
Positive Ionization |  |  |  |  | Negative Ionization
MS parameters |  |  |  |  | MS parameters
Ion Source |  |  |  |  | Ion Source
Parameters | Value |  |  |  | Parameters | Value
Ion Source Type: | H-ESI |  |  |  | Ion Source Type: | H-ESI
Spray Voltage: | Static |  |  |  | Spray Voltage: | Static
Positive Ion (V): | 3000 |  |  |  | Positive Ion (V): | 3000
Negative Ion (V): | 2500 |  |  |  | Negative Ion (V): | 2500
Gas Mode: | Static |  |  |  | Gas Mode: | Static
Sheath Gas (Arb): | 40 |  |  |  | Sheath Gas (Arb): | 40
Aux Gas (Arb): | 20 |  |  |  | Aux Gas (Arb): | 20
Sweep Gas (Arb): | 0 |  |  |  | Sweep Gas (Arb): | 0
Ion Transfer Tube Temp (°C): | 260 |  |  |  | Ion Transfer Tube Temp (°C): | 260
Vaporizer Temp (°C): | 290 |  |  |  | Vaporizer Temp (°C): | 290
MS Global Settings |  |  |  |  | MS Global Settings
Parameters | Value |  |  |  | Parameters | Value
Infusion Mode: | Liquid Chromatography |  |  |  | Infusion Mode: | Liquid Chromatography
Expected LC Peak Width (s): | 6 |  |  |  | Expected LC Peak Width (s): | 6
Advanced Peak Determination: | TRUE |  |  |  | Advanced Peak Determination: | TRUE
Mild Trapping: | TRUE |  |  |  | Mild Trapping: | TRUE
Default Charge State: | 1 |  |  |  | Default Charge State: | 1
Enable Xcalibur AcquireX method configurations: | TRUE |  |  |  | Enable Xcalibur AcquireX method configurations: | TRUE
Internal Mass Calibration: | RunStart EASY-IC™ |  |  |  | Internal Mass Calibration: | RunStart EASY-IC™
Settings for Full Scan |  |  |  |  | Settings for Full Scan
Orbitrap Resolution: | 120000 |  |  |  | Orbitrap Resolution: | 120000
Scan Range (m/z): | 180-1500 |  |  |  | Scan Range (m/z): | 180-1500
RF Lens (%): | 25 |  |  |  | RF Lens (%): | 25
AGC Target: | Custom |  |  |  | AGC Target: | Custom
Normalized AGC Target (%): | 100 |  |  |  | Normalized AGC Target (%): | 100
Maximum Injection Time Mode: | Auto |  |  |  | Maximum Injection Time Mode: | Auto
Microscans: | 1 |  |  |  | Microscans: | 1
Data Type: | Profile |  |  |  | Data Type: | Profile
Polarity: | Positive |  |  |  | Polarity: | Negative
Source Fragmentation: | Disabled |  |  |  | Source Fragmentation: | Disabled
Filters |  |  |  |  | Filters
Intensity |  |  |  |  | Intensity
Filter Type: | Intensity Threshold |  |  |  | Filter Type: | Intensity Threshold
Intensity Threshold: | 5000 |  |  |  | Intensity Threshold: | 5000
Dynamic Exclusion |  |  |  |  | Dynamic Exclusion
Dynamic Exclusion Mode: | Custom |  |  |  | Dynamic Exclusion Mode: | Custom
Exclude after n times: | 1 |  |  |  | Exclude after n times: | 1
Exclusion duration (s): | 5 |  |  |  | Exclusion duration (s): | 5
Mass Tolerance: | ppm |  |  |  | Mass Tolerance: | ppm
Low | 10 |  |  |  | Low | 10
High | 10 |  |  |  | High | 10
Exclude isotopes: | TRUE |  |  |  | Exclude isotopes: | TRUE
Perform dependent scan on single charge state per precursor only: | FALSE |  |  |  | Perform dependent scan on single charge state per precursor only: | FALSE
Targeted Mass |  |  |  |  | Targeted Mass
Mass List |  |  |  |  | Mass List
Mass List Type: | m/z |  |  |  | Mass List Type: | m/z
Time Mode: | Start/End Time |  |  |  | Time Mode: | Start/End Time
Include Intensity Threshold: | TRUE |  |  |  | Include Intensity Threshold: | TRUE
Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE |  |  |  | Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE
Compound | m/z | t start (min) | t stop (min) | Intensity Threshold | Compound | m/z | t start (min) | t stop (min) | Intensity Threshold
MRFA (Placeholder) | 524.265 | 0 | 14 | 0 | MRFA (Placeholder) | 524.265 | 0 | 14 | 0
Mass Tolerance: | ppm |  |  |  | Mass Tolerance: | ppm
Low: | 10 |  |  |  | Low: | 10
High: | 10 |  |  |  | High: | 10
Set Collision Energy per Compound: | FALSE |  |  |  | Set Collision Energy per Compound: | FALSE
Perform dependent scan on most intense ion if no targets are found: | TRUE |  |  |  | Perform dependent scan on most intense ion if no targets are found: | TRUE
Targeted Mass Exclusion |  |  |  |  | Targeted Mass Exclusion
Mass List |  |  |  |  | Mass List
Mass List Type: | m/z |  |  |  | Mass List Type: | m/z
Time Mode: | Start/End Time |  |  |  | Time Mode: | Start/End Time
Include Intensity Threshold: | TRUE |  |  |  | Include Intensity Threshold: | TRUE
Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE |  |  |  | Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE
Compound | m/z | t start (min) | t stop (min) | Intensity Threshold | Compound | m/z | t start (min) | t stop (min) | Intensity Threshold
Mass list created via AcquireX |  |  |  |  | Mass list created via AcquireX
Mass Tolerance: | ppm |  |  |  | Mass Tolerance: | ppm
Low: | 10 |  |  |  | Low: | 10
High: | 10 |  |  |  | High: | 10
Apex Detection |  |  |  |  | Apex Detection
Desired Apex Window (%): | 30 |  |  |  | Desired Apex Window (%): | 30
Data Dependent |  |  |  |  | Data Dependent
Data Dependent Mode: | Cycle Time |  |  |  | Data Dependent Mode: | Cycle Time
Time between Master Scans (sec): | 0.6 |  |  |  | Time between Master Scans (sec): | 0.6
Settings for ddMS2 |  |  |  |  | Settings for ddMS2
Multiplex Ions: | FALSE |  |  |  | Multiplex Ions: | FALSE
Isolation Window (m/z): | 1 |  |  |  | Isolation Window (m/z): | 1
Isolation Offset: | Off |  |  |  | Isolation Offset: | Off
Collision Energy Mode: | Stepped |  |  |  | Collision Energy Mode: | Stepped
Collision Energy Type: | Normalized |  |  |  | Collision Energy Type: | Normalized
HCD Collision Energies (V): | 15,30,50 |  |  |  | HCD Collision Energies (V): | 15,30,50
Orbitrap Resolution: | 15000 |  |  |  | Orbitrap Resolution: | 15000
Scan Range Mode: | Auto |  |  |  | Scan Range Mode: | Auto
AGC Target: | Custom |  |  |  | AGC Target: | Custom
Normalized AGC Target (%): | 10 |  |  |  | Normalized AGC Target (%): | 10
Maximum Injection Time Mode: | Auto |  |  |  | Maximum Injection Time Mode: | Auto
Microscans: | 1 |  |  |  | Microscans: | 1
Data Type: | Centroid |  |  |  | Data Type: | Centroid
Setting for AcquireX
Algorithm: | Background Exclusion
Description: | Create and use an exclusion list of constant background ions and peaks to reduce background fragmentation in your ID runs.
Used reference sample: | An extraced blank sample was used to create the exclusion lists for positive and negative ionization modes.
Exclusion Override Factor: | 3
Excluison List Peak Window Extension (s): | 6
Internal Standards
Isotopically labeled internal standard | Vendor | Category | Concentration of standard working solution [ng/mL] | Volume of working solution [µL] | Sample Volume [µL] | Concentration related to sample [ng/mL]
Arachidonic acid-d8 | Cayman Chemicals | D | 100 | 75 | 10 | 750
CE 18:1-d7 | Avanti Polar Lipids | J | 5000 | 75 | 10 | 37500
Cer d18:1/16:0-d7 | Avanti Polar Lipids | A | 20 | 75 | 10 | 150
Cholesterol-d7 | Avanti Polar Lipids | K | 7506.7 | 75 | 10 | 56300.25
DG 15:0/18:1-d7 | Avanti Polar Lipids | F | 300 | 75 | 10 | 2250
LacCer d18:1/17:0 | Avanti Polar Lipids | C | 60 | 75 | 10 | 450
LPC 18:1-d7 | Avanti Polar Lipids | F | 300 | 75 | 10 | 2250
LPC O-16:0-d4 | Cayman Chemicals | A | 20 | 75 | 10 | 150
LPE 18:1-d7 | Avanti Polar Lipids | A | 20 | 75 | 10 | 150
LPG 17:1 | Avanti Polar Lipids | A | 20 | 75 | 10 | 150
LPI 17:1 | Avanti Polar Lipids | A | 20 | 75 | 10 | 150
PC 15:0/18:1-d7 | Avanti Polar Lipids | I | 2000 | 75 | 10 | 15000
PC O-18:0/18:1-d9 | Avanti Polar Lipids | E | 200 | 75 | 10 | 1500
PE 15:0/18:1-d7 | Avanti Polar Lipids | D | 100 | 75 | 10 | 750
PG 15:0/18:1-d7 | Avanti Polar Lipids | D | 100 | 75 | 10 | 750
PI 15:0/18:1-d7 | Avanti Polar Lipids | D | 100 | 75 | 10 | 750
PS 15:0/18:1-d7 | Avanti Polar Lipids | B | 25.1 | 75 | 10 | 188.25
SM d18:1/18:1-d9 | Avanti Polar Lipids | G | 400 | 75 | 10 | 3000
TG 14:0/16:1/14:0-d5 | Avanti Polar Lipids | H | 600 | 75 | 10 | 4500
TG 15:0/18:1-d7/15:0 | Avanti Polar Lipids | H | 600 | 75 | 10 | 4500
TG 20:0/20:1/20:0-d5 | Avanti Polar Lipids | H | 600 | 75 | 10 | 4500
*Avanti Polar Lipids | Alabaster, AL, USA
*Cayman Chemicals | Ann Arbor, MI, USA
Acceptance criteria
The following analytes were excluded from data evaluation:
- Analytes showing a mass error > 5 ppm
- Analytes showing an isotope score < 40 (10 ppm, 90% intensity)
- Analytes without identity confirmation via spectral library matching or reference substance measurements
- Analytes containing only missing values in QC samples and study samples
- Analytes containing > 50% missing values in QC samples and study samples
- Analytes showing > 30% area RSD in QC samples
Target List for evaluation
Compound Name | Compound Type | Compound Formula | m/z | Polarity | Adduct | Charge State | Retention Time
AHexCer 56:1;O3 | Target Compound | C62H119NO10 | 1038.8907 | Positive | M+H | 1 | 8.95
AHexCer 58:1;O3 | Target Compound | C64H123NO10 | 1066.922 | Positive | M+H | 1 | 9.34
AHexCer 58:2;O3 | Target Compound | C64H121NO10 | 1064.9063 | Positive | M+H | 1 | 8.97
AHexCer 60:2;O3 | Target Compound | C66H125NO10 | 1092.9376 | Positive | M+H | 1 | 9.38
CAR 10:0 | Target Compound | C17H33NO4 | 316.2482 | Positive | M+H | 1 | 2
CAR 10:1 | Target Compound | C17H31NO4 | 314.2326 | Positive | M+H | 1 | 1.89
CAR 12:0 | Target Compound | C19H37NO4 | 344.2795 | Positive | M+H | 1 | 2.21
CAR 12:1 | Target Compound | C19H35NO4 | 342.2639 | Positive | M+H | 1 | 2.13
CAR 13:1 | Target Compound | C20H37NO4 | 356.2795 | Positive | M+H | 1 | 2.2
CAR 14:0 | Target Compound | C21H41NO4 | 372.3108 | Positive | M+H | 1 | 2.37
CAR 14:1 | Target Compound | C21H39NO4 | 370.2952 | Positive | M+H | 1 | 2.28
CAR 16:0 | Target Compound | C23H45NO4 | 400.3421 | Positive | M+H | 1 | 2.5
CAR 16:1 | Target Compound | C23H43NO4 | 398.3265 | Positive | M+H | 1 | 2.42
CAR 18:0 | Target Compound | C25H49NO4 | 428.3734 | Positive | M+H | 1 | 2.65
CAR 18:1 | Target Compound | C25H47NO4 | 426.3578 | Positive | M+H | 1 | 2.55
CAR 18:2 | Target Compound | C25H45NO4 | 424.3421 | Positive | M+H | 1 | 2.45
CAR 20:0 | Target Compound | C27H53NO4 | 456.4047 | Positive | M+H | 1 | 2.81
CAR 20:1 | Target Compound | C27H51NO4 | 454.3891 | Positive | M+H | 1 | 2.68
CAR 20:2 | Target Compound | C27H49NO4 | 452.3734 | Positive | M+H | 1 | 2.56
CAR 20:4 | Target Compound | C27H45NO4 | 448.3421 | Positive | M+H | 1 | 2.45
CAR 22:1 | Target Compound | C29H55NO4 | 482.4204 | Positive | M+H | 1 | 2.85
CAR 22:2 | Target Compound | C29H53NO4 | 480.4047 | Positive | M+H | 1 | 2.71
CAR 8:0 | Target Compound | C15H29NO4 | 288.2169 | Positive | M+H | 1 | 1.7
CE 18:1-d7 | Internal Standard | D7C45H71O2 | 675.6779 | Positive | M+NH4 | 1 | 9.42
Cer 18:0;O2/16:0 | Target Compound | C34H69NO3 | 540.535 | Positive | M+H | 1 | 5.13
Cer 18:0;O2/22:0 | Target Compound | C40H81NO3 | 624.6289 | Positive | M+H | 1 | 6.87
Cer 18:0;O2/24:0 | Target Compound | C42H85NO3 | 652.6602 | Positive | M+H | 1 | 7.48
Cer 18:0;O2/24:1 | Target Compound | C42H83NO3 | 650.6446 | Positive | M+H | 1 | 6.89
Cer 18:1;O2/16:0 | Target Compound | C34H67NO3 | 538.5194 | Positive | M+H | 1 | 4.92
Cer 18:1;O2/16:0-d7 | Internal Standard | D7C34H60NO3 | 545.5633 | Positive | M+H | 1 | 4.92
Cer 18:1;O2/18:0 | Target Compound | C36H71NO3 | 566.5507 | Positive | M+H | 1 | 5.48
Cer 18:1;O2/20:0 | Target Compound | C38H75NO3 | 594.582 | Positive | M+H | 1 | 6.16
Cer 18:1;O2/21:0 | Target Compound | C39H77NO3 | 608.5976 | Positive | M+H | 1 | 6.45
Cer 18:1;O2/22:0 | Target Compound | C40H79NO3 | 622.6133 | Positive | M+H | 1 | 6.67
Cer 18:1;O2/23:0 | Target Compound | C41H81NO3 | 636.6289 | Positive | M+H | 1 | 6.98
Cer 18:1;O2/24:0 | Target Compound | C42H83NO3 | 650.6446 | Positive | M+H | 1 | 7.27
Cer 18:1;O2/24:1 | Target Compound | C42H81NO3 | 648.6289 | Positive | M+H | 1 | 6.69
Cer 18:1;O2/26:0 | Target Compound | C44H87NO3 | 678.6759 | Positive | M+H | 1 | 7.82
Cer 18:1;O2/36:10 | Target Compound | C54H87NO3 | 798.6759 | Positive | M+H | 1 | 6.32
Cer 18:1;O2/38:10 | Target Compound | C56H91NO3 | 826.7072 | Positive | M+H | 1 | 6.92
Cer 18:2;O2/16:0 | Target Compound | C34H65NO3 | 536.5037 | Positive | M+H | 1 | 4.37
Cer 18:2;O2/22:0 | Target Compound | C40H77NO3 | 620.5976 | Positive | M+H | 1 | 6.2
Cer 18:2;O2/23:0 | Target Compound | C41H79NO3 | 634.6133 | Positive | M+H | 1 | 6.51
Cer 18:2;O2/24:1 | Target Compound | C42H79NO3 | 646.6133 | Positive | M+H | 1 | 6.21
Cholesterol-d7 | Internal Standard | D7C27H39O | 376.3955 | Positive | M-OH- | 1 | 3.93
DG 15:0/18:1-d7 | Internal Standard | D7C36H61O5 | 605.5844 | Positive | M+NH4 | 1 | 5.59
DG 30:1 | Target Compound | C33H62O5 | 556.4936 | Positive | M+NH4 | 1 | 4.8
DG 31:0 | Target Compound | C34H66O5 | 572.5249 | Positive | M+NH4 | 1 | 5.5
DG 31:1 | Target Compound | C34H64O5 | 570.5092 | Positive | M+NH4 | 1 | 5.05
DG 32:0 | Target Compound | C35H68O5 | 586.5405 | Positive | M+NH4 | 1 | 5.8
DG 32:1 | Target Compound | C35H66O5 | 584.5249 | Positive | M+NH4 | 1 | 5.35
DG 32:2 | Target Compound | C35H64O5 | 582.5092 | Positive | M+NH4 | 1 | 4.94
DG 33:0 | Target Compound | C36H70O5 | 600.5562 | Positive | M+NH4 | 1 | 6.05
DG 33:1 | Target Compound | C36H68O5 | 598.5405 | Positive | M+NH4 | 1 | 5.6
DG 34:0 | Target Compound | C37H72O5 | 614.5718 | Positive | M+NH4 | 1 | 6.35
DG 34:1 | Target Compound | C37H70O5 | 612.5562 | Positive | M+NH4 | 1 | 5.9
DG 34:2 | Target Compound | C37H68O5 | 610.5405 | Positive | M+NH4 | 1 | 5.48
DG 34:3 | Target Compound | C37H66O5 | 608.5249 | Positive | M+NH4 | 1 | 5.05
DG 35:1 | Target Compound | C38H72O5 | 626.5718 | Positive | M+NH4 | 1 | 6.2
DG 35:2 | Target Compound | C38H70O5 | 624.5562 | Positive | M+NH4 | 1 | 5.76
DG 36:0 | Target Compound | C39H76O5 | 642.6031 | Positive | M+NH4 | 1 | 6.95
DG 36:1 | Target Compound | C39H74O5 | 640.5874 | Positive | M+NH4 | 1 | 6.5
DG 36:2 | Target Compound | C39H72O5 | 638.5718 | Positive | M+NH4 | 1 | 6.05
DG 36:3 | Target Compound | C39H70O5 | 636.5562 | Positive | M+NH4 | 1 | 5.6
DG 36:4 | Target Compound | C39H68O5 | 634.5405 | Positive | M+NH4 | 1 | 5.09
DG 36:5 | Target Compound | C39H66O5 | 632.5249 | Positive | M+NH4 | 1 | 4.7
DG 38:1 | Target Compound | C41H78O5 | 668.6187 | Positive | M+NH4 | 1 | 6.8
DG 38:2 | Target Compound | C41H76O5 | 666.6031 | Positive | M+NH4 | 1 | 6.33
DG 38:3 | Target Compound | C41H74O5 | 664.5874 | Positive | M+NH4 | 1 | 5.9
DG 38:4 | Target Compound | C41H72O5 | 662.5718 | Positive | M+NH4 | 1 | 5.9
DG 38:5 | Target Compound | C41H70O5 | 660.5562 | Positive | M+NH4 | 1 | 5.38
DG 38:6 | Target Compound | C41H68O5 | 658.5405 | Positive | M+NH4 | 1 | 4.97
DG 40:6 | Target Compound | C43H72O5 | 686.5718 | Positive | M+NH4 | 1 | 5.4
DG 40:7 | Target Compound | C43H70O5 | 684.5562 | Positive | M+NH4 | 1 | 5.19
DG 40:8 | Target Compound | C43H68O5 | 682.5405 | Positive | M+NH4 | 1 | 4.76
DG 44:8 | Target Compound | C47H76O5 | 738.6031 | Positive | M+NH4 | 1 | 5.85
Hex2Cer 18:1;O2/16:0 | Target Compound | C46H87NO13 | 862.625 | Positive | M+H | 1 | 4.24
Hex2Cer 18:1;O2/24:0 | Target Compound | C54H103NO13 | 974.7502 | Positive | M+H | 1 | 6.37
Hex2Cer 18:1;O2/24:1 | Target Compound | C54H101NO13 | 972.7346 | Positive | M+H | 1 | 5.76
HexCer 18:0;O2/18:0 | Target Compound | C42H83NO8 | 730.6191 | Positive | M+H | 1 | 5.15
HexCer 18:0;O2/22:0 | Target Compound | C46H91NO8 | 786.6817 | Positive | M+H | 1 | 6.28
HexCer 18:0;O2/22:0;O | Target Compound | C46H91NO9 | 802.6767 | Positive | M+H | 1 | 6.22
HexCer 18:0;O2/23:0 | Target Compound | C47H93NO8 | 800.6974 | Positive | M+H | 1 | 6.58
HexCer 18:0;O2/24:0 | Target Compound | C48H95NO8 | 814.713 | Positive | M+H | 1 | 6.87
HexCer 18:0;O2/24:0;O | Target Compound | C48H95NO9 | 830.708 | Positive | M+H | 1 | 6.84
HexCer 18:1;O2/18:0 | Target Compound | C42H81NO8 | 728.6035 | Positive | M+H | 1 | 4.94
HexCer 18:1;O2/18:0;O | Target Compound | C42H81NO9 | 744.5984 | Positive | M+H | 1 | 4.88
HexCer 18:1;O2/20:0 | Target Compound | C44H85NO8 | 756.6348 | Positive | M+H | 1 | 5.52
HexCer 18:1;O2/20:0;O | Target Compound | C44H85NO9 | 772.6297 | Positive | M+H | 1 | 5.42
HexCer 18:1;O2/21:0;O | Target Compound | C45H87NO9 | 786.6454 | Positive | M+H | 1 | 5.7
HexCer 18:1;O2/22:0 | Target Compound | C46H89NO8 | 784.6661 | Positive | M+H | 1 | 6.07
HexCer 18:1;O2/22:0;O | Target Compound | C46H89NO9 | 800.661 | Positive | M+H | 1 | 5.95
HexCer 18:1;O2/22:1 | Target Compound | C46H87NO8 | 782.6504 | Positive | M+H | 1 | 5.51
HexCer 18:1;O2/22:1;O | Target Compound | C46H87NO9 | 798.6454 | Positive | M+H | 1 | 5.46
HexCer 18:1;O2/23:0;O | Target Compound | C47H91NO9 | 814.6767 | Positive | M+H | 1 | 6.3
HexCer 18:1;O2/23:1 | Target Compound | C47H89NO8 | 796.6661 | Positive | M+H | 1 | 5.79
HexCer 18:1;O2/23:1;O | Target Compound | C47H89NO9 | 812.661 | Positive | M+H | 1 | 5.8
HexCer 18:1;O2/24:0 | Target Compound | C48H93NO8 | 812.6974 | Positive | M+H | 1 | 6.71
HexCer 18:1;O2/24:0;O | Target Compound | C48H93NO9 | 828.6923 | Positive | M+H | 1 | 6.6
HexCer 18:1;O2/24:1 | Target Compound | C48H91NO8 | 810.6817 | Positive | M+H | 1 | 6.06
HexCer 18:1;O2/24:1;O | Target Compound | C48H91NO9 | 826.6767 | Positive | M+H | 1 | 6
HexCer 18:1;O2/24:2 | Target Compound | C48H89NO8 | 808.6661 | Positive | M+H | 1 | 5.6
HexCer 18:1;O2/24:2;O | Target Compound | C48H89NO9 | 824.661 | Positive | M+H | 1 | 5.52
HexCer 18:1;O2/25:0;O | Target Compound | C49H95NO9 | 842.708 | Positive | M+H | 1 | 6.84
HexCer 18:1;O2/25:1 | Target Compound | C49H93NO8 | 824.6974 | Positive | M+H | 1 | 6.29
HexCer 18:1;O2/25:1;O | Target Compound | C49H93NO9 | 840.6923 | Positive | M+H | 1 | 6.28
HexCer 18:1;O2/40:2;O | Target Compound | C64H121NO9 | 1048.9114 | Positive | M+H | 1 | 9.07
HexCer 18:1;O2/42:2;O | Target Compound | C66H125NO9 | 1076.9427 | Positive | M+H | 1 | 9.44
HexCer 18:2;O2/24:0 | Target Compound | C48H91NO8 | 810.6817 | Positive | M+H | 1 | 6.16
LacCer 18:1;O2/17:0 | Internal Standard | C47H89NO13 | 876.6407 | Positive | M+H | 1 | 4.46
LPC 14:0 | Target Compound | C22H46NO7P | 468.3085 | Positive | M+H | 1 | 2.45
LPC 15:0 | Target Compound | C23H48NO7P | 482.3241 | Positive | M+H | 1 | 2.53
LPC 16:0 | Target Compound | C24H50NO7P | 496.3398 | Positive | M+H | 1 | 2.61
LPC 16:1 | Target Compound | C24H48NO7P | 494.3241 | Positive | M+H | 1 | 2.49
LPC 17:0 | Target Compound | C25H52NO7P | 510.3554 | Positive | M+H | 1 | 2.68
LPC 18:0 | Target Compound | C26H54NO7P | 524.3711 | Positive | M+H | 1 | 2.78
LPC 18:1 | Target Compound | C26H52NO7P | 522.3554 | Positive | M+H | 1 | 2.65
LPC 18:1-d7 | Internal Standard | D7C26H45NO7P | 529.3994 | Positive | M+H | 1 | 2.65
LPC 18:2 | Target Compound | C26H50NO7P | 520.3398 | Positive | M+H | 1 | 2.53
LPC 18:3 | Target Compound | C26H48NO7P | 518.3241 | Positive | M+H | 1 | 2.45
LPC 19:0 | Target Compound | C27H56NO7P | 538.3867 | Positive | M+H | 1 | 2.88
LPC 19:1 | Target Compound | C27H54NO7P | 536.3711 | Positive | M+H | 1 | 2.73
LPC 20:0 | Target Compound | C28H58NO7P | 552.4024 | Positive | M+H | 1 | 3
LPC 20:1 | Target Compound | C28H56NO7P | 550.3867 | Positive | M+H | 1 | 2.81
LPC 20:2 | Target Compound | C28H54NO7P | 548.3711 | Positive | M+H | 1 | 2.68
LPC 20:3 | Target Compound | C28H52NO7P | 546.3554 | Positive | M+H | 1 | 2.78
LPC 20:4 | Target Compound | C28H50NO7P | 544.3398 | Positive | M+H | 1 | 2.52
LPC 20:5 | Target Compound | C28H48NO7P | 542.3241 | Positive | M+H | 1 | 2.4
LPC 22:0 | Target Compound | C30H62NO7P | 580.4337 | Positive | M+H | 1 | 3.3
LPC 22:4 | Target Compound | C30H54NO7P | 572.3711 | Positive | M+H | 1 | 2.64
LPC 22:5 | Target Compound | C30H52NO7P | 570.3554 | Positive | M+H | 1 | 2.53
LPC 22:6 | Target Compound | C30H50NO7P | 568.3398 | Positive | M+H | 1 | 2.49
LPC 24:0 | Target Compound | C32H66NO7P | 608.465 | Positive | M+H | 1 | 3.65
LPC 24:1 | Target Compound | C32H64NO7P | 606.4493 | Positive | M+H | 1 | 3.3
LPC O-16:0-d4 | Internal Standard | D4C24H48NO6P | 486.3856 | Positive | M+H | 1 | 2.69
LPC O-24:1 | Target Compound | C32H66NO6P | 592.4701 | Positive | M+H | 1 | 3.48
LPE 18:1-d7 | Internal Standard | D7C23H39NO7P | 487.3524 | Positive | M+H | 1 | 2.67
LPG 17:1 | Internal Standard | C23H45O9P | 514.3139 | Positive | M+NH4 | 1 | 2.48
LPI 17:1 | Internal Standard | C26H49O12P | 602.33 | Positive | M+NH4 | 1 | 2.45
PC 15:0/18:1-d7 | Internal Standard | D7C41H73NO8P | 753.6134 | Positive | M+H | 1 | 4.47
PC 25:0 | Target Compound | C33H66NO8P | 636.4599 | Positive | M+H | 1 | 3.2
PC 28:0 | Target Compound | C36H72NO8P | 678.5068 | Positive | M+H | 1 | 3.8
PC 28:1 | Target Compound | C36H70NO8P | 676.4912 | Positive | M+H | 1 | 3.52
PC 29:0 | Target Compound | C37H74NO8P | 692.5225 | Positive | M+H | 1 | 3.98
PC 30:0 | Target Compound | C38H76NO8P | 706.5381 | Positive | M+H | 1 | 4.17
PC 30:1 | Target Compound | C38H74NO8P | 704.5225 | Positive | M+H | 1 | 3.87
PC 30:2 | Target Compound | C38H72NO8P | 702.5068 | Positive | M+H | 1 | 3.6
PC 31:0 | Target Compound | C39H78NO8P | 720.5538 | Positive | M+H | 1 | 4.39
PC 31:1 | Target Compound | C39H76NO8P | 718.5381 | Positive | M+H | 1 | 4.05
PC 31:2 | Target Compound | C39H74NO8P | 716.5225 | Positive | M+H | 1 | 3.72
PC 32:0 | Target Compound | C40H80NO8P | 734.5694 | Positive | M+H | 1 | 4.59
PC 32:1 | Target Compound | C40H78NO8P | 732.5538 | Positive | M+H | 1 | 4.26
PC 32:2 | Target Compound | C40H76NO8P | 730.5381 | Positive | M+H | 1 | 3.93
PC 33:0 | Target Compound | C41H82NO8P | 748.5851 | Positive | M+H | 1 | 4.8
PC 33:1 | Target Compound | C41H80NO8P | 746.5694 | Positive | M+H | 1 | 4.47
PC 33:2 | Target Compound | C41H78NO8P | 744.5538 | Positive | M+H | 1 | 4.14
PC 33:3 | Target Compound | C41H76NO8P | 742.5381 | Positive | M+H | 1 | 3.77
PC 34:0 | Target Compound | C42H84NO8P | 762.6007 | Positive | M+H | 1 | 5.1
PC 34:0;O | Target Compound | C42H82NO9P | 776.58 | Positive | M+H | 1 | 4.2
PC 34:1 | Target Compound | C42H82NO8P | 760.5851 | Positive | M+H | 1 | 4.69
PC 34:2 | Target Compound | C42H80NO8P | 758.5694 | Positive | M+H | 1 | 4.33
PC 34:3 | Target Compound | C42H78NO8P | 756.5538 | Positive | M+H | 1 | 4.03
PC 34:4 | Target Compound | C42H76NO8P | 754.5381 | Positive | M+H | 1 | 3.9
PC 34:5 | Target Compound | C42H74NO8P | 752.5225 | Positive | M+H | 1 | 3.62
PC 35:0 | Target Compound | C43H86NO8P | 776.6164 | Positive | M+H | 1 | 5.3
PC 35:1 | Target Compound | C43H84NO8P | 774.6007 | Positive | M+H | 1 | 4.94
PC 35:2 | Target Compound | C43H82NO8P | 772.5851 | Positive | M+H | 1 | 4.56
PC 35:3 | Target Compound | C43H80NO8P | 770.5694 | Positive | M+H | 1 | 4.23
PC 35:4 | Target Compound | C43H78NO8P | 768.5538 | Positive | M+H | 1 | 4.08
PC 36:0 | Target Compound | C44H88NO8P | 790.632 | Positive | M+H | 1 | 5.62
PC 36:1 | Target Compound | C44H86NO8P | 788.6164 | Positive | M+H | 1 | 5.21
PC 36:2 | Target Compound | C44H84NO8P | 786.6007 | Positive | M+H | 1 | 4.81
PC 36:3 | Target Compound | C44H82NO8P | 784.5851 | Positive | M+H | 1 | 4.46
PC 36:4 | Target Compound | C44H80NO8P | 782.5694 | Positive | M+H | 1 | 4.28
PC 36:4;O | Target Compound | C44H80NO9P | 798.5643 | Positive | M+H | 1 | 3.84
PC 36:5 | Target Compound | C44H78NO8P | 780.5538 | Positive | M+H | 1 | 3.98
PC 36:6 | Target Compound | C44H76NO8P | 778.5381 | Positive | M+H | 1 | 3.74
PC 37:1 | Target Compound | C45H88NO8P | 802.632 | Positive | M+H | 1 | 5.48
PC 37:2 | Target Compound | C45H86NO8P | 800.6164 | Positive | M+H | 1 | 5.06
PC 37:3 | Target Compound | C45H84NO8P | 798.6007 | Positive | M+H | 1 | 4.66
PC 37:4 | Target Compound | C45H82NO8P | 796.5851 | Positive | M+H | 1 | 4.51
PC 37:6 | Target Compound | C45H78NO8P | 792.5538 | Positive | M+H | 1 | 3.93
PC 38:0 | Target Compound | C46H92NO8P | 818.6633 | Positive | M+H | 1 | 6.21
PC 38:1 | Target Compound | C46H90NO8P | 816.6477 | Positive | M+H | 1 | 5.68
PC 38:2 | Target Compound | C46H88NO8P | 814.632 | Positive | M+H | 1 | 5.24
PC 38:3 | Target Compound | C46H86NO8P | 812.6164 | Positive | M+H | 1 | 4.93
PC 38:4 | Target Compound | C46H84NO8P | 810.6007 | Positive | M+H | 1 | 4.75
PC 38:5 | Target Compound | C46H82NO8P | 808.5851 | Positive | M+H | 1 | 4.39
PC 38:6 | Target Compound | C46H80NO8P | 806.5694 | Positive | M+H | 1 | 4.14
PC 38:7 | Target Compound | C46H78NO8P | 804.5538 | Positive | M+H | 1 | 3.85
PC 39:1 | Target Compound | C47H92NO8P | 830.6633 | Positive | M+H | 1 | 6.03
PC 39:2 | Target Compound | C47H90NO8P | 828.6477 | Positive | M+H | 1 | 5.6
PC 39:3 | Target Compound | C47H88NO8P | 826.632 | Positive | M+H | 1 | 5.19
PC 39:4 | Target Compound | C47H86NO8P | 824.6164 | Positive | M+H | 1 | 5.01
PC 39:5 | Target Compound | C47H84NO8P | 822.6007 | Positive | M+H | 1 | 4.55
PC 39:6 | Target Compound | C47H82NO8P | 820.5851 | Positive | M+H | 1 | 4.31
PC 39:7 | Target Compound | C47H80NO8P | 818.5694 | Positive | M+H | 1 | 4.02
PC 40:0 | Target Compound | C48H96NO8P | 846.6946 | Positive | M+H | 1 | 6.81
PC 40:1 | Target Compound | C48H94NO8P | 844.679 | Positive | M+H | 1 | 6.34
PC 40:2 | Target Compound | C48H92NO8P | 842.6633 | Positive | M+H | 1 | 5.9
PC 40:4 | Target Compound | C48H88NO8P | 838.632 | Positive | M+H | 1 | 5.25
PC 40:6 | Target Compound | C48H84NO8P | 834.6007 | Positive | M+H | 1 | 4.56
PC 40:7 | Target Compound | C48H82NO8P | 832.5851 | Positive | M+H | 1 | 4.2
PC 40:8 | Target Compound | C48H80NO8P | 830.5694 | Positive | M+H | 1 | 3.93
PC 41:1 | Target Compound | C49H96NO8P | 858.6946 | Positive | M+H | 1 | 6.6
PC 41:2 | Target Compound | C49H94NO8P | 856.679 | Positive | M+H | 1 | 6.18
PC 41:6 | Target Compound | C49H86NO8P | 848.6164 | Positive | M+H | 1 | 4.76
PC 42:1 | Target Compound | C50H98NO8P | 872.7103 | Positive | M+H | 1 | 6.94
PC 42:10 | Target Compound | C50H80NO8P | 854.5694 | Positive | M+H | 1 | 3.81
PC 42:2 | Target Compound | C50H96NO8P | 870.6946 | Positive | M+H | 1 | 6.47
PC 42:4 | Target Compound | C50H92NO8P | 866.6633 | Positive | M+H | 1 | 5.6
PC 42:5 | Target Compound | C50H90NO8P | 864.6477 | Positive | M+H | 1 | 5.1
PC 42:6 | Target Compound | C50H88NO8P | 862.632 | Positive | M+H | 1 | 4.86
PC 42:7 | Target Compound | C50H86NO8P | 860.6164 | Positive | M+H | 1 | 4.62
PC 42:8 | Target Compound | C50H84NO8P | 858.6007 | Positive | M+H | 1 | 4.23
PC 42:9 | Target Compound | C50H82NO8P | 856.5851 | Positive | M+H | 1 | 4
PC 43:1 | Target Compound | C51H100NO8P | 886.7259 | Positive | M+H | 1 | 7.19
PC 43:2 | Target Compound | C51H98NO8P | 884.7103 | Positive | M+H | 1 | 6.78
PC 44:1 | Target Compound | C52H102NO8P | 900.7416 | Positive | M+H | 1 | 7.47
PC 44:10 | Target Compound | C52H84NO8P | 882.6007 | Positive | M+H | 1 | 4.23
PC 44:11 | Target Compound | C52H82NO8P | 880.5851 | Positive | M+H | 1 | 3.91
PC 44:12 | Target Compound | C52H80NO8P | 878.5694 | Positive | M+H | 1 | 3.6
PC 44:2 | Target Compound | C52H100NO8P | 898.7259 | Positive | M+H | 1 | 7.05
PC 44:4 | Target Compound | C52H96NO8P | 894.6946 | Positive | M+H | 1 | 6.35
PC 44:5 | Target Compound | C52H94NO8P | 892.679 | Positive | M+H | 1 | 5.81
PC 44:8 | Target Compound | C52H88NO8P | 886.632 | Positive | M+H | 1 | 4.58
PC O-16:0_16:0 | Target Compound | C40H82NO7P | 720.5902 | Positive | M+H | 1 | 4.94
PC O-16:0_18:2 | Target Compound | C42H82NO7P | 744.5902 | Positive | M+H | 1 | 5.02
PC O-16:0_20:4 | Target Compound | C44H82NO7P | 768.5902 | Positive | M+H | 1 | 4.59
PC O-16:1_16:0 | Target Compound | C40H80NO7P | 718.5745 | Positive | M+H | 1 | 4.89
PC O-16:1_18:1 | Target Compound | C42H82NO7P | 744.5902 | Positive | M+H | 1 | 4.66
PC O-16:1_18:2 | Target Compound | C42H80NO7P | 742.5745 | Positive | M+H | 1 | 4.6
PC O-16:1_20:4 | Target Compound | C44H80NO7P | 766.5745 | Positive | M+H | 1 | 4.53
PC O-18:0/18:1-d9 | Internal Standard | D9C44H77NO7P | 781.678 | Positive | M+H | 1 | 5.51
PC O-18:0_20:4 | Target Compound | C46H86NO7P | 796.6215 | Positive | M+H | 1 | 5.1
PC O-18:1_18:1 | Target Compound | C44H86NO7P | 772.6215 | Positive | M+H | 1 | 5.16
PC O-18:1_18:2 | Target Compound | C44H84NO7P | 770.6058 | Positive | M+H | 1 | 4.75
PC O-18:1_20:4 | Target Compound | C46H84NO7P | 794.6058 | Positive | M+H | 1 | 4.65
PC O-30:0 | Target Compound | C38H78NO7P | 692.5589 | Positive | M+H | 1 | 5.84
PC O-30:1 | Target Compound | C38H76NO7P | 690.5432 | Positive | M+H | 1 | 5.34
PC O-31:0 | Target Compound | C39H80NO7P | 706.5745 | Positive | M+H | 1 | 5.26
PC O-31:1 | Target Compound | C39H78NO7P | 704.5589 | Positive | M+H | 1 | 6.47
PC O-32:2 | Target Compound | C40H78NO7P | 716.5589 | Positive | M+H | 1 | 4.6
PC O-33:0 | Target Compound | C41H84NO7P | 734.6058 | Positive | M+H | 1 | 4.35
PC O-33:1 | Target Compound | C41H82NO7P | 732.5902 | Positive | M+H | 1 | 4.59
PC O-34:0 | Target Compound | C42H86NO7P | 748.6215 | Positive | M+H | 1 | 4.86
PC O-34:1 | Target Compound | C42H84NO7P | 746.6058 | Positive | M+H | 1 | 5.05
PC O-34:4 | Target Compound | C42H78NO7P | 740.5589 | Positive | M+H | 1 | 4.34
PC O-36:0 | Target Compound | C44H90NO7P | 776.6528 | Positive | M+H | 1 | 4.9
PC O-36:1 | Target Compound | C44H88NO7P | 774.6371 | Positive | M+H | 1 | 5.51
PC O-36:6 | Target Compound | C44H78NO7P | 764.5589 | Positive | M+H | 1 | 4.4
PC O-37:4 | Target Compound | C45H84NO7P | 782.6058 | Positive | M+H | 1 | 4.86
PC O-37:6 | Target Compound | C45H80NO7P | 778.5745 | Positive | M+H | 1 | 4.5
PC O-38:1 | Target Compound | C46H92NO7P | 802.6684 | Positive | M+H | 1 | 4.9
PC O-38:10 | Target Compound | C46H74NO7P | 784.5276 | Positive | M+H | 1 | 4.15
PC O-38:2_1 | Target Compound | C46H90NO7P | 800.6528 | Positive | M+H | 1 | 6.29
PC O-38:2_2 | Target Compound | C46H90NO7P | 800.6528 | Positive | M+H | 1 | 4.6
PC O-38:3 | Target Compound | C46H88NO7P | 798.6371 | Positive | M+H | 1 | 4.9
PC O-38:6 | Target Compound | C46H82NO7P | 792.5902 | Positive | M+H | 1 | 4.55
PC O-38:7 | Target Compound | C46H80NO7P | 790.5745 | Positive | M+H | 1 | 4.45
PC O-38:9 | Target Compound | C46H76NO7P | 786.5432 | Positive | M+H | 1 | 4.25
PC O-39:1 | Target Compound | C47H94NO7P | 816.6841 | Positive | M+H | 1 | 5.05
PC O-39:6 | Target Compound | C47H84NO7P | 806.6058 | Positive | M+H | 1 | 4.6
PC O-39:8 | Target Compound | C47H80NO7P | 802.5745 | Positive | M+H | 1 | 4.55
PC O-40:10 | Target Compound | C48H78NO7P | 812.5589 | Positive | M+H | 1 | 4.25
PC O-40:11 | Target Compound | C48H76NO7P | 810.5432 | Positive | M+H | 1 | 4.15
PC O-40:2 | Target Compound | C48H94NO7P | 828.6841 | Positive | M+H | 1 | 5.05
PC O-40:4 | Target Compound | C48H90NO7P | 824.6528 | Positive | M+H | 1 | 5.36
PC O-40:5 | Target Compound | C48H88NO7P | 822.6371 | Positive | M+H | 1 | 4.77
PC O-40:6 | Target Compound | C48H86NO7P | 820.6215 | Positive | M+H | 1 | 4.65
PC O-40:7 | Target Compound | C48H84NO7P | 818.6058 | Positive | M+H | 1 | 4.55
PC O-40:8 | Target Compound | C48H82NO7P | 816.5902 | Positive | M+H | 1 | 4.45
PC O-40:9 | Target Compound | C48H80NO7P | 814.5745 | Positive | M+H | 1 | 4.35
PC O-42:10 | Target Compound | C50H82NO7P | 840.5902 | Positive | M+H | 1 | 4.5
PC O-42:3 | Target Compound | C50H96NO7P | 854.6997 | Positive | M+H | 1 | 5.2
PC O-42:4 | Target Compound | C50H94NO7P | 852.6841 | Positive | M+H | 1 | 5.62
PC O-42:5 | Target Compound | C50H92NO7P | 850.6684 | Positive | M+H | 1 | 4.89
PC O-42:6 | Target Compound | C50H90NO7P | 848.6528 | Positive | M+H | 1 | 4.75
PC O-42:7 | Target Compound | C50H88NO7P | 846.6371 | Positive | M+H | 1 | 4.65
PC O-44:5 | Target Compound | C52H96NO7P | 878.6997 | Positive | M+H | 1 | 5.01
PC O-44:6 | Target Compound | C52H94NO7P | 876.6841 | Positive | M+H | 1 | 4.8
PC O-44:7 | Target Compound | C52H92NO7P | 874.6684 | Positive | M+H | 1 | 4.8
PC O-46:7 | Target Compound | C54H96NO7P | 902.6997 | Positive | M+H | 1 | 4.9
PE 15:0/18:1-d7 | Internal Standard | D7C38H67NO8P | 711.5664 | Positive | M+H | 1 | 4.62
PG 15:0/18:1-d7 | Internal Standard | D7C39H68O10P | 759.5875 | Positive | M+NH4 | 1 | 4.03
PI 15:0/18:1-d7 | Internal Standard | D7C42H72O13P | 847.6036 | Positive | M+NH4 | 1 | 3.91
PS 15:0/18:1-d7 | Internal Standard | D7C39H67NO10P | 755.5562 | Positive | M+H | 1 | 4.07
SE 27:1/14:0 | Target Compound | C41H72O2 | 614.5871 | Positive | M+NH4 | 1 | 9.05
SE 27:1/15:0 | Target Compound | C42H74O2 | 628.6027 | Positive | M+NH4 | 1 | 9.25
SE 27:1/16:2 | Target Compound | C43H72O2 | 638.5871 | Positive | M+NH4 | 1 | 8.62
SE 27:1/17:0 | Target Compound | C44H78O2 | 656.634 | Positive | M+NH4 | 1 | 9.65
SE 27:1/17:1 | Target Compound | C44H76O2 | 654.6184 | Positive | M+NH4 | 1 | 9.3
SE 27:1/18:1 | Target Compound | C45H78O2 | 668.634 | Positive | M+NH4 | 1 | 9.44
SE 27:1/18:2 | Target Compound | C45H76O2 | 666.6184 | Positive | M+NH4 | 1 | 9.04
SE 27:1/18:3 | Target Compound | C45H74O2 | 664.6027 | Positive | M+NH4 | 1 | 8.61
SE 27:1/20:1 | Target Compound | C47H82O2 | 696.6653 | Positive | M+NH4 | 1 | 9.8
SE 27:1/20:2 | Target Compound | C47H80O2 | 694.6497 | Positive | M+NH4 | 1 | 9.46
SE 27:1/20:3 | Target Compound | C47H78O2 | 692.634 | Positive | M+NH4 | 1 | 9.08
SE 27:1/20:4 | Target Compound | C47H76O2 | 690.6184 | Positive | M+NH4 | 1 | 8.76
SE 27:1/20:5 | Target Compound | C47H74O2 | 688.6027 | Positive | M+NH4 | 1 | 8.34
SE 27:1/22:4 | Target Compound | C49H80O2 | 718.6497 | Positive | M+NH4 | 1 | 9.14
SE 27:1/22:5 | Target Compound | C49H78O2 | 716.634 | Positive | M+NH4 | 1 | 8.77
SE 27:1/22:6 | Target Compound | C49H76O2 | 714.6184 | Positive | M+NH4 | 1 | 8.46
SE 27:1/24:6 | Target Compound | C51H80O2 | 742.6497 | Positive | M+NH4 | 1 | 8.9
SM 18:1;O2/18:1-d9 | Internal Standard | D9C41H72N2O6P | 738.647 | Positive | M+H | 1 | 4.36
SM 30:0;O2 | Target Compound | C35H73N2O6P | 649.5279 | Positive | M+H | 1 | 3.67
SM 30:1;O2 | Target Compound | C35H71N2O6P | 647.5122 | Positive | M+H | 1 | 3.52
SM 32:0;O2 | Target Compound | C37H77N2O6P | 677.5592 | Positive | M+H | 1 | 4.03
SM 32:1;O2 | Target Compound | C37H75N2O6P | 675.5435 | Positive | M+H | 1 | 3.86
SM 32:2;O2 | Target Compound | C37H73N2O6P | 673.5279 | Positive | M+H | 1 | 3.57
SM 33:1;O2 | Target Compound | C38H77N2O6P | 689.5592 | Positive | M+H | 1 | 4.06
SM 33:2;O2 | Target Compound | C38H75N2O6P | 687.5435 | Positive | M+H | 1 | 3.73
SM 34:0;O2 | Target Compound | C39H81N2O6P | 705.5905 | Positive | M+H | 1 | 4.45
SM 34:1;O2 | Target Compound | C39H79N2O6P | 703.5748 | Positive | M+H | 1 | 4.25
SM 34:1;O3 | Target Compound | C39H79N2O7P | 719.5698 | Positive | M+H | 1 | 4.08
SM 34:2;O2 | Target Compound | C39H77N2O6P | 701.5592 | Positive | M+H | 1 | 3.93
SM 35:0;O2 | Target Compound | C40H83N2O6P | 719.6061 | Positive | M+H | 1 | 4.66
SM 35:1;O2 | Target Compound | C40H81N2O6P | 717.5905 | Positive | M+H | 1 | 4.48
SM 35:2;O2 | Target Compound | C40H79N2O6P | 715.5748 | Positive | M+H | 1 | 4.11
SM 36:0;O2 | Target Compound | C41H85N2O6P | 733.6218 | Positive | M+H | 1 | 4.92
SM 36:1;O2 | Target Compound | C41H83N2O6P | 731.6061 | Positive | M+H | 1 | 4.73
SM 36:2;O2 | Target Compound | C41H81N2O6P | 729.5905 | Positive | M+H | 1 | 4.36
SM 36:3;O2 | Target Compound | C41H79N2O6P | 727.5748 | Positive | M+H | 1 | 4.03
SM 37:1;O2 | Target Compound | C42H85N2O6P | 745.6218 | Positive | M+H | 1 | 5.03
SM 37:2;O2 | Target Compound | C42H83N2O6P | 743.6061 | Positive | M+H | 1 | 4.59
SM 38:0;O2 | Target Compound | C43H89N2O6P | 761.6531 | Positive | M+H | 1 | 5.51
SM 38:1;O2 | Target Compound | C43H87N2O6P | 759.6374 | Positive | M+H | 1 | 5.32
SM 38:2;O2 | Target Compound | C43H85N2O6P | 757.6218 | Positive | M+H | 1 | 4.86
SM 38:3;O2 | Target Compound | C43H83N2O6P | 755.6061 | Positive | M+H | 1 | 4.38
SM 39:1;O2 | Target Compound | C44H89N2O6P | 773.6531 | Positive | M+H | 1 | 5.6
SM 39:2;O2 | Target Compound | C44H87N2O6P | 771.6374 | Positive | M+H | 1 | 5.12
SM 40:0;O2 | Target Compound | C45H93N2O6P | 789.6844 | Positive | M+H | 1 | 6.03
SM 40:1;O2 | Target Compound | C45H91N2O6P | 787.6688 | Positive | M+H | 1 | 5.82
SM 40:2;O2 | Target Compound | C45H89N2O6P | 785.6531 | Positive | M+H | 1 | 5.37
SM 40:3;O2 | Target Compound | C45H87N2O6P | 783.6374 | Positive | M+H | 1 | 4.86
SM 40:4;O2 | Target Compound | C45H85N2O6P | 781.6218 | Positive | M+H | 1 | 4.48
SM 41:0;O2 | Target Compound | C46H95N2O6P | 803.7001 | Positive | M+H | 1 | 6.34
SM 41:1;O2 | Target Compound | C46H93N2O6P | 801.6844 | Positive | M+H | 1 | 6.12
SM 41:2;O2 | Target Compound | C46H91N2O6P | 799.6688 | Positive | M+H | 1 | 5.67
SM 41:3;O2 | Target Compound | C46H89N2O6P | 797.6531 | Positive | M+H | 1 | 5.12
SM 42:0;O2 | Target Compound | C47H97N2O6P | 817.7157 | Positive | M+H | 1 | 6.65
SM 42:0;O3 | Target Compound | C47H97N2O7P | 833.7106 | Positive | M+H | 1 | 6.16
SM 42:1;O2 | Target Compound | C47H95N2O6P | 815.7001 | Positive | M+H | 1 | 6.43
SM 42:1;O3 | Target Compound | C47H95N2O7P | 831.695 | Positive | M+H | 1 | 5.57
SM 42:2;O2 | Target Compound | C47H93N2O6P | 813.6844 | Positive | M+H | 1 | 5.84
SM 42:3;O2 | Target Compound | C47H91N2O6P | 811.6688 | Positive | M+H | 1 | 5.37
SM 42:4;O2 | Target Compound | C47H89N2O6P | 809.6531 | Positive | M+H | 1 | 4.92
SM 42:5;O2 | Target Compound | C47H87N2O6P | 807.6374 | Positive | M+H | 1 | 4.62
SM 43:1;O2 | Target Compound | C48H97N2O6P | 829.7157 | Positive | M+H | 1 | 6.64
SM 43:2;O2 | Target Compound | C48H95N2O6P | 827.7001 | Positive | M+H | 1 | 6.05
SM 43:3;O2 | Target Compound | C48H93N2O6P | 825.6844 | Positive | M+H | 1 | 5.65
SM 44:1;O2 | Target Compound | C49H99N2O6P | 843.7314 | Positive | M+H | 1 | 7.03
SM 44:2;O2 | Target Compound | C49H97N2O6P | 841.7157 | Positive | M+H | 1 | 6.38
SM 44:3;O2 | Target Compound | C49H95N2O6P | 839.7001 | Positive | M+H | 1 | 5.92
ST 27:1;O | Target Compound | C27H46O | 369.3516 | Positive | M-OH | 1 | 3.96
TG 14:0/16:1/14:0-d5 | Internal Standard | C47H83D5O6 | 771.7233 | Positive | M+NH4 | 1 | 8.35
TG 15:0/18:1-d7/15:0 | Internal Standard | D7C51H89O6 | 829.7985 | Positive | M+NH4 | 1 | 9.19
TG 20:0/20:1/20:0-d5 | Internal Standard | C63H115D5O6 | 995.9737 | Positive | M+NH4 | 1 | 11
TG 24:0 | Target Compound | C27H50O6 | 488.3946 | Positive | M+NH4 | 1 | 3.7
TG 26:0 | Target Compound | C29H54O6 | 516.4259 | Positive | M+NH4 | 1 | 4.1
TG 28:0 | Target Compound | C31H58O6 | 544.4572 | Positive | M+NH4 | 1 | 4.57
TG 34:0 | Target Compound | C37H70O6 | 628.5511 | Positive | M+NH4 | 1 | 6.23
TG 36:1 | Target Compound | C39H72O6 | 654.5667 | Positive | M+NH4 | 1 | 6.33
TG 42:0 | Target Compound | C45H86O6 | 740.6763 | Positive | M+NH4 | 1 | 8.34
TG 42:1 | Target Compound | C45H84O6 | 738.6606 | Positive | M+NH4 | 1 | 7.93
TG 42:2 | Target Compound | C45H82O6 | 736.645 | Positive | M+NH4 | 1 | 7.49
TG 43:0 | Target Compound | C46H88O6 | 754.6919 | Positive | M+NH4 | 1 | 8.54
TG 44:0 | Target Compound | C47H90O6 | 768.7076 | Positive | M+NH4 | 1 | 8.77
TG 44:1 | Target Compound | C47H88O6 | 766.6919 | Positive | M+NH4 | 1 | 8.4
TG 44:2 | Target Compound | C47H86O6 | 764.6763 | Positive | M+NH4 | 1 | 7.99
TG 44:3 | Target Compound | C47H84O6 | 762.6606 | Positive | M+NH4 | 1 | 7.57
TG 45:0 | Target Compound | C48H92O6 | 782.7232 | Positive | M+NH4 | 1 | 8.95
TG 45:1 | Target Compound | C48H90O6 | 780.7076 | Positive | M+NH4 | 1 | 8.62
TG 46:0 | Target Compound | C49H94O6 | 796.7389 | Positive | M+NH4 | 1 | 9.16
TG 46:1 | Target Compound | C49H92O6 | 794.7232 | Positive | M+NH4 | 1 | 8.82
TG 46:2 | Target Compound | C49H90O6 | 792.7076 | Positive | M+NH4 | 1 | 8.44
TG 46:3 | Target Compound | C49H88O6 | 790.6919 | Positive | M+NH4 | 1 | 8.04
TG 47:0 | Target Compound | C50H96O6 | 810.7545 | Positive | M+NH4 | 1 | 9.35
TG 47:1 | Target Compound | C50H94O6 | 808.7389 | Positive | M+NH4 | 1 | 9.01
TG 47:2 | Target Compound | C50H92O6 | 806.7232 | Positive | M+NH4 | 1 | 8.66
TG 48:0 | Target Compound | C51H98O6 | 824.7702 | Positive | M+NH4 | 1 | 9.53
TG 48:1 | Target Compound | C51H96O6 | 822.7545 | Positive | M+NH4 | 1 | 9.2
TG 48:2 | Target Compound | C51H94O6 | 820.7389 | Positive | M+NH4 | 1 | 8.86
TG 48:3 | Target Compound | C51H92O6 | 818.7232 | Positive | M+NH4 | 1 | 8.48
TG 48:4 | Target Compound | C51H90O6 | 816.7076 | Positive | M+NH4 | 1 | 8.09
TG 49:0 | Target Compound | C52H100O6 | 838.7858 | Positive | M+NH4 | 1 | 9.7
TG 49:1 | Target Compound | C52H98O6 | 836.7702 | Positive | M+NH4 | 1 | 9.39
TG 49:2 | Target Compound | C52H96O6 | 834.7545 | Positive | M+NH4 | 1 | 9.05
TG 49:3 | Target Compound | C52H94O6 | 832.7389 | Positive | M+NH4 | 1 | 8.7
TG 49:4 | Target Compound | C52H92O6 | 830.7232 | Positive | M+NH4 | 1 | 8.43
TG 50:0 | Target Compound | C53H102O6 | 852.8015 | Positive | M+NH4 | 1 | 9.88
TG 50:1 | Target Compound | C53H100O6 | 850.7858 | Positive | M+NH4 | 1 | 9.57
TG 50:2 | Target Compound | C53H98O6 | 848.7702 | Positive | M+NH4 | 1 | 9.25
TG 50:3 | Target Compound | C53H96O6 | 846.7545 | Positive | M+NH4 | 1 | 8.9
TG 50:4 | Target Compound | C53H94O6 | 844.7389 | Positive | M+NH4 | 1 | 8.54
TG 50:5 | Target Compound | C53H92O6 | 842.7232 | Positive | M+NH4 | 1 | 8.15
TG 51:0 | Target Compound | C54H104O6 | 866.8171 | Positive | M+NH4 | 1 | 10.02
TG 51:1 | Target Compound | C54H102O6 | 864.8015 | Positive | M+NH4 | 1 | 9.74
TG 51:2 | Target Compound | C54H100O6 | 862.7858 | Positive | M+NH4 | 1 | 9.45
TG 51:3 | Target Compound | C54H98O6 | 860.7702 | Positive | M+NH4 | 1 | 9.11
TG 51:4 | Target Compound | C54H96O6 | 858.7545 | Positive | M+NH4 | 1 | 8.75
TG 52:0 | Target Compound | C55H106O6 | 880.8328 | Positive | M+NH4 | 1 | 10.17
TG 52:1 | Target Compound | C55H104O6 | 878.8171 | Positive | M+NH4 | 1 | 9.91
TG 52:2 | Target Compound | C55H102O6 | 876.8015 | Positive | M+NH4 | 1 | 9.61
TG 52:3 | Target Compound | C55H100O6 | 874.7858 | Positive | M+NH4 | 1 | 9.29
TG 52:4 | Target Compound | C55H98O6 | 872.7702 | Positive | M+NH4 | 1 | 8.95
TG 52:5 | Target Compound | C55H96O6 | 870.7545 | Positive | M+NH4 | 1 | 8.6
TG 52:6 | Target Compound | C55H94O6 | 868.7389 | Positive | M+NH4 | 1 | 8.35
TG 52:7 | Target Compound | C55H92O6 | 866.7232 | Positive | M+NH4 | 1 | 8.05
TG 53:0 | Target Compound | C56H108O6 | 894.8484 | Positive | M+NH4 | 1 | 10.33
TG 53:1 | Target Compound | C56H106O6 | 892.8328 | Positive | M+NH4 | 1 | 10.06
TG 53:2 | Target Compound | C56H104O6 | 890.8171 | Positive | M+NH4 | 1 | 9.77
TG 53:3 | Target Compound | C56H102O6 | 888.8015 | Positive | M+NH4 | 1 | 9.47
TG 53:4 | Target Compound | C56H100O6 | 886.7858 | Positive | M+NH4 | 1 | 9.1
TG 53:5 | Target Compound | C56H98O6 | 884.7702 | Positive | M+NH4 | 1 | 8.8
TG 54:0 | Target Compound | C57H110O6 | 908.8641 | Positive | M+NH4 | 1 | 10.44
TG 54:1 | Target Compound | C57H108O6 | 906.8484 | Positive | M+NH4 | 1 | 10.23
TG 54:2 | Target Compound | C57H106O6 | 904.8328 | Positive | M+NH4 | 1 | 9.95
TG 54:3 | Target Compound | C57H104O6 | 902.8171 | Positive | M+NH4 | 1 | 9.67
TG 54:4 | Target Compound | C57H102O6 | 900.8015 | Positive | M+NH4 | 1 | 9.33
TG 54:5 | Target Compound | C57H100O6 | 898.7858 | Positive | M+NH4 | 1 | 9.05
TG 54:6 | Target Compound | C57H98O6 | 896.7702 | Positive | M+NH4 | 1 | 8.76
TG 54:7 | Target Compound | C57H96O6 | 894.7545 | Positive | M+NH4 | 1 | 8.46
TG 55:0 | Target Compound | C58H112O6 | 922.8797 | Positive | M+NH4 | 1 | 10.66
TG 55:1 | Target Compound | C58H110O6 | 920.8641 | Positive | M+NH4 | 1 | 10.37
TG 55:2 | Target Compound | C58H108O6 | 918.8484 | Positive | M+NH4 | 1 | 10.1
TG 55:3 | Target Compound | C58H106O6 | 916.8328 | Positive | M+NH4 | 1 | 9.82
TG 55:4 | Target Compound | C58H104O6 | 914.8171 | Positive | M+NH4 | 1 | 9.53
TG 55:5 | Target Compound | C58H102O6 | 912.8015 | Positive | M+NH4 | 1 | 9.28
TG 55:6 | Target Compound | C58H100O6 | 910.7858 | Positive | M+NH4 | 1 | 9
TG 56:0 | Target Compound | C59H114O6 | 936.8954 | Positive | M+NH4 | 1 | 10.79
TG 56:1 | Target Compound | C59H112O6 | 934.8797 | Positive | M+NH4 | 1 | 10.52
TG 56:2 | Target Compound | C59H110O6 | 932.8641 | Positive | M+NH4 | 1 | 10.26
TG 56:3 | Target Compound | C59H108O6 | 930.8484 | Positive | M+NH4 | 1 | 9.98
TG 56:4 | Target Compound | C59H106O6 | 928.8328 | Positive | M+NH4 | 1 | 9.69
TG 56:5 | Target Compound | C59H104O6 | 926.8171 | Positive | M+NH4 | 1 | 9.53
TG 56:6 | Target Compound | C59H102O6 | 924.8015 | Positive | M+NH4 | 1 | 9.2
TG 56:7 | Target Compound | C59H100O6 | 922.7858 | Positive | M+NH4 | 1 | 8.92
TG 56:8 | Target Compound | C59H98O6 | 920.7702 | Positive | M+NH4 | 1 | 8.57
TG 56:9 | Target Compound | C59H96O6 | 918.7545 | Positive | M+NH4 | 1 | 8.21
TG 57:1 | Target Compound | C60H114O6 | 948.8954 | Positive | M+NH4 | 1 | 10.66
TG 57:2 | Target Compound | C60H112O6 | 946.8797 | Positive | M+NH4 | 1 | 10.4
TG 57:3 | Target Compound | C60H110O6 | 944.8641 | Positive | M+NH4 | 1 | 10.13
TG 57:4 | Target Compound | C60H108O6 | 942.8484 | Positive | M+NH4 | 1 | 9.85
TG 58:1 | Target Compound | C61H116O6 | 962.911 | Positive | M+NH4 | 1 | 10.8
TG 58:10 | Target Compound | C61H98O6 | 944.7702 | Positive | M+NH4 | 1 | 8.4
TG 58:11 | Target Compound | C61H96O6 | 942.7545 | Positive | M+NH4 | 1 | 8
TG 58:2 | Target Compound | C61H114O6 | 960.8954 | Positive | M+NH4 | 1 | 10.55
TG 58:3 | Target Compound | C61H112O6 | 958.8797 | Positive | M+NH4 | 1 | 10.28
TG 58:4 | Target Compound | C61H110O6 | 956.8641 | Positive | M+NH4 | 1 | 10
TG 58:5 | Target Compound | C61H108O6 | 954.8484 | Positive | M+NH4 | 1 | 9.74
TG 58:6 | Target Compound | C61H106O6 | 952.8328 | Positive | M+NH4 | 1 | 9.47
TG 58:7 | Target Compound | C61H104O6 | 950.8171 | Positive | M+NH4 | 1 | 9.17
TG 58:8 | Target Compound | C61H102O6 | 948.8015 | Positive | M+NH4 | 1 | 9
TG 58:9 | Target Compound | C61H100O6 | 946.7858 | Positive | M+NH4 | 1 | 8.66
TG 59:2 | Target Compound | C62H116O6 | 974.911 | Positive | M+NH4 | 1 | 10.69
TG 59:3 | Target Compound | C62H114O6 | 972.8954 | Positive | M+NH4 | 1 | 10.42
TG 59:4 | Target Compound | C62H112O6 | 970.8797 | Positive | M+NH4 | 1 | 10.15
TG 60:10 | Target Compound | C63H102O6 | 972.8015 | Positive | M+NH4 | 1 | 8.85
TG 60:11 | Target Compound | C63H100O6 | 970.7858 | Positive | M+NH4 | 1 | 8.53
TG 60:12 | Target Compound | C63H98O6 | 968.7702 | Positive | M+NH4 | 1 | 8.2
TG 60:2 | Target Compound | C63H118O6 | 988.9267 | Positive | M+NH4 | 1 | 10.82
TG 60:3 | Target Compound | C63H116O6 | 986.911 | Positive | M+NH4 | 1 | 10.55
TG 60:4 | Target Compound | C63H114O6 | 984.8954 | Positive | M+NH4 | 1 | 10.29
TG 60:5 | Target Compound | C63H112O6 | 982.8797 | Positive | M+NH4 | 1 | 10.02
TG 60:6 | Target Compound | C63H110O6 | 980.8641 | Positive | M+NH4 | 1 | 9.75
TG 60:7 | Target Compound | C63H108O6 | 978.8484 | Positive | M+NH4 | 1 | 9.59
TG 60:8 | Target Compound | C63H106O6 | 976.8328 | Positive | M+NH4 | 1 | 9.28
TG 62:2 | Target Compound | C65H122O6 | 1016.958 | Positive | M+NH4 | 1 | 11.07
TG 63:3 | Target Compound | C65H120O6 | 1014.9423 | Positive | M+NH4 | 1 | 10.85
TG 70:3 | Target Compound | C73H136O6 | 1127.0675 | Positive | M+NH4 | 1 | 11.71
TG 72:3 | Target Compound | C75H140O6 | 1155.0988 | Positive | M+NH4 | 1 | 11.93
TG O-50:1 | Target Compound | C53H102O5 | 836.8066 | Positive | M+NH4 | 1 | 10.04
TG O-50:2 | Target Compound | C55H104O5 | 862.8222 | Positive | M+NH4 | 1 | 10.06
TG O-50:3 | Target Compound | C55H102O5 | 860.8066 | Positive | M+NH4 | 1 | 9.78
TG O-52:3 | Target Compound | C57H106O5 | 888.8379 | Positive | M+NH4 | 1 | 10.09
Tributyl citrate | Target Compound | C18H32O7 | 361.2221 | Positive | M+H | 1 | 2.47
Ubiquinone-10 | Target Compound | C59H90O4 | 863.6912 | Positive | M+H | 1 | 8.34
Arachidonic acid-d8 | Internal Standard | D8C20H24O2 | 311.2832 | Negative | M-H | 1 | 2.89
Cer 18:1;O2/16:0-d7 | Internal Standard | D7C34H60NO3 | 589.5542 | Negative | M+HCO2 | 1 | 4.9
FA 14:0 | Target Compound | C14H28O2 | 227.2017 | Negative | M-H | 1 | 2.28
FA 15:0 | Target Compound | C15H30O2 | 241.2173 | Negative | M-H | 1 | 2.49
FA 16:0 | Target Compound | C16H32O2 | 255.2329 | Negative | M-H | 1 | 3.08
FA 16:1 | Target Compound | C16H30O2 | 253.2173 | Negative | M-H | 1 | 2.88
FA 17:0 | Target Compound | C17H34O2 | 269.2486 | Negative | M-H | 1 | 3.23
FA 17:1 | Target Compound | C17H32O2 | 267.233 | Negative | M-H | 1 | 3.03
FA 18:0 | Target Compound | C18H36O2 | 283.2643 | Negative | M-H | 1 | 3.41
FA 18:1 | Target Compound | C18H34O2 | 281.2486 | Negative | M-H | 1 | 3.14
FA 18:2 | Target Compound | C18H32O2 | 279.233 | Negative | M-H | 1 | 2.94
FA 18:3 | Target Compound | C18H30O2 | 277.2173 | Negative | M-H | 1 | 2.78
FA 18:4 | Target Compound | C18H28O2 | 275.2016 | Negative | M-H | 1 | 2.64
FA 20:1 | Target Compound | C20H38O2 | 309.2799 | Negative | M-H | 1 | 3.46
FA 20:2 | Target Compound | C20H36O2 | 307.2643 | Negative | M-H | 1 | 3.21
FA 20:4 | Target Compound | C20H32O2 | 303.233 | Negative | M-H | 1 | 2.89
FA 20:5 | Target Compound | C20H30O2 | 301.2173 | Negative | M-H | 1 | 2.75
FA 22:4 | Target Compound | C22H36O2 | 331.2643 | Negative | M-H | 1 | 3.09
FA 22:6 | Target Compound | C22H32O2 | 327.233 | Negative | M-H | 1 | 2.83
FA 24:0 | Target Compound | C24H48O2 | 367.3582 | Negative | M-H | 1 | 5.29
FA 24:1 | Target Compound | C24H46O2 | 365.3425 | Negative | M-H | 1 | 4.4
LPC 18:1-d7 | Internal Standard | D7C26H45NO7P | 573.3903 | Negative | M+HCO2 | 1 | 2.65
LPC O-16:0 | Target Compound | C24H52NO6P | 526.3514 | Negative | M+HCO2 | 1 | 2.69
LPC O-16:0-d4 | Internal Standard | D4C24H48NO6P | 530.3765 | Negative | M+HCO2 | 1 | 2.69
LPC O-16:1 | Target Compound | C24H50NO6P | 524.3358 | Negative | M+HCO2 | 1 | 2.69
LPC O-17:1 | Target Compound | C25H52NO6P | 538.3514 | Negative | M+HCO2 | 1 | 2.76
LPC O-18:0 | Target Compound | C26H56NO6P | 554.3827 | Negative | M+HCO2 | 1 | 2.89
LPC O-18:1 | Target Compound | C26H54NO6P | 552.3671 | Negative | M+HCO2 | 1 | 2.76
LPC O-18:2 | Target Compound | C26H52NO6P | 550.3514 | Negative | M+HCO2 | 1 | 2.72
LPC O-20:0 | Target Compound | C28H58NO6P | 580.3984 | Negative | M+HCO2 | 1 | 2.93
LPE 16:0 | Target Compound | C21H44NO7P | 452.2783 | Negative | M-H | 1 | 2.64
LPE 16:1 | Target Compound | C21H42NO7P | 450.2626 | Negative | M-H | 1 | 2.53
LPE 17:0 | Target Compound | C22H46NO7P | 466.2939 | Negative | M-H | 1 | 2.73
LPE 17:2 | Target Compound | C22H42NO7P | 462.2626 | Negative | M-H | 1 | 2.5
LPE 18:0 | Target Compound | C23H48NO7P | 480.3096 | Negative | M-H | 1 | 2.82
LPE 18:1 | Target Compound | C23H46NO7P | 478.2939 | Negative | M-H | 1 | 2.69
LPE 18:1-d7 | Internal Standard | D7C23H39NO7P | 485.3379 | Negative | M-H | 1 | 2.68
LPE 18:2 | Target Compound | C23H44NO7P | 476.2783 | Negative | M-H | 1 | 2.57
LPE 18:3 | Target Compound | C23H42NO7P | 474.2626 | Negative | M-H | 1 | 2.47
LPE 20:0 | Target Compound | C25H52NO7P | 508.3409 | Negative | M-H | 1 | 3.06
LPE 20:1 | Target Compound | C25H50NO7P | 506.3252 | Negative | M-H | 1 | 2.86
LPE 20:2 | Target Compound | C25H48NO7P | 504.3096 | Negative | M-H | 1 | 2.74
LPE 20:3 | Target Compound | C25H46NO7P | 502.2939 | Negative | M-H | 1 | 2.62
LPE 20:4 | Target Compound | C25H44NO7P | 500.2783 | Negative | M-H | 1 | 2.54
LPE 20:5 | Target Compound | C25H42NO7P | 498.2626 | Negative | M-H | 1 | 2.45
LPE 22:0 | Target Compound | C27H56NO7P | 536.3722 | Negative | M-H | 1 | 3.15
LPE 22:3 | Target Compound | C27H50NO7P | 530.3252 | Negative | M-H | 1 | 2.7
LPE 22:4 | Target Compound | C27H48NO7P | 528.3096 | Negative | M-H | 1 | 2.68
LPE 22:5 | Target Compound | C27H46NO7P | 526.2939 | Negative | M-H | 1 | 2.55
LPE 22:6 | Target Compound | C27H44NO7P | 524.2783 | Negative | M-H | 1 | 2.51
LPE 24:0 | Target Compound | C29H60NO7P | 564.4035 | Negative | M-H | 1 | 3.29
LPE O-16:1 | Target Compound | C21H44NO6P | 436.2833 | Negative | M-H | 1 | 2.73
LPE O-18:1 | Target Compound | C23H48NO6P | 464.3146 | Negative | M-H | 1 | 2.95
LPE O-18:2 | Target Compound | C23H46NO6P | 462.299 | Negative | M-H | 1 | 2.76
LPE O-20:1 | Target Compound | C25H52NO6P | 492.346 | Negative | M-H | 1 | 3.21
LPG 16:0 | Target Compound | C22H45O9P | 483.2728 | Negative | M-H | 1 | 2.53
LPG 17:1 | Internal Standard | C23H45O9P | 495.2728 | Negative | M-H | 1 | 2.49
LPG 18:0 | Target Compound | C24H49O9P | 511.3041 | Negative | M-H | 1 | 2.69
LPG 18:1 | Target Compound | C24H47O9P | 509.2885 | Negative | M-H | 1 | 2.57
LPG 18:2 | Target Compound | C24H45O9P | 507.2728 | Negative | M-H | 1 | 2.47
LPI 16:0 | Target Compound | C25H49O12P | 571.2889 | Negative | M-H | 1 | 2.5
LPI 17:1 | Internal Standard | C26H49O12P | 583.2889 | Negative | M-H | 1 | 2.46
LPI 18:0 | Target Compound | C27H53O12P | 599.3202 | Negative | M-H | 1 | 2.65
LPI 18:1 | Target Compound | C27H51O12P | 597.3045 | Negative | M-H | 1 | 2.53
LPI 18:2 | Target Compound | C27H49O12P | 595.2889 | Negative | M-H | 1 | 2.44
LPI 20:3 | Target Compound | C29H51O12P | 621.3045 | Negative | M-H | 1 | 2.49
LPI 20:4 | Target Compound | C29H49O12P | 619.2889 | Negative | M-H | 1 | 2.43
LPS 18:0 | Target Compound | C24H48NO9P | 524.2994 | Negative | M-H | 1 | 2.62
LPS 18:1 | Target Compound | C24H46NO9P | 522.2837 | Negative | M-H | 1 | 2.45
LPS 20:4 | Target Compound | C26H44NO9P | 544.2681 | Negative | M-H | 1 | 2.37
LPS 22:6 | Target Compound | C28H44NO9P | 568.2681 | Negative | M-H | 1 | 2.35
NeuAcHex2Cer 34:1;O2 | Target Compound | C57H104N2O21 | 1151.7059 | Negative | M-H | 1 | 3.69
PA 34:1 | Target Compound | C37H71O8P | 673.4814 | Negative | M-H | 1 | 4.61
PA 36:1 | Target Compound | C39H75O8P | 701.5127 | Negative | M-H | 1 | 5.14
PA 36:2 | Target Compound | C39H73O8P | 699.497 | Negative | M-H | 1 | 4.7
PC 15:0/18:1-d7 | Internal Standard | D7C41H73NO8P | 797.6043 | Negative | M+HCO2 | 1 | 4.47
PC 34:1;O | Target Compound | C42H80NO9P | 818.5553 | Negative | M+HCO2 | 1 | 3.27
PC 34:2;O2 | Target Compound | C42H80NO10P | 834.5502 | Negative | M+HCO2 | 1 | 3.23
PC 36:1;O | Target Compound | C44H84NO9P | 846.5866 | Negative | M+HCO2 | 1 | 3.6
PC 36:2;O | Target Compound | C44H82NO9P | 844.5709 | Negative | M+HCO2 | 1 | 3.38
PC 36:3;O2 | Target Compound | C44H80NO10P | 858.5502 | Negative | M+HCO2 | 1 | 2.89
PC 36:3;O | Target Compound | C44H80NO9P | 842.5553 | Negative | M+HCO2 | 1 | 3.5
PC 38:5;O | Target Compound | C46H80NO9P | 866.5553 | Negative | M+HCO2 | 1 | 3.28
PC O-18:0/18:1-d9 | Internal Standard | D9C44H77NO7P | 825.6689 | Negative | M+HCO2 | 1 | 5.51
PE 15:0/18:1-d7 | Internal Standard | D7C38H67NO8P | 709.5519 | Negative | M-H | 1 | 4.63
PE 32:0 | Target Compound | C37H74NO8P | 690.5079 | Negative | M-H | 1 | 4.82
PE 32:1 | Target Compound | C37H72NO8P | 688.4923 | Negative | M-H | 1 | 4.42
PE 34:1 | Target Compound | C39H76NO8P | 716.5236 | Negative | M-H | 1 | 4.9
PE 34:2 | Target Compound | C39H74NO8P | 714.5079 | Negative | M-H | 1 | 4.51
PE 34:3 | Target Compound | C39H72NO8P | 712.4923 | Negative | M-H | 1 | 4.16
PE 35:2 | Target Compound | C40H76NO8P | 728.5236 | Negative | M-H | 1 | 4.74
PE 36:1 | Target Compound | C41H80NO8P | 744.5549 | Negative | M-H | 1 | 5.41
PE 36:2 | Target Compound | C41H78NO8P | 742.5392 | Negative | M-H | 1 | 4.99
PE 36:3 | Target Compound | C41H76NO8P | 740.5236 | Negative | M-H | 1 | 4.59
PE 36:4 | Target Compound | C41H74NO8P | 738.5079 | Negative | M-H | 1 | 4.43
PE 36:5 | Target Compound | C41H72NO8P | 736.4923 | Negative | M-H | 1 | 4.1
PE 37:4 | Target Compound | C42H76NO8P | 752.5236 | Negative | M-H | 1 | 4.66
PE 38:2 | Target Compound | C43H82NO8P | 770.5705 | Negative | M-H | 1 | 5.52
PE 38:4 | Target Compound | C43H78NO8P | 766.5392 | Negative | M-H | 1 | 4.92
PE 38:6 | Target Compound | C43H74NO8P | 762.5079 | Negative | M-H | 1 | 4.27
PE 40:1 | Target Compound | C45H88NO8P | 800.6175 | Negative | M-H | 1 | 6.4
PE 40:3 | Target Compound | C45H84NO8P | 796.5862 | Negative | M-H | 1 | 5.45
PE 40:4 | Target Compound | C45H82NO8P | 794.5705 | Negative | M-H | 1 | 5.16
PE 40:5 | Target Compound | C45H80NO8P | 792.5549 | Negative | M-H | 1 | 4.89
PE 40:6 | Target Compound | C45H78NO8P | 790.5392 | Negative | M-H | 1 | 4.74
PE 40:7 | Target Compound | C45H76NO8P | 788.5236 | Negative | M-H | 1 | 4.36
PE 40:8 | Target Compound | C45H74NO8P | 786.5079 | Negative | M-H | 1 | 4.07
PE 41:1 | Target Compound | C46H90NO8P | 814.6331 | Negative | M-H | 1 | 6.6
PE 42:1 | Target Compound | C47H92NO8P | 828.6488 | Negative | M-H | 1 | 6.89
PE 42:10 | Target Compound | C47H74NO8P | 810.5079 | Negative | M-H | 1 | 3.95
PE 42:7 | Target Compound | C47H80NO8P | 816.5549 | Negative | M-H | 1 | 4.81
PE 42:8 | Target Compound | C47H78NO8P | 814.5392 | Negative | M-H | 1 | 4.23
PE 44:10 | Target Compound | C49H78NO8P | 838.5392 | Negative | M-H | 1 | 4.12
PE 44:11 | Target Compound | C49H76NO8P | 836.5236 | Negative | M-H | 1 | 3.83
PE O-16:1_16:0 | Target Compound | C37H74NO7P | 674.513 | Negative | M-H | 1 | 5.15
PE O-16:1_18:1 | Target Compound | C39H76NO7P | 700.5287 | Negative | M-H | 1 | 5.26
PE O-16:1_18:2 | Target Compound | C39H74NO7P | 698.513 | Negative | M-H | 1 | 4.8
PE O-16:1_20:4 | Target Compound | C41H74NO7P | 722.513 | Negative | M-H | 1 | 4.71
PE O-16:1_22:6 | Target Compound | C43H74NO7P | 746.513 | Negative | M-H | 1 | 4.54
PE O-18:1_16:0 | Target Compound | C39H78NO7P | 702.5443 | Negative | M-H | 1 | 5.67
PE O-18:1_18:1 | Target Compound | C41H80NO7P | 728.56 | Negative | M-H | 1 | 5.77
PE O-18:1_18:2 | Target Compound | C41H78NO7P | 726.5443 | Negative | M-H | 1 | 5.32
PE O-18:1_20:4 | Target Compound | C43H78NO7P | 750.5443 | Negative | M-H | 1 | 5.24
PE O-18:1_22:6 | Target Compound | C45H78NO7P | 774.5443 | Negative | M-H | 1 | 5.06
PE O-18:2_14:0 | Target Compound | C37H72NO7P | 672.4974 | Negative | M-H | 1 | 4.74
PE O-18:2_18:2 | Target Compound | C41H76NO7P | 724.5287 | Negative | M-H | 1 | 4.89
PE O-18:2_20:4 | Target Compound | C43H76NO7P | 748.5287 | Negative | M-H | 1 | 4.81
PE O-35:2 | Target Compound | C40H78NO7P | 714.5443 | Negative | M-H | 1 | 5.52
PE O-35:3 | Target Compound | C40H76NO7P | 712.5287 | Negative | M-H | 1 | 5.05
PE O-36:1 | Target Compound | C41H82NO7P | 730.5756 | Negative | M-H | 1 | 5.84
PE O-36:6 | Target Compound | C41H72NO7P | 720.4974 | Negative | M-H | 1 | 4.35
PE O-38:1 | Target Compound | C43H86NO7P | 758.6069 | Negative | M-H | 1 | 6.24
PE O-38:2 | Target Compound | C43H84NO7P | 756.5913 | Negative | M-H | 1 | 6.3
PE O-38:3 | Target Compound | C43H82NO7P | 754.5756 | Negative | M-H | 1 | 5.9
PE O-38:4 | Target Compound | C43H80NO7P | 752.56 | Negative | M-H | 1 | 5.45
PE O-40:2 | Target Compound | C45H88NO7P | 784.6226 | Negative | M-H | 1 | 6.85
PE O-40:3 | Target Compound | C45H86NO7P | 782.6069 | Negative | M-H | 1 | 6.34
PE O-40:5 | Target Compound | C45H82NO7P | 778.5756 | Negative | M-H | 1 | 5.79
PE O-40:6 | Target Compound | C45H80NO7P | 776.56 | Negative | M-H | 1 | 5.24
PE O-40:8 | Target Compound | C45H76NO7P | 772.5287 | Negative | M-H | 1 | 4.58
PE O-42:3 | Target Compound | C47H90NO7P | 810.6382 | Negative | M-H | 1 | 6.91
PE O-42:5 | Target Compound | C47H86NO7P | 806.6069 | Negative | M-H | 1 | 6.38
PE O-42:6 | Target Compound | C47H84NO7P | 804.5913 | Negative | M-H | 1 | 5.76
PG 15:0/18:1-d7 | Internal Standard | D7C39H68O10P | 740.5465 | Negative | M-H | 1 | 4
PG 30:0 | Target Compound | C36H71O10P | 693.4712 | Negative | M-H | 1 | 3.7
PG 32:0 | Target Compound | C38H75O10P | 721.5025 | Negative | M-H | 1 | 4.07
PG 32:1 | Target Compound | C38H73O10P | 719.4869 | Negative | M-H | 1 | 3.87
PG 34:1 | Target Compound | C40H77O10P | 747.5182 | Negative | M-H | 1 | 4.52
PG 34:2 | Target Compound | C40H75O10P | 745.5025 | Negative | M-H | 1 | 3.95
PG 36:1 | Target Compound | C42H81O10P | 775.5495 | Negative | M-H | 1 | 5.02
PG 36:2 | Target Compound | C42H79O10P | 773.5338 | Negative | M-H | 1 | 4.6
PG 36:3 | Target Compound | C42H77O10P | 771.5182 | Negative | M-H | 1 | 4.23
PG 36:4 | Target Compound | C42H75O10P | 769.5025 | Negative | M-H | 1 | 3.84
PG 38:4 | Target Compound | C44H79O10P | 797.5338 | Negative | M-H | 1 | 4.23
PG 38:5 | Target Compound | C44H77O10P | 795.5182 | Negative | M-H | 1 | 3.76
PG 38:6 | Target Compound | C44H75O10P | 793.5025 | Negative | M-H | 1 | 3.5
PG 40:6 | Target Compound | C46H79O10P | 821.5338 | Negative | M-H | 1 | 3.95
PG 40:7 | Target Compound | C46H77O10P | 819.5182 | Negative | M-H | 1 | 3.65
PG 40:8 | Target Compound | C46H75O10P | 817.5025 | Negative | M-H | 1 | 3.46
PI 15:0/18:1-d7 | Internal Standard | D7C42H72O13P | 828.5625 | Negative | M-H | 1 | 3.94
PI 32:0 | Target Compound | C41H79O13P | 809.5186 | Negative | M-H | 1 | 4.04
PI 32:1 | Target Compound | C41H77O13P | 807.5029 | Negative | M-H | 1 | 3.78
PI 33:1 | Target Compound | C42H79O13P | 821.5186 | Negative | M-H | 1 | 3.95
PI 34:0 | Target Compound | C43H83O13P | 837.5498 | Negative | M-H | 1 | 4.43
PI 34:1 | Target Compound | C43H81O13P | 835.5342 | Negative | M-H | 1 | 4.12
PI 34:2 | Target Compound | C43H79O13P | 833.5186 | Negative | M-H | 1 | 3.83
PI 35:1 | Target Compound | C44H83O13P | 849.5498 | Negative | M-H | 1 | 4.31
PI 35:2 | Target Compound | C44H81O13P | 847.5342 | Negative | M-H | 1 | 4.01
PI 36:1 | Target Compound | C45H85O13P | 863.5655 | Negative | M-H | 1 | 4.54
PI 36:2 | Target Compound | C45H83O13P | 861.5498 | Negative | M-H | 1 | 4.21
PI 36:3 | Target Compound | C45H81O13P | 859.5342 | Negative | M-H | 1 | 3.93
PI 36:4 | Target Compound | C45H79O13P | 857.5186 | Negative | M-H | 1 | 3.79
PI 36:5 | Target Compound | C45H77O13P | 855.5029 | Negative | M-H | 1 | 3.55
PI 37:4 | Target Compound | C46H81O13P | 871.5342 | Negative | M-H | 1 | 3.97
PI 38:3 | Target Compound | C47H85O13P | 887.5655 | Negative | M-H | 1 | 4.31
PI 38:4 | Target Compound | C47H83O13P | 885.5498 | Negative | M-H | 1 | 4.17
PI 38:5 | Target Compound | C47H81O13P | 883.5342 | Negative | M-H | 1 | 3.87
PI 38:6 | Target Compound | C47H79O13P | 881.5186 | Negative | M-H | 1 | 3.68
PI 40:5 | Target Compound | C49H85O13P | 911.5655 | Negative | M-H | 1 | 4.14
PI 40:6 | Target Compound | C49H83O13P | 909.5498 | Negative | M-H | 1 | 4.02
PI O-18:1_20:4 | Target Compound | C47H83O12P | 869.5549 | Negative | M-H | 1 | 4.12
PS 15:0/18:1-d7 | Internal Standard | D7C39H67NO10P | 753.5417 | Negative | M-H | 1 | 3.99
PS 34:1 | Target Compound | C40 H76 N O10 P | 760.5134 | Negative | M-H | 1 | 4.17
PS 36:1 | Target Compound | C42H80NO10P | 788.5447 | Negative | M-H | 1 | 4.59
PS 36:2 | Target Compound | C42H78NO10P | 786.5291 | Negative | M-H | 1 | 4.27
PS 36:4 | Target Compound | C42 H74 N O10 P | 782.4978 | Negative | M-H | 1 | 3.84
PS 38:1 | Target Compound | C44 H84 N O10 P | 816.576 | Negative | M-H | 1 | 5.02
PS 38:2 | Target Compound | C44H82NO10P | 814.5604 | Negative | M-H | 1 | 4.71
PS 38:3 | Target Compound | C44 H80 N O10 P | 812.5447 | Negative | M-H | 1 | 4.39
PS 38:4 | Target Compound | C44 H78 N O10 P | 810.5291 | Negative | M-H | 1 | 4.22
PS 38:5 | Target Compound | C44 H76 N O10 P | 808.5134 | Negative | M-H | 1 | 3.92
PS 38:6 | Target Compound | C44H74NO10P | 806.4978 | Negative | M-H | 1 | 3.72
PS 40:1 | Target Compound | C46H88NO10P | 844.6073 | Negative | M-H | 1 | 5.64
PS 40:2 | Target Compound | C46H86NO10P | 842.5917 | Negative | M-H | 1 | 5.1
PS 40:4 | Target Compound | C46 H82 N O10 P | 838.5604 | Negative | M-H | 1 | 4.5
PS 40:5 | Target Compound | C46 H80 N O10 P | 836.5447 | Negative | M-H | 1 | 4.36
PS 40:6 | Target Compound | C46 H78 N O10 P | 834.5291 | Negative | M-H | 1 | 4.06
PS 40:7 | Target Compound | C46 H76 N O10 P | 832.5134 | Negative | M-H | 1 | 3.77
PS 42:1 | Target Compound | C48H92NO10P | 872.6386 | Negative | M-H | 1 | 6.4
PS 42:2 | Target Compound | C48H90NO10P | 870.623 | Negative | M-H | 1 | 5.69
PS 44:10 | Target Compound | C50 H78 N O10 P | 882.5291 | Negative | M-H | 1 | 3.66
PS 44:12 | Target Compound | C50 H74 N O10 P | 878.4978 | Negative | M-H | 1 | 3.31
SHexCer 36:1;O2 | Target Compound | C42H81NO11S | 806.5458 | Negative | M-H | 1 | 4.16
SHexCer 36:1;O3 | Target Compound | C42H81NO12S | 822.5407 | Negative | M-H | 1 | 4.05
SHexCer 38:1;O2 | Target Compound | C44H85NO11S | 834.5771 | Negative | M-H | 1 | 4.6
SHexCer 38:1;O3 | Target Compound | C44H85NO12S | 850.572 | Negative | M-H | 1 | 4.48
SHexCer 40:1;O2 | Target Compound | C46H89NO11S | 862.6084 | Negative | M-H | 1 | 5.07
SHexCer 40:1;O3 | Target Compound | C46H89NO12S | 878.6033 | Negative | M-H | 1 | 4.96
SHexCer 40:2;O2 | Target Compound | C46H87NO11S | 860.5927 | Negative | M-H | 1 | 4.61
SHexCer 41:1;O3 | Target Compound | C47H91NO12S | 892.6189 | Negative | M-H | 1 | 5.24
SHexCer 42:0;O2 | Target Compound | C48H95NO11S | 892.6553 | Negative | M-H | 1 | 5.9
SHexCer 42:1;O2 | Target Compound | C48H93NO11S | 890.6397 | Negative | M-H | 1 | 5.62
SHexCer 42:2;O2 | Target Compound | C48H91NO11S | 888.624 | Negative | M-H | 1 | 5.05
SHexCer 42:2;O3 | Target Compound | C48H91NO12S | 904.6189 | Negative | M-H | 1 | 5
SHexCer 42:3;O2 | Target Compound | C48H89NO11S | 886.6084 | Negative | M-H | 1 | 4.64
SHexCer 44:2;O2 | Target Compound | C50H95NO11S | 916.6553 | Negative | M-H | 1 | 5.6
SM 18:1;O2/18:1-d9 | Internal Standard | D9C41H72N2O6P | 782.6379 | Negative | M+HCO2 | 1 | 4.35
ST 19:2;O2;S | Target Compound | C19H28O5S | 367.1585 | Negative | M-H | 1 | 1.72
ST 19:2;O3;S | Target Compound | C19H28O6S | 383.1534 | Negative | M-H | 1 | 1.11
ST 27:1;O;S | Target Compound | C27H46O4S | 465.3044 | Negative | M-H | 1 | 2.88
Thyroxin | Target Compound | C15H11I4NO4 | 775.6794 | Negative | M-H | 1 | 1.88

## Lipidomics_plasma
Lipidomic analysis of human plasma
LC-MS/MS Instrumentation and Analysis
Instrumentation
HPLC: | Shimadzu Nexera X2
MS: | Sciex TripleTOF 6600
Mode | Duospray Ion source in positive and negative ESI mode
Column | Zorbax RRHD Eclipse Plus C8 2.1 x 50 mm, 1.8µm (Agilent)
Pre-column | Zorbax Eclipse Plus C8 2.1 mm, 1.8 µm (Agilent)
Analysis
Data Acquisition | Analyst TF 1.7.1
Data Evaluation | MultiQuant 3.02
Method | Scan + ddMS2
LC parameters
Solvents
A (positive ion mode) | Water + 0,1 % formic acid + 10 mM ammonium formate
A (negative ion mode) | Water + 0,1 % formic acid + 1 mM ammonium formate
B | Acetonitrile:isopropanol (2:3) + 0.1% formic acid
Gradient Elution
Time [min] | A [%] | B [%] | Flowrate [mL/min]
0 | 80 | 20 | 0.3
0.3 | 80 | 20 | 0.3
3.5 | 20 | 80 | 0.3
10.5 | 0 | 100 | 0.3
13 | 0 | 100 | 0.3
14.5 | 80 | 20 | 0.3
17 | 80 | 20 | 0.3
Injection Volume |  | Column Temperature
Positive Ionization Mode | 2 µL | 40 °C
Negative Ionization Mode | 5 µL
MS parameters
Parameter | Abbreviation | Negative Ionisation | Positive Ionisation
Accumulation time MS-Scan | - | 250 ms | 250 ms
Accumulation time MS/MS | - | 50 ms | 50 ms
Collission gas (Q2) | CAD | 4 psi | 4 psi
Curtain Gas | CUR | 40 psi | 40 psi
Temperature heater | TEM | 400 °C | 400 °C
Ionisatio voltage | ISVF | -3500 V | 4500 V
Gas 1 | GS 1 | 50 psi | 25 psi
Gas 2 | GS 2 | 50 psi | 25 psi
Mass range MS-Scan | - | 100–1000 m/z | 100–1000 m/z
Mass range MS/MS | - | 50–1000 m/z | 50–1000 m/z
Declustering Potential Scan | DP | -80 V | 80 V
Declustering Potential MS/MS | DP | -80 V | 80 V
Collision energy MS-Scan | CE | -5 V | 5 V
Collision energy MS/MS | CE/CES | -60 – -20 V | 20–60 V
Q1 resolution |  | unit | unit
MS/MS-Mode | - | high sensitivity | high sensitivity
IDA channels | - | 6 | 6
Accumulated detection windows | - | 4 | 4
Ion impulse time | IRD | 11 ms | 11 ms
Internal Standards
Isotopically labeled internal standard | Vendor |  | Concentration of standard working solution [ng/mL] | Volume of working solution [µL] | Sample Volume [µL]
Arachidonic acid-d8 | Cayman Chemicals |  | 8.33 | 150 | 20
CE 18:1-d7 | Avanti Polar Lipids |  | 4650 | 150 | 20
Cer d18:1/16:0-d7 | Avanti Polar Lipids |  | 6.67 | 150 | 20
Cholesterol-d7 | Avanti Polar Lipids |  | 1330 | 150 | 20
DG 15:0/18:1-d7 | Avanti Polar Lipids |  | 157 | 150 | 20
LacCer d18:1/17:0 | Avanti Polar Lipids |  | 53.3 | 150 | 20
LPC 18:1-d7 | Avanti Polar Lipids |  | 317 | 150 | 20
LPC O-16:0-d4 | Cayman Chemicals |  | 16.7 | 150 | 20
LPE 18:1-d7 | Avanti Polar Lipids |  | 6.67 | 150 | 20
LPG 17:1 | Avanti Polar Lipids |  | 31.7 | 150 | 20
LPI 17:1 | Avanti Polar Lipids |  | 14.2 | 150 | 20
PC 15:0/18:1-d7 | Avanti Polar Lipids |  | 2110 | 150 | 20
PC O-18:0/18:1-d9 | Avanti Polar Lipids |  | 104 | 150 | 20
PE 15:0/18:1-d7 | Avanti Polar Lipids |  | 66.7 | 150 | 20
PG 15:0/18:1-d7 | Avanti Polar Lipids |  | 63.3 | 150 | 20
PI 15:0/18:1-d7 | Avanti Polar Lipids |  | 113 | 150 | 20
PS 15:0/18:1-d7 | Avanti Polar Lipids |  | 261 | 150 | 20
SM d18:1/18:1-d9 | Avanti Polar Lipids |  | 393 | 150 | 20
TG 15:0/18:1-d7/15:0 | Avanti Polar Lipids |  | 759 | 150 | 20
*Avanti Polar Lipids | Alabaster, AL, USA
*Cayman Chemicals | Ann Arbor, MI, USA
Acceptance criteria
The following analytes were excluded from data evaluation:
- Analytes showing a mass error > 5 ppm
- Analytes without identity confirmation via spectral library matching or reference substance measurements
- Analytes containing only missing values in QC samples and study samples
- Analytes containing > 50% missing values in QC samples and study samples
- Analytes showing > 30% area RSD in QC samples
Target List for evaluation
Analyte | Formula | Adduct | m/z [Da] | RT [min]
Thyroxin | C15H11I4NO4 | -H | 775.6794 | 2.7
Sterole
ST 19:2_OH3_S | C19H28O6S | -H | 383.1534 | 1.9
ST 19:2_OH2_S | C19H28O5S | -H | 367.1585 | 2.5
ST 27:1_OH | C27H46O | -OH | 369.3516 | 5.8
ST 27:1_OH_S | C27H46O4S | -H | 465.3044 | 5
Sterylester
SE 27:1/14:1 | C41H70O2 | +NH4 | 612.5714 | 9.5
SE 27:1/16:1 | C43H74O2 | +NH4 | 640.6027 | 9.9
SE 27:1/17:1 | C44H76O2 | +NH4 | 654.6184 | 10.1
SE 27:1/18:1 | C45H78O2 | +NH4 | 668.634 | 10.2
SE 27:1/18:2 | C45H76O2 | +NH4 | 666.6184 | 9.9
SE 27:1/20:3 | C47H78O2 | +NH4 | 692.634 | 10
SE 27:1/20:4 | C47H76O2 | +NH4 | 690.6184 | 9.7
SE 27:1/20:5 | C47H74O2 | +NH4 | 688.6027 | 9.4
SE 27:1/22:4 | C49H80O2 | +NH4 | 718.6497 | 10
SE 27:1/22:6 | C49H76O2 | +NH4 | 714.6184 | 9.5
SE 27:1/24:1 | C51H90O2 | +NH4 | 752.7279 | 11.1
Ceramide
Cer d18:1/16:0 | C34H67NO3 | +H | 538.5194 | 6.7
Cer d18:1/22:0 | C40H79NO3 | +H | 622.6133 | 8.2
Cer d18:1/23:0 | C41H81NO3 | +H | 636.6289 | 8.4
Cer d18:1/24:0 | C42H83NO3 | +H | 650.6446 | 8.6
Cer d18:0/22:0 | C40H81NO3 | +H | 624.6289 | 8.3
Cer d18:0/24:0 | C42H85NO3 | +H | 652.6602 | 8.8
Cer d18:2/22:0 | C40H77NO3 | +H | 620.5976 | 7.8
Glycosylceramide
HexCer d18:1/16:0 | C40H77NO8 | +H | 700.5722 | 6.3
HexCer d18:1/24:0 | C48H93NO8 | +H | 812.6974 | 8.1
Hex2Cer d18:1/16:0 | C46H87NO13 | +H | 862.625 | 6.1
Hex2Cer d18:1/24:1 | C54H101NO13 | +H | 972.7346 | 7.4
Diglyceride
DG 32:1 | C35H66O5 | +NH4 | 584.5249 | 7.1
DG 34:1 | C37H70O5 | +NH4 | 612.5562 | 7.5
DG 34:2 | C37H68O5 | +NH4 | 610.5405 | 7.1
DG 36:1 | C39H74O5 | +NH4 | 640.5875 | 8
DG 36:2 | C39H72O5 | +NH4 | 638.5718 | 7.6
DG 36:3 | C39H70O5 | +NH4 | 636.5562 | 7.2
DG 36:4 | C39H68O5 | +NH4 | 634.5405 | 6.9
DG 38:5 | C41H70O5 | +NH4 | 660.5562 | 7.2
DG 38:6 | C41H68O5 | +NH4 | 658.5405 | 6.8
freie Fettsäuren
FA 14:0 | C14H28O2 | -H | 227.2017 | 4.5
FA 15:0 | C15H30O2 | -H | 241.2173 | 4.7
FA 16:0 | C16H32O2 | -H | 255.233 | 4.8
FA 16:1 | C16H30O2 | -H | 253.2173 | 4.6
FA 17:0 | C17H34O2 | -H | 269.2486 | 5
FA 17:1 | C17H32O2 | -H | 267.233 | 4.8
FA 18:0 | C18H36O2 | -H | 283.2643 | 5.2
FA 18:1 | C18H34O2 | -H | 281.2486 | 4.9
FA 18:2 | C18H32O2 | -H | 279.233 | 4.7
FA 18:3 | C18H30O2 | -H | 277.2173 | 4.5
FA 18:4 | C18H28O2 | -H | 275.2017 | 4.3
FA 20:1 | C20H38O2 | -H | 309.2799 | 5.3
FA 20:2 | C20H36O2 | -H | 307.2643 | 5
FA 20:4 | C20H32O2 | -H | 303.233 | 4.7
FA 20:5 | C20H30O2 | -H | 301.2173 | 4.5
FA 22:4 | C22H36O2 | -H | 331.2643 | 4.9
FA 22:6 | C22H32O2 | -H | 327.233 | 4.6
FA 24:0 | C24H48O2 | -H | 367.3582 | 6.6
FA 24:1 | C24H46O2 | -H | 365.3425 | 6.1
Lysophosphatidylcholine
LPC 14:0 | C22H46NO7P | +H | 468.3085 | 3.9
LPC 15:0 | C23H48NO7P | +H | 482.3241 | 4.1
LPC 16:0 | C24H50NO7P | +H | 496.3398 | 4.2
LPC 16:1 | C24H48NO7P | +H | 494.3241 | 4
LPC 17:0 | C25H52NO7P | +H | 510.3554 | 4.4
LPC 18:0 | C26H54NO7P | +H | 524.3711 | 4.5
LPC 18:1 | C26H52NO7P | +H | 522.3554 | 4.3
LPC 18:2 | C26H50NO7P | +H | 520.3398 | 4.1
LPC 18:3 | C26H48NO7P | +H | 518.3241 | 3.9
LPC 20:0 | C28H58NO7P | +H | 552.4024 | 4.7
LPC 20:1 | C28H56NO7P | +H | 550.3867 | 4.6
LPC 20:2 | C28H54NO7P | +H | 548.3711 | 4.4
LPC 20:3 | C28H52NO7P | +H | 546.3554 | 4.2
LPC 20:4 | C28H50NO7P | +H | 544.3398 | 4.1
LPC 20:5 | C28H48NO7P | +H | 542.3241 | 3.9
LPC 22:0 | C30H62NO7P | +H | 580.4337 | 5.1
LPC 22:4 | C30H54NO7P | +H | 572.3711 | 4.3
LPC 22:5 | C30H52NO7P | +H | 570.3554 | 4.1
LPC 22:6 | C30H50NO7P | +H | 568.3398 | 4
LPC 24:0 | C32H66NO7P | +H | 608.465 | 5.5
LPC O-16:0 | C24H52NO6P | +HCO2 | 526.3503 | 4.3
LPC O-16:1 | C24H50NO6P | +HCO2 | 524.3347 | 4.3
LPC O-18:0 | C26H56NO6P | +HCO2 | 554.3816 | 4.6
LPC O-18:1 | C26H54NO6P | +HCO2 | 552.366 | 4.4
Lysophosphatidylethanolamine
LPE 16:0 | C21H44NO7P | -H | 452.2783 | 4.2
LPE 18:0 | C23H48NO7P | -H | 480.3096 | 4.5
LPE 18:1 | C23H46NO7P | -H | 478.2939 | 4.3
LPE 18:2 | C23H44NO7P | -H | 476.2783 | 4.1
LPE 20:3 | C25H46NO7P | -H | 502.2939 | 4.2
LPE 20:4 | C25H44NO7P | -H | 500.2783 | 4.1
LPE 22:6 | C27H44NO7P | -H | 524.2783 | 4
LPE O-16:1 | C21H44NO6P | -H | 436.2834 | 4.4
LPE O-18:1 | C23H48NO6P | -H | 464.3147 | 4.7
Lysophosphatidylinositole
LPI 18:0 | C27H53O12P | -H | 599.3202 | 4.5
LPI 18:2 | C27H49O12P | -H | 595.2889 | 4
LPI 20:3 | C29H51O12P | -H | 621.3045 | 4.2
LPI 20:4 | C29H49O12P | -H | 619.2889 | 4
Alkyl- & Alkenyl-phosphatidylcholine
PC O-16:0_16:0 | C40H82NO7P | +HCO2 | 764.58 | 6.7
PC O-16:0_18:2 | C42H82NO7P | +HCO2 | 788.58 | 6.4
PC O-16:0_20:4 | C44H82NO7P | +HCO2 | 812.58 | 6.4
PC O-16:1_16:0 | C40H80NO7P | +HCO2 | 762.5644 | 6.7
PC O-16:1_18:1 | C42H82NO7P | +HCO2 | 788.58 | 6.8
PC O-16:1_18:2 | C42H80NO7P | +HCO2 | 786.5644 | 6.4
PC O-16:1_20:4 | C44H80NO7P | +HCO2 | 810.5644 | 6.3
PC O-18:0_20:4 | C46H86NO7P | +HCO2 | 840.6113 | 6.9
PC O-18:1_20:4 | C46H84NO7P | +HCO2 | 838.5957 | 6.4
PC O-34:1 | C42H84NO7P | +HCO2 | 790.5957 | 6.8
Phosphatidylcholine
PC 30:0 | C38H76NO8P | +H | 706.5381 | 6
PC 30:1 | C38H74NO8P | +H | 704.5225 | 5.8
PC 30:2 | C38H72NO8P | +H | 702.5068 | 5.4
PC 32:0 | C40H80NO8P | +H | 734.5694 | 6.4
PC 32:1 | C40H78NO8P | +H | 732.5538 | 6.1
PC 32:2 | C40H76NO8P | +H | 730.5381 | 5.8
PC 33:2 | C41H78NO8P | +H | 744.5538 | 6
PC 34:0 | C42H84NO8P | +H | 762.6007 | 6.9
PC 34:1 | C42H82NO8P | +H | 760.5851 | 6.5
PC 34:2 | C42H80NO8P | +H | 758.5694 | 6.2
PC 34:3 | C42H78NO8P | +H | 756.5538 | 5.9
PC 34:4 | C42H76NO8P | +H | 754.5381 | 5.8
PC 34:5 | C42H74NO8P | +H | 752.5225 | 5.5
PC 36:1 | C44H86NO8P | +H | 788.6164 | 7
PC 36:2 | C44H84NO8P | +H | 786.6007 | 6.6
PC 36:3 | C44H82NO8P | +H | 784.5851 | 6.3
PC 36:4 | C44H80NO8P | +H | 782.5694 | 6.1
PC 36:5 | C44H78NO8P | +H | 780.5538 | 5.8
PC 38:2 | C46H88NO8P | +H | 814.632 | 7
PC 38:3 | C46H86NO8P | +H | 812.6164 | 6.7
PC 38:4 | C46H84NO8P | +H | 810.6007 | 6.6
PC 38:6 | C46H80NO8P | +H | 806.5694 | 6
PC 40:4 | C48H88NO8P | +H | 838.632 | 6.9
PC 40:6 | C48H84NO8P | +H | 834.6007 | 6.4
PC 40:7 | C48H82NO8P | +H | 832.5851 | 6.1
Phosphatidylethanolamine
PE 32:1 | C37H72NO8P | -H | 688.4923 | 6.2
PE 34:1 | C39H76NO8P | -H | 716.5236 | 6.7
PE 34:2 | C39H74NO8P | -H | 714.5079 | 6.3
PE 34:3 | C39H72NO8P | -H | 712.4923 | 6
PE 36:1 | C41H80NO8P | -H | 744.5549 | 7.1
PE 36:2 | C41H78NO8P | -H | 742.5392 | 6.8
PE 36:3 | C41H76NO8P | -H | 740.5236 | 6.4
PE 36:4 | C41H74NO8P | -H | 738.5079 | 6.3
PE 36:5 | C41H72NO8P | -H | 736.4923 | 6
PE 38:4 | C43H78NO8P | -H | 766.5392 | 6.7
PE 38:5 | C43H76NO8P | -H | 764.5236 | 6.3
PE 38:6 | C43H74NO8P | -H | 762.5079 | 6.1
PE 40:6 | C45H78NO8P | -H | 790.5392 | 6.5
PE 40:7 | C45H76NO8P | -H | 788.5236 | 6.2
PE O-16:1_18:2 | C39H74NO7P | -H | 698.513 | 6.6
PE O-16:1_20:4 | C43H78NO7P | -H | 750.5443 | 6.5
PE O-18:1_18:1 | C41H80NO7P | -H | 728.56 | 7.4
PE O-18:1_18:2 | C41H78NO7P | -H | 726.5443 | 7
PE O-18:1_20:4 | C43H78NO7P | -H | 750.5443 | 7
PE O-18:2_20:4 | C43H78NO7P | -H | 750.5443 | 6.6
Phosphatidylinositole
PI 32:1 | C41H77O13P | -H | 807.5029 | 6.4
PI 34:1 | C43H81O13P | -H | 835.5342 | 6.9
PI 34:2 | C43H79O13P | -H | 833.5186 | 6.5
PI 36:1 | C45H85O13P | -H | 863.5655 | 7.5
PI 36:2 | C45H83O13P | -H | 861.5499 | 7
PI 36:3 | C45H81O13P | -H | 859.5342 | 6.7
PI 36:4 | C45H79O13P | -H | 857.5186 | 6.5
PI 38:3 | C47H85O13P | -H | 887.5655 | 7.2
PI 38:4 | C47H83O13P | -H | 885.5499 | 6.9
PI 38:6 | C47H79O13P | -H | 881.5186 | 6.3
PI 40:5 | C49H85O13P | -H | 911.5655 | 6.9
PI 40:6 | C35H71N2O6P | -H | 645.4977 | 6.8
Phosphatidylserine
PS 36:1 | C42H80NO10P | -H | 788.5447 | 7.3
Sphingomyeline
SM 30:1 | C35H71N2O6P | +H | 647.5123 | 5.4
SM 32:1 | C37H75N2O6P | +H | 675.5436 | 5.7
SM 32:2 | C37H73N2O6P | +H | 673.5279 | 5.4
SM 33:1 | C38H77N2O6P | +H | 689.5592 | 5.9
SM 33:2 | C38H75N2O6P | +H | 687.5436 | 5.6
SM 34:0 | C39H81N2O6P | +H | 705.5905 | 6.3
SM 34:1 | C39H79N2O6P | +H | 703.5749 | 6.1
SM 34:2 | C39H77N2O6P | +H | 701.5592 | 5.8
SM 35:1 | C40H81N2O6P | +H | 717.5905 | 6.3
SM 36:0 | C41H85N2O6P | +H | 733.6218 | 6.7
SM 36:1 | C41H83N2O6P | +H | 731.6062 | 6.5
SM 36:2 | C41H81N2O6P | +H | 729.5905 | 6.2
SM 36:3 | C41H79N2O6P | +H | 727.5749 | 5.9
SM 37:1 | C42H85N2O6P | +H | 745.6218 | 6.8
SM 38:1 | C43H87N2O6P | +H | 759.6375 | 7.1
SM 38:2 | C43H85N2O6P | +H | 757.6218 | 6.6
SM 39:1 | C44H89N2O6P | +H | 773.6531 | 7.3
SM 40:1 | C45H91N2O6P | +H | 787.6688 | 7.5
SM 40:2 | C45H89N2O6P | +H | 785.6531 | 7.1
SM 40:3 | C45H87N2O6P | +H | 783.6375 | 6.7
SM 41:1 | C46H93N2O6P | +H | 801.6844 | 7.7
SM 41:2 | C46H91N2O6P | +H | 799.6688 | 7.4
SM 42:1 | C47H95N2O6P | +H | 815.7001 | 8
SM 42:2 | C47H93N2O6P | +H | 813.6844 | 7.5
SM 42:3 | C47H91N2O6P | +H | 811.6688 | 7.1
SM 43:1 | C48H97N2O6P | +H | 829.7157 | 8.2
SM 43:2 | C48H95N2O6P | +H | 827.7001 | 7.7
Triglyceride
TG 42:1 | C45H84O6 | +NH4 | 738.6606 | 9.1
TG 42:2 | C45H82O6 | +NH4 | 736.645 | 8.7
TG 44:0 | C47H90O6 | +NH4 | 768.7076 | 9.7
TG 44:1 | C47H88O6 | +NH4 | 766.6919 | 9.4
TG 44:2 | C47H86O6 | +NH4 | 764.6763 | 9.1
TG 46:1 | C49H92O6 | +NH4 | 794.7232 | 9.7
TG 46:2 | C49H90O6 | +NH4 | 792.7076 | 9.4
TG 46:3 | C49H88O6 | +NH4 | 790.6919 | 9.1
TG 48:1 | C51H96O6 | +NH4 | 822.7545 | 10
TG 48:2 | C51H94O6 | +NH4 | 820.7389 | 9.7
TG 48:3 | C51H92O6 | +NH4 | 818.7232 | 9.5
TG 48:4 | C51H90O6 | +NH4 | 816.7076 | 9.2
TG 50:1 | C53H100O6 | +NH4 | 850.7858 | 10.3
TG 50:2 | C53H98O6 | +NH4 | 848.7702 | 10
TG 50:3 | C53H96O6 | +NH4 | 846.7545 | 9.8
TG 50:4 | C53H94O6 | +NH4 | 844.7389 | 9.5
TG 52:1 | C55H104O6 | +NH4 | 878.8171 | 10.5
TG 52:2 | C55H104O7 | +NH4 | 894.812 | 10.3
TG 52:3 | C55H100O6 | +NH4 | 874.7858 | 10.1
TG 52:4 | C55H98O6 | +NH4 | 872.7702 | 9.8
TG 52:5 | C55H96O6 | +NH4 | 870.7545 | 9.5
TG 54:1 | C57H108O6 | +NH4 | 906.8484 | 10.8
TG 54:2 | C57H106O6 | +NH4 | 904.8328 | 10.6
TG 54:3 | C57H104O6 | +NH4 | 902.8171 | 10.3
TG 54:4 | C57H102O6 | +NH4 | 900.8015 | 10.1
TG 54:5 | C57H100O6 | +NH4 | 898.7858 | 9.8
TG 54:6 | C57H98O6 | +NH4 | 896.7702 | 9.7
TG 56:1 | C59H112O6 | +NH4 | 934.8797 | 11
TG 56:2 | C59H110O6 | +NH4 | 932.8641 | 10.8
TG 56:3 | C59H108O6 | +NH4 | 930.8484 | 10.6
TG 56:4 | C59H106O6 | +NH4 | 928.8328 | 10.4
TG 56:5 | C59H104O6 | +NH4 | 926.8171 | 10.2
TG 56:6 | C59H102O6 | +NH4 | 924.8015 | 10
TG 58:1 | C61H116O6 | +NH4 | 962.911 | 11.2
TG 58:2 | C61H114O6 | +NH4 | 960.8954 | 11
TG 58:3 | C61H112O6 | +NH4 | 958.8797 | 10.8
TG 58:4 | C61H110O6 | +NH4 | 956.8641 | 10.6
TG 58:6 | C61H106O6 | +NH4 | 952.8328 | 10.2

## Metabolomics
Metabolomic analyses
LC-MS/MS Instrumentation and Analysis
Instrumentation
HPLC: | Thermo Vanquish Horizon
MS: | Thermo Orbitrap Exploris 480
Mode | H-ESI - positive and negative electrospray ionization
Column | SeQuant ZIC-HILIC 2.1 x 100 mm 3.5 µm (Merck)
Pre-column | SeQuant ZIC-HILIC 2.1 x 20 mm Guard Kit (Merck)
Analysis
Data Acquisition | Xcalibur 4.4
Data Evaluation | TraceFinder 5.1
Method | Scan + ddMS2
LC parameters
Solvents
A | Water + 0.1% formic acid
B | Acetonitrile + 0.1% formic acid
Gradient Elution
Time [min] | A [%] | B [%] | Flowrate [mL/min]
0 | 5 | 95 | 0.3
1.5 | 5 | 95 | 0.3
12 | 60 | 40 | 0.3
14 | 60 | 40 | 0.3
14.2 | 75 | 25 | 0.3
17 | 75 | 25 | 0.3
19 | 5 | 95 | 0.5
26 | 5 | 95 | 0.5
Injection Volume |  | Column Temperature
Positive Ionization Mode | 3 µL | 30 °C
Negative Ionization Mode | 5 µL
Positive Ionization |  |  |  |  | Negative Ionization
MS parameters |  |  |  |  | MS parameters
Ion Source |  |  |  |  | Ion Source
Parameters | Value |  |  |  | Parameters | Value
Ion Source Type: | H-ESI |  |  |  | Ion Source Type: | H-ESI
Spray Voltage: | Static |  |  |  | Spray Voltage: | Static
Positive Ion (V): | 3400 |  |  |  | Positive Ion (V): | 3400
Negative Ion (V): | 3000 |  |  |  | Negative Ion (V): | 3000
Gas Mode: | Static |  |  |  | Gas Mode: | Static
Sheath Gas (Arb): | 25 |  |  |  | Sheath Gas (Arb): | 25
Aux Gas (Arb): | 7 |  |  |  | Aux Gas (Arb): | 7
Sweep Gas (Arb): | 0 |  |  |  | Sweep Gas (Arb): | 0
Ion Transfer Tube Temp (°C): | 320 |  |  |  | Ion Transfer Tube Temp (°C): | 320
Vaporizer Temp (°C): | 180 |  |  |  | Vaporizer Temp (°C): | 180
MS Global Settings |  |  |  |  | MS Global Settings
Parameters | Value |  |  |  | Parameters | Value
Infusion Mode: | Liquid Chromatography |  |  |  | Infusion Mode: | Liquid Chromatography
Expected LC Peak Width (s): | 6 |  |  |  | Expected LC Peak Width (s): | 6
Advanced Peak Determination: | TRUE |  |  |  | Advanced Peak Determination: | TRUE
Mild Trapping: | TRUE |  |  |  | Mild Trapping: | TRUE
Default Charge State: | 1 |  |  |  | Default Charge State: | 1
Enable Xcalibur AcquireX method configurations: | TRUE |  |  |  | Enable Xcalibur AcquireX method configurations: | TRUE
Internal Mass Calibration: | RunStart EASY-IC™ |  |  |  | Internal Mass Calibration: | RunStart EASY-IC™
Settings for Full Scan |  |  |  |  | Settings for Full Scan
Orbitrap Resolution: | 120000 |  |  |  | Orbitrap Resolution: | 120000
Scan Range (m/z): | 70-700 |  |  |  | Scan Range (m/z): | 59-590
RF Lens (%): | 20 |  |  |  | RF Lens (%): | 30
AGC Target: | Custom |  |  |  | AGC Target: | Custom
Normalized AGC Target (%): | 100 |  |  |  | Normalized AGC Target (%): | 100
Maximum Injection Time Mode: | Auto |  |  |  | Maximum Injection Time Mode: | Auto
Microscans: | 1 |  |  |  | Microscans: | 1
Data Type: | Profile |  |  |  | Data Type: | Profile
Polarity: | Positive |  |  |  | Polarity: | Negative
Source Fragmentation: | Disabled |  |  |  | Source Fragmentation: | Disabled
Filters |  |  |  |  | Filters
Intensity |  |  |  |  | Intensity
Filter Type: | Intensity Threshold |  |  |  | Filter Type: | Intensity Threshold
Intensity Threshold: | 5000 |  |  |  | Intensity Threshold: | 5000
Dynamic Exclusion |  |  |  |  | Dynamic Exclusion
Dynamic Exclusion Mode: | Custom |  |  |  | Dynamic Exclusion Mode: | Custom
Exclude after n times: | 1 |  |  |  | Exclude after n times: | 1
Exclusion duration (s): | 5 |  |  |  | Exclusion duration (s): | 5
Mass Tolerance: | ppm |  |  |  | Mass Tolerance: | ppm
Low | 10 |  |  |  | Low | 10
High | 10 |  |  |  | High | 10
Exclude isotopes: | TRUE |  |  |  | Exclude isotopes: | TRUE
Perform dependent scan on single charge state per precursor only: | FALSE |  |  |  | Perform dependent scan on single charge state per precursor only: | FALSE
Targeted Mass |  |  |  |  | Targeted Mass
Mass List |  |  |  |  | Mass List
Mass List Type: | m/z |  |  |  | Mass List Type: | m/z
Time Mode: | Start/End Time |  |  |  | Time Mode: | Start/End Time
Include Intensity Threshold: | TRUE |  |  |  | Include Intensity Threshold: | TRUE
Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE |  |  |  | Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE
Compound | m/z | t start (min) | t stop (min) | Intensity Threshold | Compound | m/z | t start (min) | t stop (min) | Intensity Threshold
MRFA (Placeholder) | 524.265 | 0 | 26 | 0 | MRFA (Placeholder) | 524.265 | 0 | 26 | 0
Mass Tolerance: | ppm |  |  |  | Mass Tolerance: | ppm
Low: | 10 |  |  |  | Low: | 10
High: | 10 |  |  |  | High: | 10
Set Collision Energy per Compound: | FALSE |  |  |  | Set Collision Energy per Compound: | FALSE
Perform dependent scan on most intense ion if no targets are found: | TRUE |  |  |  | Perform dependent scan on most intense ion if no targets are found: | TRUE
Targeted Mass Exclusion |  |  |  |  | Targeted Mass Exclusion
Mass List |  |  |  |  | Mass List
Mass List Type: | m/z |  |  |  | Mass List Type: | m/z
Time Mode: | Start/End Time |  |  |  | Time Mode: | Start/End Time
Include Intensity Threshold: | TRUE |  |  |  | Include Intensity Threshold: | TRUE
Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE |  |  |  | Add Mass List Targets Determined by Xcalibur AcquireX: | TRUE
Compound | m/z | t start (min) | t stop (min) | Intensity Threshold | Compound | m/z | t start (min) | t stop (min) | Intensity Threshold
MRFA (Placeholder) | 524.265 | 0 | 26 | 1e+20 | MRFA (Placeholder) | 524.265 | 0 | 26 | 1e+20
Mass Tolerance: | ppm |  |  |  | Mass Tolerance: | ppm
Low: | 10 |  |  |  | Low: | 10
High: | 10 |  |  |  | High: | 10
Apex Detection |  |  |  |  | Apex Detection
Desired Apex Window (%): | 30 |  |  |  | Desired Apex Window (%): | 30
Data Dependent |  |  |  |  | Data Dependent
Data Dependent Mode: | Cycle Time |  |  |  | Data Dependent Mode: | Cycle Time
Time between Master Scans (sec): | 0.6 |  |  |  | Time between Master Scans (sec): | 0.6
Settings for ddMS2 |  |  |  |  | Settings for ddMS2
Multiplex Ions: | FALSE |  |  |  | Multiplex Ions: | FALSE
Isolation Window (m/z): | 1 |  |  |  | Isolation Window (m/z): | 1
Isolation Offset: | Off |  |  |  | Isolation Offset: | Off
Collision Energy Mode: | Stepped |  |  |  | Collision Energy Mode: | Stepped
Collision Energy Type: | Absolute |  |  |  | Collision Energy Type: | Absolute
HCD Collision Energies (V): | 15,30,60 |  |  |  | HCD Collision Energies (V): | 15,30,60
Orbitrap Resolution: | 15000 |  |  |  | Orbitrap Resolution: | 15000
Scan Range Mode: | Auto |  |  |  | Scan Range Mode: | Auto
AGC Target: | Custom |  |  |  | AGC Target: | Custom
Normalized AGC Target (%): | 10 |  |  |  | Normalized AGC Target (%): | 10
Maximum Injection Time Mode: | Auto |  |  |  | Maximum Injection Time Mode: | Auto
Microscans: | 1 |  |  |  | Microscans: | 1
Data Type: | Profile |  |  |  | Data Type: | Profile
Internal Standard
Isotopically labeled internal standard | Vendor | Category | Concentration of standard working solution [ng/mL] | Volume of working solution [µL] | Sample Volume [µL] | Concentration related to sample [ng/mL]
Cholic acid-2,2,4,4-d4 | Cambridge Isotope Laboratories | A | 66.7 | 75 | 10 | 500
D-Fructose-U13C6 | Cambridge Isotope Laboratories | A | 66.7 | 75 | 10 | 500
Glycine-d5 | Sigma Aldrich | A | 66.7 | 75 | 10 | 500
Hypoxanthine-13C5 | Cambridge Isotope Laboratories | B | 6.7 | 75 | 10 | 50
L-Carnitine-d9 | Cambridge Isotope Laboratories | A | 66.7 | 75 | 10 | 500
L-Isoleucine-d10 | Cambridge Isotope Laboratories | A | 66.7 | 75 | 10 | 500
L-Leucine-d10 | Cambridge Isotope Laboratories | A | 66.7 | 75 | 10 | 500
L-Lysine-d9 | Cambridge Isotope Laboratories | A | 66.7 | 75 | 10 | 500
Succinic acid-2,2,3,3-d4 | Sigma-Aldrich | A | 66.7 | 75 | 10 | 500
*Cambridge Isotopes Laboratories | Tewksbury, MA, USA
*Sigma Aldrich | St. Louis, MO, USA
Acceptance criteria
The following analytes were excluded from data evaluation:
- Analytes showing a mass error > 5 ppm
- Analytes showing an isotope score < 40 (10 ppm, 90% intensity)
- Analytes without identity confirmation via spectral library matching or reference substance measurements
- Analytes containing only missing values in QC samples and study samples
- Analytes containing > 50% missing values in QC samples and study samples
- Analytes showing > 30% area RSD in QC samples
Target List for evaluation
Compound Name | Compound Type | Compound Formula | m/z | Polarity | Adduct | Charge State | Retention Time
5-Hydroxylysine | TargetCompound | C6H14N2O3 | 163.1077 | Positive | M+H | 1 | 10.7
Acetyllysine | TargetCompound | C8H16N2O3 | 189.1234 | Positive | M+H | 1 | 7.73
Adenosine | TargetCompound | C10H13N5O4 | 268.104 | Positive | M+H | 1 | 5.63
Aminobutyric acid | TargetCompound | C4H9NO2 | 104.0706 | Positive | M+H | 1 | 7.73
Aminooctanoic acid | TargetCompound | C8H17NO2 | 160.1332 | Positive | M+H | 1 | 3.21
Anserine | TargetCompound | C10H16N4O3 | 241.1295 | Positive | M+H | 1 | 12.48
Arginine | TargetCompound | C6H14N4O2 | 175.119 | Positive | M+H | 1 | 11.37
Asparagine | TargetCompound | C4H8N2O3 | 133.0608 | Positive | M+H | 1 | 8.52
Aspartic acid | TargetCompound | C4H7NO4 | 134.0448 | Positive | M+H | 1 | 8.32
Betaine | TargetCompound | C5H11NO2 | 118.0863 | Positive | M+H | 1 | 7.52
CAR 2:0 | TargetCompound | C9H17NO4 | 204.123 | Positive | M+H | 1 | 2.78
CAR 3:0 | TargetCompound | C10H19NO4 | 218.1387 | Positive | M+H | 1 | 2.7
CAR 4:0 | TargetCompound | C11H21NO4 | 232.1543 | Positive | M+H | 1 | 1.93
CAR 5:0 | TargetCompound | C12H23NO4 | 246.17 | Positive | M+H | 1 | 1.85
CAR 6:0 | TargetCompound | C13H25NO4 | 260.1856 | Positive | M+H | 1 | 1.79
CAR 8:0 | TargetCompound | C15H29NO4 | 288.2169 | Positive | M+H | 1 | 1.71
CAR DC3:0;2Me | TargetCompound | C11H19NO6 | 262.1285 | Positive | M+H | 1 | 6.03
CAR DC5:0 | TargetCompound | C12H21NO6 | 276.1442 | Positive | M+H | 1 | 6.82
CAR DC5:0;3Me | TargetCompound | C13H23NO6 | 290.1598 | Positive | M+H | 1 | 6.22
Carnitine | TargetCompound | C7H15NO3 | 162.1125 | Positive | M+H | 1 | 7.68
Carnitine-d9 | InternalStandard | C7H6NO3D9 | 171.169 | Positive | M+H | 1 | 7.1
Carnosine | TargetCompound | C9H14N4O3 | 227.1139 | Positive | M+H | 1 | 11.8
Choline | TargetCompound | C5H14NO | 104.107 | Positive | M+ | 1 | 4.4
cis-3-Hydroxyproline | TargetCompound | C5H9NO3 | 132.0655 | Positive | M+H | 1 | 8.02
Citrulline | TargetCompound | C6H13N3O3 | 176.103 | Positive | M+H | 1 | 8.66
Creatine | TargetCompound | C4H9N3O2 | 132.0768 | Positive | M+H | 1 | 7.4
Creatinine | TargetCompound | C4H7N3O | 114.0662 | Positive | M+H | 1 | 6.41
Cystathionine | TargetCompound | C7H14N2O4S | 223.0747 | Positive | M+H | 1 | 9.95
Cysteine | TargetCompound | C3H7NO2S | 122.027 | Positive | M+H | 1 | 7.68
Cysteine-S-sulfate | TargetCompound | C3H7NO5S2 | 201.9838 | Positive | M+H | 1 | 7.25
Cystine | TargetCompound | C6H12N2O4S2 | 241.0311 | Positive | M+H | 1 | 10.04
Cytidine | TargetCompound | C9H13N3O5 | 244.0928 | Positive | M+H | 1 | 7.89
Dimethylarginine | TargetCompound | C8H18N4O2 | 203.1503 | Positive | M+H | 1 | 9.51
Dimethylglycine | TargetCompound | C4H9NO2 | 104.0706 | Positive | M+H | 1 | 6.89
DOPA | TargetCompound | C9H11NO4 | 198.0761 | Positive | M+H | 1 | 7.67
Dopamine | TargetCompound | C8H11NO2 | 154.0863 | Positive | M+H | 1 | 7.4
Ergothioneine | TargetCompound | C9H15N3O2S | 230.0958 | Positive | M+H | 1 | 8.83
Glutamic acid | TargetCompound | C5H9NO4 | 148.0604 | Positive | M+H | 1 | 7.94
Glutamic acid-d5 | InternalStandard | C5H4NO4D5 | 153.0918 | Positive | M+H | 1 | 7.94
Glutamine | TargetCompound | C5H10N2O3 | 147.0764 | Positive | M+H | 1 | 8.38
Glycine | TargetCompound | C2H5NO2 | 76.0393 | Positive | M+H | 1 | 8.44
Glycine-d2 | InternalStandard | C2H3NO2D2 | 78.0519 | Positive | M+H | 1 | 8.44
Guanidinobutanoic acid | TargetCompound | C5H11N3O2 | 146.0924 | Positive | M+H | 1 | 6.32
Guanidinosuccinic acid | TargetCompound | C5H9N3O4 | 176.0666 | Positive | M+H | 1 | 6.42
Guanine | TargetCompound | C5H5N5O | 152.0567 | Positive | M+H | 1 | 6.9
Histamine | TargetCompound | C5H9N3 | 112.0869 | Positive | M+H | 1 | 13.06
Histidine | TargetCompound | C6H9N3O2 | 156.0768 | Positive | M+H | 1 | 10.04
Homoarginine | TargetCompound | C7H16N4O2 | 189.1346 | Positive | M+H | 1 | 12.1
Homocysteine | TargetCompound | C4H9NO2S | 136.0427 | Positive | M+H | 1 | 7.09
Hydroxybenzaldehyde | TargetCompound | C7H6O2 | 123.0441 | Positive | M+H | 1 | 6.97
Hypotaurine | TargetCompound | C2H7NO2S | 110.027 | Positive | M+H | 1 | 8.42
Hypoxanthine | TargetCompound | C5H4N4O | 137.0458 | Positive | M+H | 1 | 5.09
Hypoxanthine-13C5 | InternalStandard | [13]C5H4N4O | 142.0626 | Positive | M+H | 1 | 5.09
Indoleacrylic acid | TargetCompound | C11H9NO2 | 188.0706 | Positive | M+H | 1 | 6.16
Indolepropionic acid | TargetCompound | C11H11NO2 | 190.0863 | Positive | M+H | 1 | 5.81
Inosine | TargetCompound | C10H12N4O5 | 269.0881 | Positive | M+H | 1 | 5.84
Kynurenic acid | TargetCompound | C10H7NO3 | 190.0499 | Positive | M+H | 1 | 4.89
Kynurenine | TargetCompound | C10H12N2O3 | 209.0921 | Positive | M+H | 1 | 6.01
Leucine | TargetCompound | C6H13NO2 | 132.1019 | Positive | M+H | 1 | 5.84
Leucine-d10 | InternalStandard | C6H3NO2D10 | 142.1647 | Positive | M+H | 1 | 5.84
Leu-Pro | TargetCompound | C11H20N2O3 | 229.1547 | Positive | M+H | 1 | 4
Lysine | TargetCompound | C6H14N2O2 | 147.1128 | Positive | M+H | 1 | 10.37
Lysine-d9 | InternalStandard | C6H5N2O2D9 | 156.1693 | Positive | M+H | 1 | 11.63
Methionine | TargetCompound | C5H11NO2S | 150.0583 | Positive | M+H | 1 | 6.51
Methionine sulfoxide | TargetCompound | C5H11NO3S | 166.0532 | Positive | M+H | 1 | 8.51
Methyladenosine | TargetCompound | C11 H15 N5 O4 | 282.1197 | Positive | M+H | 1 | 7.12
Methylhistamine | TargetCompound | C6H11N3 | 126.1026 | Positive | M+H | 1 | 15.62
Methylhistidine | TargetCompound | C7H11N3O2 | 170.0924 | Positive | M+H | 1 | 10.1
Methylimidazoleacetic acid | TargetCompound | C6H8N2O2 | 141.0659 | Positive | M+H | 1 | 6.58
Methylxanthine | TargetCompound | C6H6N4O2 | 167.0564 | Positive | M+H | 1 | 2.59
N4-Acetylcytidine | TargetCompound | C11 H15 N3 O6 | 286.1034 | Positive | M+H | 1 | 6.31
N-6-Trimethyllysine | TargetCompound | C9H20N2O2 | 189.1598 | Positive | M+H | 1 | 10.65
N-Acetylarginine | TargetCompound | C8H16N4O3 | 217.1295 | Positive | M+H | 1 | 7.25
N-Acetylcarnosine | TargetCompound | C11H16N4O4 | <multiple> | Positive | M+H | 1 | 8
N-Acetylornithine | TargetCompound | C7H14N2O3 | 175.1077 | Positive | M+H | 1 | 8.02
Niacinamide | TargetCompound | C6H6N2O | 123.0553 | Positive | M+H | 1 | 2.175
O-Acetylserine | TargetCompound | C5H9NO4 | 148.0604 | Positive | M+H | 1 | 7.09
Ornithine | TargetCompound | C5H12N2O2 | 133.0972 | Positive | M+H | 1 | 10.52
Pantothenic acid | TargetCompound | C9H17NO5 | 220.118 | Positive | M+H | 1 | 2.5
Paraxanthine / Theobromine / Theophylline | TargetCompound | C7H8N4O2 | 181.072 | Positive | M+H | 1 | 1.66
Phenylalanine | TargetCompound | C9H11NO2 | 166.0863 | Positive | M+H | 1 | 5.8
Phenylethylamine | TargetCompound | C8H11N | 122.0964 | Positive | M+H | 1 | 1.98
Phenylglycine | TargetCompound | C8H9NO2 | 152.0706 | Positive | M+H | 1 | 4.6
Pipecolic acid | TargetCompound | C6H11NO2 | 130.0863 | Positive | M+H | 1 | 10.37
Proline | TargetCompound | C5H9NO2 | 116.0706 | Positive | M+H | 1 | 7.55
Proline betaine | TargetCompound | C7H13NO2 | 144.1019 | Positive | M+H | 1 | 7.62
Pyroglutamic acid | TargetCompound | C5H7NO3 | 130.0499 | Positive | M+H | 1 | 8.175
Sarcosine | TargetCompound | C3H7NO2 | 90.055 | Positive | M+H | 1 | 8
Serine | TargetCompound | C3H7NO3 | 106.0499 | Positive | M+H | 1 | 8.52
Serotonin | TargetCompound | C10H12N2O | 177.1022 | Positive | M+H | 1 | 6.57
Taurine | TargetCompound | C2H7NO3S | 126.0219 | Positive | M+H | 1 | 8.05
Threonine | TargetCompound | C4H9NO3 | 120.0655 | Positive | M+H | 1 | 8.04
Trigonelline | TargetCompound | C7H7NO2 | 138.055 | Positive | M+H | 1 | 7.44
Trimethylamine N-oxide | TargetCompound | C3H9NO | 76.0757 | Positive | M+H | 1 | 5.8
Tryptophan | TargetCompound | C11H12N2O2 | 205.0972 | Positive | M+H | 1 | 6.16
Tyrosine | TargetCompound | C9H11NO3 | 182.0812 | Positive | M+H | 1 | 6.97
Uracil | TargetCompound | C4H4N2O2 | 113.0346 | Positive | M+H | 1 | 5.09
Urocanic acid | TargetCompound | C6H6N2O2 | 139.0502 | Positive | M+H | 1 | 5.69
Valine | TargetCompound | C5H11NO2 | 118.0863 | Positive | M+H | 1 | 6.68
Xanthine | TargetCompound | C5H4N4O2 | 153.0407 | Positive | M+H | 1 | 5.03
Acetamidobutanoic acid | TargetCompound | C6H11NO3 | 144.0666 | Negative | M-H | 1 | 1.5
Allantoin | TargetCompound | C4H6N4O3 | 157.0367 | Negative | M-H | 1 | 6.67
alpha-N-Phenylacetylglutamine | TargetCompound | C13H16N2O4 | 263.1037 | Negative | M-H | 1 | 1.61
Aminohippuric acid | TargetCompound | C9H10N2O3 | 193.0619 | Negative | M-H | 1 | 1.92
AMP | TargetCompound | C10H14N5O7P | 346.0558 | Negative | M-H | 1 | 9
Ascorbic acid sulfate | TargetCompound | C6H8O9S | 254.9816 | Negative | M-H | 1 | 6.51
BA 24:1;O3;G | TargetCompound | C26H43NO4 | 432.3119 | Negative | M-H | 1 | 1.27
BA 24:1;O3;G;S | TargetCompound | C26H43NO7S | 512.2688 | Negative | M-H | 1 | 4.7
BA 24:1;O3;T | TargetCompound | C26H45NO5S | 482.2946 | Negative | M-H | 1 | 4.62
BA 24:1;O4 | TargetCompound | C24H40O4 | 391.2854 | Negative | M-H | 1 | 1
BA 24:1;O4;G | TargetCompound | C26H43NO5 | 448.3069 | Negative | M-H | 1 | 1.57
BA 24:1;O4;T | TargetCompound | C26H45NO6S | 498.2895 | Negative | M-H | 1 | 4.99
BA 24:1;O5 | TargetCompound | C24H40O5 | 407.2803 | Negative | M-H | 1 | 1.2
BA 24:1;O5;G | TargetCompound | C26H43NO6 | 464.3018 | Negative | M-H | 1 | 3.77
BA 24:1;O5;T | TargetCompound | C26H45NO7S | 514.2844 | Negative | M-H | 1 | 5.46
Bilirubin | TargetCompound | C33H36N4O6 | 583.2562 | Negative | M-H | 1 | 2.86
cis-Aconitic acid | TargetCompound | C6H6O6 | 173.0092 | Negative | M-H | 1 | 5.5
Citric acid | TargetCompound | C6H8O7 | 191.0197 | Negative | M-H | 1 | 7.8
Citric acid-13C3 | InternalStandard | [13]C3C3H8O7 | 194.0298 | Negative | M-H | 1 | 4.71
Cysteic acid | TargetCompound | C3H7NO5S | 167.9972 | Negative | M-H | 1 | 8.25
Dihexose | TargetCompound | C12H22O11 | 387.1144 | Negative | M+HCO2 | 1 | 8.45
Dihydroxybenzeneacetic acid | TargetCompound | C8H8O4 | 167.035 | Negative | M-H | 1 | 1.22
Dihydroxybenzoic acid | TargetCompound | C7H6O4 | 153.0193 | Negative | M-H | 1 | 1.66
Dimethyluric acid | TargetCompound | C7H8N4O3 | 195.0524 | Negative | M-H | 1 | 2.43
Ethylmalonic acid / Methylsuccinic acid / Glutaric acid | TargetCompound | C5H8O4 | 131.035 | Negative | M-H | 1 | 1.58
Fructose-13C6 | InternalStandard | [13]C6H12O6 | 221.0529 | Negative | M+Cl | 1 | 6.77
Fumaric acid | TargetCompound | C4H4O4 | 115.0037 | Negative | M-H | 1 | 1.6
Galacturonic acid | TargetCompound | C5H10O5 | 149.0456 | Negative | M-H | 1 | 5.51
Gluconic acid | TargetCompound | C6H12O7 | 195.051 | Negative | M-H | 1 | 8.49
Glucuronic acid | TargetCompound | C6H10O7 | 193.0354 | Negative | M-H | 1 | 7.47
Glucuronolactone | TargetCompound | C6H8O6 | 175.0248 | Negative | M-H | 1 | 4.86
Glutamic acid-d5 | InternalStandard | C5H4NO4D5 | 151.0773 | Negative | M-H | 1 | 8.06
Glutathione | TargetCompound | C10H17N3O6S | 306.0765 | Negative | M-H | 1 | 7.83
Glycerophospho-N-palmitoyl ethanolamine | TargetCompound | C21H44NO7P | 452.2783 | Negative | M-H | 1 | 5.54
Gly-Leu | TargetCompound | C8H16N2O3 | 187.1088 | Negative | M-H | 1 | 6.86
GMP | TargetCompound | C10H14N5O8P | 362.0507 | Negative | M-H | 1 | 9.47
Guanidoacetic acid | TargetCompound | C3H7N3O2 | 116.0466 | Negative | M-H | 1 | 7.86
Guanosine | TargetCompound | C10H13N5O5 | 282.0844 | Negative | M-H | 1 | 6.5
Hexose | TargetCompound | C6H12O6 | 215.0328 | Negative | M+Cl | 1 | 7.57
Hexose 6-phosphate | TargetCompound | C6H13O9P | 259.0224 | Negative | M-H | 1 | 9.87
Hippuric acid | TargetCompound | C9H9NO3 | 178.051 | Negative | M-H | 1 | 1.47
Hydroxymethylglutaric acid | TargetCompound | C6H10O5 | 161.0456 | Negative | M-H | 1 | 1.66
Hydroxyphenyllactic acid | TargetCompound | C9H10O4 | 181.0506 | Negative | M-H | 1 | 1.66
Hypoxanthine-13C5 | InternalStandard | [13]C5H4N4O | 140.048 | Negative | M-H | 1 | 5.2
IMP | TargetCompound | C10H13N4O8P | 347.0398 | Negative | M-H | 1 | 8.77
Indolelactic acid | TargetCompound | C11H11NO3 | 204.0666 | Negative | M-H | 1 | 2.275
Indoxyl sulfate | TargetCompound | C8H7NO4S | 212.0023 | Negative | M-H | 1 | 5.28
Lactic acid | TargetCompound | C3H6O3 | 89.0244 | Negative | M-H | 1 | 3.25
Maleic acid | TargetCompound | C4H4O4 | 115.0037 | Negative | M-H | 1 | 2.4
Malic acid | TargetCompound | C4H6O5 | 133.0143 | Negative | M-H | 1 | 6.1
Malonic acid | TargetCompound | C3H4O4 | 103.0037 | Negative | M-H | 1 | 4.5
Mandelic acid / Hydroxyphenylacetic acid | TargetCompound | C8H8O3 | 151.0401 | Negative | M-H | 1 | 1.375
Mannitol / Iditol | TargetCompound | C6H14O6 | 217.0484 | Negative | M+Cl | 1 | 7.36
Mesaconic acid / citraconic acid / itaconic acid / glutaconic acid | TargetCompound | C5H6O4 | 129.0193 | Negative | M-H | 1 | 1.61
Methoxyhydroxyphenylethyleneglycol sulfate | TargetCompound | C9H12O7S | 263.0231 | Negative | M-H | 1 | 5.09
Methylpentose | TargetCompound | C6H12O5 | 209.0667 | Negative | M+HCO2 | 1 | 5.87
Methyluric acid | TargetCompound | C6H6N4O3 | 181.0367 | Negative | M-H | 1 | 5.25
N-Acetylalanine | TargetCompound | C5H9NO3 | 130.051 | Negative | M-H | 1 | 8.01
N-Acetylaspartic acid | TargetCompound | C6H9NO5 | 174.0408 | Negative | M-H | 1 | 4.64
N-Acetyl-Asp-Glu | TargetCompound | C11H16N2O8 | 303.0834 | Negative | M-H | 1 | 5.5
N-Acetyl-D-galactosamine 4-sulphate | TargetCompound | C8H15NO9S | 300.0395 | Negative | M-H | 1 | 7.26
N-Acetylglutamic acid | TargetCompound | C7H11NO5 | 188.0565 | Negative | M-H | 1 | 4.73
N-Acetylglutamine | TargetCompound | C7H12N2O4 | 187.0724 | Negative | M-H | 1 | 4.93
N-Acetylneuraminic acid | TargetCompound | C11H19NO9 | 308.0987 | Negative | M-H | 1 | 7.46
N-Acetyl-tryptophan | TargetCompound | C13H14N2O3 | 245.0932 | Negative | M-H | 1 | 1.55
Orotic acid | TargetCompound | C5H4N2O4 | 155.0098 | Negative | M-H | 1 | 6
Orotidine | TargetCompound | C10H12N2O8 | 287.0521 | Negative | M-H | 1 | 7.72
Oxoproline | TargetCompound | C5H7NO3 | 128.0353 | Negative | M-H | 1 | 2.6
P-Cresol | TargetCompound | C7H8O | 107.0502 | Negative | M-H | 1 | 1.62
p-Cresol sulfate | TargetCompound | C7H8O4S | 187.0071 | Negative | M-H | 1 | 1.62
Pentose | TargetCompound | C5H10O5 | 149.0456 | Negative | M-H | 1 | 5.51
Pseudouridine | TargetCompound | C9H12N2O6 | 243.0623 | Negative | M-H | 1 | 5.09
Quinic acid | TargetCompound | C7H12O6 | 191.0561 | Negative | M-H | 1 | 5.17
Ribose 5-phosphate | TargetCompound | C5H11O8P | 229.0119 | Negative | M-H | 1 | 10.5
Saccharopine | TargetCompound | C11H20N2O6 | 275.1249 | Negative | M-H | 1 | 10.03
S-adenosylhomocysteine | TargetCompound | C14H20N6O5S | 383.1143 | Negative | M-H | 1 | 9.42
Salicyluric acid | TargetCompound | C9H9NO4 | 194.0459 | Negative | M-H | 1 | 4.15
ST 19:2;O2;S | TargetCompound | C19H28O5S | 367.1585 | Negative | M-H | 1 | 1.64
Succinic acid | TargetCompound | C4H6O4 | 117.0193 | Negative | M-H | 1 | 1.66
Succinic acid-d4 | InternalStandard | C4H2O4D4 | 121.0444 | Negative | M-H | 1 | 1.62
Threonic acid | TargetCompound | C4H8O5 | 135.0299 | Negative | M-H | 1 | 7.5
Trimethyluric acid | TargetCompound | C8H10N4O3 | 209.068 | Negative | M-H | 1 | 1.85
UMP | TargetCompound | C9H13N2O9P | 323.0286 | Negative | M-H | 1 | 9.96
Ureidopropionic acid | TargetCompound | C4H8N2O3 | 131.0462 | Negative | M-H | 1 | 4.66
Uric acid | TargetCompound | C5H4N4O3 | 167.0211 | Negative | M-H | 1 | 6.56
Uridine | TargetCompound | C9H12N2O6 | 243.0623 | Negative | M-H | 1 | 5.1
Vitamin C | TargetCompound | C6H8O6 | 175.0248 | Negative | M-H | 1 | 4.8
Xanthosine | TargetCompound | C10H12N4O6 | 283.0684 | Negative | M-H | 1 | 6.13
Xanthurenic acid | TargetCompound | C10H7NO4 | 204.0302 | Negative | M-H | 1 | 5.37
Xylitol / Arabitol | TargetCompound | C5H12O5 | 151.0612 | Negative | M-H | 1 | 6.58

## Sphingolipids
Targeted sphingolipid analyses
LC-MS/MS Instrumentation and Analysis
Instrumentation
HPLC: | Agilent 1290 Infinity II
MS: | Sciex QTRAP 6500+
Mode | ESI - positive electrospray ionization
Column | UHPLC Zorbax Eclipse Plus C8 2.1 x 30 mm 1.8 µm (Agilent)
Pre-column | UHPLC guard C8 pre-column 2.1 x 3 mm 1.8 µm (Agilent)
Column Cermides | UHPLC Zorbax Eclipse Plus C18 2.1 x 50 mm 1.8 µm column (Agilent)
Pre-column Ceramides | UHPLC guard C18 pre-column 2.1 x 3 mm 1.8 µm (Agilent)
Sample preparation | LLE (Liquid Liquid Extraction)
Analysis
Application | Analyst 1.7.1 - Multi Quant 3.0.3 (Sciex)
Method | Internal standard method (isotope diution mass spectrometry)
Calibration curve | linear regression
weighting | 1/x
LC parameters
Solvents
A | Water + 0,5% formic acid
B | Acetonitrile/Isopropanol/Acetone (50:30:20) + 1% formic acid
Gradient Elution
Time [min] | A [%] | B [%] | Flowrate [mL/min]
0 | 55 | 45 | 0.4
0.5 | 55 | 45 | 0.4
1.5 | 0 | 100 | 0.4
3 | 0 | 100 | 0.4
3.1 | 55 | 45 | 0.4
4.5 | 55 | 45 | 0.4
Sphingolipids (Ceramides)
Solvents
A | Water + 10 mM Ammoniumformiate + 0,2% formic acid
B | Acetonitrile/Isopropanol/Acetone (50:30:20) + 0.2% formic acid
Gradient Elution
Time [min] | A [%] | B [%] | Flowrate [mL/min]
0 | 50 | 50 | 0.35
0.2 | 50 | 50 | 0.35
0.6 | 10 | 90 | 0.35
4 | 0 | 100 | 0.35
5.5 | 0 | 100 | 0.35
6 | 50 | 50 | 0.35
7.5 | 50 | 50 | 0.35
MS parameters
Parameters | Value
Temperature (TEM) | 500 °C
Curtain Gas (CUR) | 40 psi
Collision Gas (CAD) | Medium
IonSpray Voltage (ISV) | 5000 V
Ion Source Gas 1 (GS1) | 50 psi
Ion Source Gas 2 (GS2) | 70 psi
Settings for multiple reaction monitoring (MRM)
Analyte | Q1 | Q3 | Dwell time | DP | EP | CXP | CE
| [m/z] | [m/z] | [ms] | [V] | [V] | [V] | [V]
SPH d18:1 | 300.282 | 252.3 | 10 | 40 | 10 | 14 | 25
SPH d18:1 | 300.282 | 282.3 | 10 | 40 | 10 | 22 | 17
SPH d18:0 | 302.298 | 284.3 | 10 | 57.5 | 10 | 10 | 19
SPH d18:0 | 302.298 | 254.2 | 10 | 57.5 | 10 | 22 | 29
S1P d18:1 | 380.249 | 264.2 | 10 | 55 | 10 | 10 | 23
S1P d18:1 | 380.249 | 362.3 | 10 | 55 | 10 | 16 | 15
S1P d18:0 | 382.264 | 284.1 | 10 | 40 | 10 | 12 | 19
S1P d18:0 | 382.246 | 266.3 | 10 | 40 | 10 | 20 | 25
SPH d18:1-d7 | 307.282 | 289.3 | 10 | 40 | 10 | 22 | 17
SPH d18:1-d7 | 307.282 | 259.3 | 10 | 40 | 10 | 14 | 25
SPH d18:0-d7 | 309.298 | 291.3 | 10 | 57.5 | 10 | 10 | 19
SPH d18:0-d7 | 309.298 | 261.2 | 10 | 57.5 | 10 | 22 | 29
S1P d18:1-d7 | 387.249 | 271.3 | 10 | 55 | 10 | 10 | 23
S1P d18:1-d7 | 387.249 | 369.3 | 10 | 55 | 10 | 16 | 15
S1P d20:1 | 408.3 | 292.4 | 10 | 50 | 10 | 20 | 22
S1P d20:1 | 408.3 | 82.3 | 10 | 50 | 10 | 7 | 48
SPH d20:0 | 330.3 | 312.4 | 10 | 57.5 | 10 | 10 | 19
SPH d20:0 | 330.3 | 60 | 10 | 57.5 | 10 | 22 | 37
SPH d20:1 | 328.31 | 310.1 | 10 | 30 | 10 | 20 | 17
SPH d20:1 | 328.31 | 280.1 | 10 | 30 | 10 | 18 | 27
S1P d18:0-d7 | 389.32 | 291.2 | 10 | 50.5 | 10 | 20 | 17
S1P d18:0-d7 | 389.32 | 273.1 | 10 | 50.5 | 10 | 18 | 25
S1P d16:1 | 352.218 | 236.2 | 10 | 50 | 10 | 18 | 21
S1P d16:1 | 352.218 | 334.2 | 10 | 50 | 10 | 28 | 14
S1P-d18:2 | 378.191 | 262.2 | 10 | 56 | 10 | 14 | 21
S1P-d18:2 | 378.191 | 360.1 | 10 | 56 | 10 | 18 | 15
S1P-d18:2 | 378.191 | 82.1 | 10 | 56 | 10 | 10 | 39
Sphingolipids (Ceramides)
Parameters | Value
S1P-d18:2 3 | 350 °C
Curtain Gas (CUR) | 35 psi
Collision Gas (CAD) | Medium
IonSpray Voltage (ISV) | 5000 V
Ion Source Gas 1 (GS1) | 50 psi
Ion Source Gas 2 (GS2) | 50 psi
Settings for multiple reaction monitoring (MRM)
Analyte | Q1 | Q3 | Dwell time | DP | EP | CXP | CE
| [m/z] | [m/z] | [ms] | [V] | [V] | [V] | [V]
Cer d18:0/16:0 | 540.528 | 284.3 | 3.03 | 48 | 10 | 14 | 37
Cer d18:0/16:0 | 540.528 | 522.5 | 3.03 | 48 | 10 | 22 | 25
Cer d18:0/18:0 | 568.559 | 284.3 | 3.4 | 38 | 10 | 10 | 41
Cer d18:0/18:0 | 568.559 | 550.6 | 3.4 | 38 | 10 | 22 | 29
Cer d18:0/18:0-d3 | 571.631 | 284.2 | 3.4 | 40 | 10 | 10 | 40
Cer d18:0/18:0-d3 | 571.631 | 553.4 | 3.4 | 40 | 10 | 22 | 29
Cer d18:0/18:1 | 566.6 | 284.3 | 3.06 | 32.5 | 10 | 25 | 41
Cer d18:0/18:1 | 566.6 | 548.6 | 3.06 | 32.5 | 10 | 26 | 26
Cer d18:0/24:0 | 652.653 | 284.2 | 4.54 | 40.5 | 10 | 20 | 41
Cer d18:0/24:0 | 652.653 | 634.2 | 4.54 | 40.5 | 10 | 28 | 29
Cer d18:0/24:0-d7 | 659.71 | 273.1 | 4.54 | 28 | 10 | 8 | 49
Cer d18:0/24:0-d7 | 659.71 | 641.3 | 4.54 | 28 | 10 | 22 | 27
Cer d18:0/24:1 | 650.637 | 284.2 | 4.13 | 45.5 | 10 | 20 | 43
Cer d18:0/24:1 | 650.637 | 632.6 | 4.13 | 45.5 | 10 | 26 | 31
Cer d18:1/14:0 | 510.481 | 264.2 | 2.56 | 38 | 10 | 16 | 33
Cer d18:1/14:0 | 510.481 | 492.5 | 2.56 | 38 | 10 | 20 | 13
Cer d18:1/16:0 | 538.512 | 264.2 | 2.88 | 50 | 10 | 18 | 40
Cer d18:1/16:0 | 538.512 | 520.2 | 2.88 | 50 | 10 | 20 | 17
Cer d18:1/16:0-d7 | 545.5 | 271.2 | 2.88 | 25.5 | 10 | 15 | 39
Cer d18:1/16:0-d7 | 545.5 | 527.5 | 2.88 | 25.5 | 10 | 14 | 20
Cer d18:1/18:0 | 566.54 | 264.4 | 3.24 | 33 | 10 | 12 | 35
Cer d18:1/18:0 | 566.54 | 548.5 | 3.24 | 33 | 10 | 20 | 20
Cer d18:1/18:0-d7 | 573.598 | 271.1 | 3.24 | 33 | 10 | 20 | 35
Cer d18:1/18:0-d7 | 573.598 | 555.1 | 3.24 | 33 | 10 | 26 | 15
Cer d18:1/18:1 | 564.528 | 264.3 | 2.92 | 30 | 10 | 14 | 37
Cer d18:1/18:1 | 564.528 | 546.4 | 2.92 | 30 | 10 | 16 | 17
Cer d18:1/18:1-d7 | 571.58 | 271.3 | 2.92 | 20 | 10 | 18 | 35
Cer d18:1/18:1-d7 | 571.58 | 553.5 | 2.92 | 20 | 10 | 42 | 17
Cer d18:1/20:0 | 594.575 | 264.4 | 3.62 | 35.5 | 10 | 20 | 37
Cer d18:1/20:0 | 594.575 | 576.5 | 3.62 | 35.5 | 10 | 26 | 17
Cer d18:1/20:1-d7 | 599.6 | 271.2 | 3.25 | 40 | 10 | 18 | 35
Cer d18:1/20:1-d7 | 599.6 | 581.6 | 3.25 | 40 | 10 | 42 | 17
Cer d18:1/22:0 | 622.6 | 264.4 | 4 | 35 | 10 | 25 | 40
Cer d18:1/22:0 | 622.6 | 604.6 | 4 | 35 | 10 | 35 | 23
Cer d18:1/22:1-d7 | 627.65 | 271.3 | 3.64 | 16 | 10 | 20 | 38
Cer d18:1/22:1-d7 | 627.65 | 609.6 | 3.64 | 16 | 10 | 16 | 17
Cer d18:1/24:0 13C Isotope | 651.637 | 264.2 | 4.37 | 65.5 | 10 | 28 | 41
Cer d18:1/24:0 13C Isotope | 651.637 | 633.6 | 4.37 | 65.5 | 10 | 24 | 19
Cer d18:1/24:0-d4 | 654.669 | 264.4 | 4.37 | 65 | 10 | 28 | 41
Cer d18:1/24:0-d4 | 654.669 | 636.6 | 4.37 | 65 | 10 | 24 | 19
Cer d18:1/24:1 13C Isotope | 649.622 | 264.2 | 3.97 | 40.5 | 10 | 28 | 41
Cer d18:1/24:1 13C Isotope | 649.622 | 631.6 | 3.97 | 40.5 | 10 | 26 | 19
Cer d18:1/24:1-d7 | 655.6 | 271.3 | 3.97 | 40.5 | 10 | 15 | 46
Cer d18:1/24:1-d7 | 655.6 | 637.6 | 3.97 | 40.5 | 10 | 17 | 23
GlcCer d18:1/16:0 13C Isotope | 701.57 | 264.2 | 2.46 | 50.5 | 10 | 18 | 45
GlcCer d18:1/16:0 13C Isotope | 701.57 | 683.5 | 2.46 | 50.5 | 10 | 30 | 19
GlcCer d18:1/18:0 | 728.581 | 264.2 | 2.78 | 43 | 10 | 24 | 49
GlcCer d18:1/18:0 | 728.581 | 710.6 | 2.78 | 43 | 10 | 30 | 21
GlcCer d18:1/18:0-d5 | 733.676 | 269.3 | 2.78 | 43 | 10 | 24 | 49
GlcCer d18:1/18:0-d5 | 733.676 | 715.6 | 2.78 | 43 | 10 | 30 | 21
GlcCer d18:1/18:1 | 726.58 | 264.2 | 2.5 | 45.5 | 10 | 26 | 51
GlcCer d18:1/18:1 | 726.58 | 708.5 | 2.5 | 45.5 | 10 | 20 | 23
GlcCer d18:1/18:1-d5 | 731.36 | 269.2 | 2.5 | 70 | 10 | 18 | 49
GlcCer d18:1/18:1-d5 | 731.36 | 551.5 | 2.5 | 70 | 10 | 40 | 25
GlcCer d18:1/24:1 13C Isotope | 811.67 | 264.3 | 3.45 | 45 | 10 | 12 | 53
GlcCer d18:1/24:1 13C Isotope | 811.67 | 793.7 | 3.45 | 45 | 10 | 20 | 23
GlcCer d18:1/24:1-d5 | 815.71 | 269.2 | 3.45 | 45 | 10 | 26 | 51
GlcCer d18:1/24:1-d5 | 815.71 | 797.3 | 3.45 | 45 | 10 | 30 | 25
LacCer d18:1/16:0 13C Isotope | 863.62 | 264.3 | 2.29 | 65.5 | 10 | 28 | 59
LacCer d18:1/16:0 13C Isotope | 863.62 | 521.5 | 2.29 | 65.5 | 10 | 8 | 33
LacCer d18:1/16:0-d3 | 865.641 | 264.4 | 2.29 | 65 | 10 | 28 | 59
LacCer d18:1/16:0-d3 | 865.641 | 523.3 | 2.29 | 65 | 10 | 8 | 33
LacCer d18:1/17:0 | 876.598 | 264.4 | 2.43 | 60 | 10 | 14 | 59
LacCer d18:1/17:0 | 876.598 | 534.5 | 2.43 | 60 | 10 | 16 | 35
LacCer d18:1/18:0 | 890.628 | 264.2 | 2.57 | 60 | 10 | 14 | 59
LacCer d18:1/18:0 | 890.628 | 548.5 | 2.57 | 60 | 10 | 24 | 31
LacCer d18:1/18:1 | 888.7 | 264.3 | 2.33 | 50 | 10 | 14 | 60
LacCer d18:1/18:1 | 888.7 | 546.7 | 2.33 | 50 | 10 | 20 | 35
LacCer d18:1/24:0 | 974.7 | 264.3 | 3.58 | 45.5 | 10 | 10 | 67
LacCer d18:1/24:0 | 974.7 | 632.6 | 3.58 | 45.5 | 10 | 24 | 33
LacCer d18:1/24:0-d7 | 981.8 | 271.1 | 3.58 | 45 | 10 | 20 | 61
LacCer d18:1/24:0-d7 | 981.8 | 639.3 | 3.58 | 45 | 10 | 24 | 33
LacCer d18:1/24:1 13C Isotope | 973.7 | 264.2 | 3.21 | 45.5 | 10 | 22 | 63
LacCer d18:1/24:1 13C Isotope | 973.7 | 631.5 | 3.21 | 45.5 | 10 | 28 | 33
Standards and quality controls (STDs, QCs)
Concentration of standard/QC working solution [ng/mL]
| Kategorie A | Kategorie B | Kategorie C |  |  |  | Kategorie A
STD1 | 0.2 | 0.5 | 2 |  |  |  | S1P d18:1
STD2 | 0.4 | 1 | 4 |  |  |  | S1P d18:0
STD3 | 1.2 | 3 | 12 |  |  |  | S1P d20:1
STD4 | 2.4 | 6 | 24 |  |  |  | SPH d20:0
STD5 | 9 | 22.5 | 90 |  |  |  | SPH d20:1
STD6 | 36 | 90 | 360
STD7 | 120 | 300 | 1200
STD8 | 320 | 800 | 3200
STD9 | 800 | 2000 | 8000
STD10 | 1000 | 2500 | 10000
LLQC | 0.6 | 1.5 | 6
LQC | 3.6 | 9 | 36
MQC | 300 | 750 | 3000
HQC | 750 | 1875 | 7500
Internal Standard
Concentration of standard/QC working solution [ng/mL]
| Kategorie A | Kategorie B |  |  |  |  | Kategorie A
IS | 35 | 200 |  |  |  |  | SPH d18:1 d7
SPH d18:0 d7
Sphingolipids (Ceramides)
Standards and quality controls (STDs, QCs)
Concentration of standard/QC working solution [ng/mL]
| Kategorie A | Kategorie B | Kategorie C | Kategorie D | Kategorie E |  | Kategorie A
STD1 | 0.75 | 3.125 | 3.75 | 7.5 | 28.75 |  | Cer d18:1/14:0
STD2 | 1.5 | 6.25 | 7.5 | 15 | 57.5 |  | GlcCer d18:1/18:1
STD3 | 3 | 12.5 | 15 | 30 | 115 |  | GlcCer d18:1/18:0
STD4 | 7.5 | 31.25 | 37.5 | 75 | 287.5
STD5 | 18 | 75 | 90 | 180 | 690
STD6 | 36 | 150 | 180 | 360 | 1380
STD7 | 60 | 250 | 300 | 600 | 2300
STD8 | 120 | 500 | 600 | 1200 | 4600
STD9 | 240 | 1000 | 1200 | 2400 | 9200
STD10 | 300 | 1250 | 1500 | 3000 | 11500
LLQC | 2.25 | 9.375 | 11.25 | 22.5 | 86.25
LQC | 9 | 37.5 | 45 | 90 | 345
MQC | 102 | 425 | 510 | 1020 | 3910
HQC | 225 | 937.5 | 1125 | 2250 | 8625
Purchasing
Reference compound | Vendor
SPH d18:1 | Avanti Polar Lipids
SPH d18:0 | Avanti Polar Lipids
SPH d20:1 | Avanti Polar Lipids
SPH d20:0 | Avanti Polar Lipids
S1P d16:1 | Avanti Polar Lipids
S1P d18:1 | Avanti Polar Lipids
S1P d18:0 | Avanti Polar Lipids
Cer d18:0/16:0 | Avanti Polar Lipids
Cer d18:0/18:0 | Avanti Polar Lipids
Cer d18:0/24:0 | Avanti Polar Lipids
Cer d18:0/24:1 | Avanti Polar Lipids
Cer d18:1/14:0 | Avanti Polar Lipids
Cer d18:1/16:0 | Avanti Polar Lipids
Cer d18:1/18:0 | Avanti Polar Lipids
Cer d18:0/18:1 | Avanti Polar Lipids
Cer d18:1/20:0 | Avanti Polar Lipids
Cer d18:1/22:0 | Avanti Polar Lipids
Cer d18:1/24:0 | Avanti Polar Lipids
Cer d18:1/24:1 | Avanti Polar Lipids
GlcCer d18:1/16:0 | Avanti Polar Lipids
GlcCer d18:1/18:0 | Avanti Polar Lipids
GlcCer d18:1/18:1 | Avanti Polar Lipids
GlcCer d18:1/24:1 | Avanti Polar Lipids
LacCer d18:1/16:0 | Avanti Polar Lipids
LacCer d18:1/18:0 | Avanti Polar Lipids
LacCer d18:1/18:1 | Avanti Polar Lipids
LacCer d18:1/24:0 | Avanti Polar Lipids
LacCer d18:1/24:1 | Avanti Polar Lipids
*Avanti Polar Lipids | Alabaster, AL, USA
*Chiroblock GmbH | Wolfen, Germany
Internal Standard
Isotopically labeled internal standard | Vendor | Category | Concentration of standard working solution [ng/mL] | Volume of working solution [µL] | Volume of Z-buffer [µL] | Concentration related to sample [ng/mL]
SPH d18:1-d7 | Avanti Polar Lipids | A | 35 | 20 | 10 | 70
SPH d18:0-d7 | Avanti Polar Lipids | A | 35 | 20 | 10 | 70
S1P d18:1-d7 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
S1P d18:0-d7 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
Cer d18:0/18:0-d3 | Avanti Polar Lipids | C | 175 | 20 | 10 | 350
Cer d18:0/24:0-d7 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
Cer d18:1/16:0-d7 | Avanti Polar Lipids | A | 35 | 20 | 10 | 70
Cer d18:1/18:0-d7 | Avanti Polar Lipids | C | 175 | 20 | 10 | 350
Cer d18:1/18:1-d7 | Avanti Polar Lipids | C | 175 | 20 | 10 | 350
Cer d18:1/20:1-d7 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
Cer d18:1/22:1-d7 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
Cer d18:1/24:0-d4 | Chiroblock GmbH | D | 200 | 20 | 10 | 400
Cer d18:1/24:1-d7 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
GlcCer d18:1/18:0-d5 | Avanti Polar Lipids | A | 35 | 20 | 10 | 70
GlcCer d18:1/18:1-d5 | Avanti Polar Lipids | B | 50 | 20 | 10 | 100
GlcCer d18:1/24:1-d5 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
LacCer d18:1/16:0-d3 | Avanti Polar Lipids | E | 1000 | 20 | 10 | 2000
LacCer d18:1/17:0 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
LacCer d18:1/24:0-d7 | Avanti Polar Lipids | D | 200 | 20 | 10 | 400
*Avanti Polar Lipids | Alabaster, AL, USA
*Chiroblock GmbH | Wolfen, Germany
Acceptance criteria measurment
Procedure | acceptance criteria | accurancy
standards | 75% (6 /8) | ±15%
QC-Sets (LQC, MQC, HQC) | 67% (2/3) | ±15%
QC sample at each concentration | 50% (1/2)
Chromatogram
Sphingoid bases
Ceramides