Extended Data Figure/Table Legends

Figure S1: AutoCMR reduces required scan interactions and scan time variability.

a, User Interface (UI) mouse tracking (green) during a conventional 2DCMR exam (left) and an AutoCMR exam (right). While the conventional 2DCMR exam requires hundreds of user interactions, the AutoCMR exam requires minimal user input, requiring only an initial start click and the entry of the type and amount of contrast given. b, The number of mouse clicks (mean ± standard deviation) recorded during 12 conventional 2DCMR exams during localization, cine, and delayed enhancement imaging (120 ± 80 clicks) is significantly higher (p<0.0001) than the number of clicks recorded during 33 AutoCMR exams (5.3 ±0.7 clicks). c,d, Clinical 2DCMR and AutoCMR scan duration were calculated for the 100 AutoCMR patients scanned at an academic hospital (c) and the 25 AutoCMR patients scanned at a CHC (d). All clinical 2DCMR was performed at an Academic hospital, and scan time was retrospectively calculated counting only localizers, cine, and DE acquisitions. Scan duration (mean ± standard deviation) was significantly lower in AutoCMR acquisitions than in 2DCMR acquisitions at both the academic hospital (2DCMR: 30.6 ± 8.4 min; ACMR: 25.0 ± 2.3 min; p<0.001) and the CHC (2DCMR: 37.7 ± 8.4 min; ACMR: 26.3 ± 2.3 min; p<0.001). Only 5 total AutoCMR exams (Academic Hospital: 3%, CHC: 8%) were longer than 30 minutes and all AutoCMR exams were less than 35 minutes. In contrast, 60 (48%) of the 125 total 2DCMR exams were longer than 30 minutes, with 36 (29%) exams spanning longer than 35 minutes.

Figure S2: Validation of AutoCMR cine self-navigation and LV function.

a, b, Heart rate (a) and total number of detected peaks (b) from AutoCMR self-gating (SG) and extracted ECG signals (ECG) in 100 patients scanned at an Academic Hospital are highly correlated (ICC=0.97, 0.99) and demonstrate low bias (-0.257 bpm, 4.79 peaks) in Bland-Altman analysis. c-f. LV function from single slice images at the middle of the LV in the short axis plane is compared for 2DCMR and AutoCMR acquired in a single scanning session. Because AutoCMR images have been reformatted into 2D views, this single slice comparison avoids the confounding influence of comparing 3D AutoCMR data to volumes calculated from thick-slice, 2D image stacks with gaps between slices. Slice-matched 2DCMR and AutoCMR images were semi-automatically segmented using CVI42 Circle and manually corrected. LV end diastolic area (c) and fractional area change (d), calculated as FAC = 100 × (end diastolic area – end systolic area) / end diastolic area, demonstrate excellent correlation (ICC=0.90, 0.95) and low bias (-2.48 cm2, 1.63%) in Bland-Altman analysis. e, A line projection (blue) through cine images from a representative patient plotted over time is used to generate m-mode plots, which demonstrate qualitatively comparable change in imaging features over time for 2DCMR and AutoCMR images. f, LV endocardial area of the representative patient measured for each cardiac frame using 2DCMR and AutoCMR cine demonstrate agreement, with overlapping error as determined by calculating a 1 pixel offset on endocardial borders.

Figure S3: AutoCMR phantom validation.

a, A custom resolution phantom was printed with 3x3 grids providing 2.4 mm (blue), 2.0 mm (orange), 1.6 mm (green), and 1.0 mm (red) resolution. A reference image was acquired using fully sampled Cartesian data acquired with a spoiled gradient echo acquisition (matrix=192x192x192, resolution=1.6mm isotropic, FOV=307mm, TR/TE=4.0/2.9ms, FA=12°). Cine data was acquired with the proposed sequence (spoiled gradient echo, matrix=192x192x192, resolution=1.6mm isotropic, FOV=307mm, TR/TE=4.0/2.9ms, FA=12°) and was respiratory and cardiac binned into 30 cardiac frames using the self-gating signals from a representative patient. DE data was also collected using the proposed sequence (inversion recovery spoiled gradient echo, matrix=192x192x192, resolution=1.6mm isotropic, FOV=307mm, TR/TE=3.7/2.6ms, FA=10°) and was respiratory binned using the self-gating signals from the same representative patient. Both magnitude (DE MAG) and phase-sensitive inversion recovery (DE PSIR) like contrasts were generated. All images were interpolated by a factor of 2. b, Line plots of signal intensity drawn through each of the resolution grids show resolved peaks for 2.4mm, 2.0mm, and 1.6mm in Reference, Cine (cardiac frame 0), DE MAG (TI=250ms), and DE PSIR (TI=250ms) images. c, d, The 3T T1mes50 phantom was scanned using the AutoCMR DE acquisition and images were reconstructed with TI = [80, 340] in steps of 10 ms using an 80ms reconstruction window. d, Signal intensity from PSIR images and voxels in each of the 9 regions of interest is plotted over all reconstructed inversion times, demonstrated expected inversion recovery curves.

Figure S4: 3D isotropic AutoCMR images can be reformatted into any desired viewing plane.

a, c, AutoCMR cine images acquired at an academic hospital (a) and a community health center (c) with a temporal resolution of 30 frames per heartbeat. All frames were automatically segmented, providing LV, RV, LA, RA, myocardium, and aorta segmentations (bottom right). b, d, AutoCMR DE images acquired at an academic hospital (b) and a community health center (d) providing isotropic, whole heart coverage and multiple inversion times (bottom right).

Figure S5: Auto-segmentation and view finding validation.

a, b, Manual (a) and auto (b) cine image segmentations for a representative patient are observed to be qualitatively comparable in a 4-chamber view in LVED (top) and LVES (bottom). c-l, Auto-segmentations of the LV, RV, LA, RA, Myocardium, and Aorta are compared to cardiologist-reviewed manual segmentations in a 10-patient test set. Dice scores calculated for segmentations in (c) LVED and (d) LVES are > 0.8 for all segmentations across all patients. In Bland-Altman analysis, volumes calculated from manual and auto segmentations of the four chambers of the heart (e, i, LV, f, j, RV, g, k, LA, h, l, RA) in LVED (top) and LVES (bottom) are observed to have low bias. Atrial volumes are noted to demonstrate higher bias in LVES than LVED. m-o, Automated segmentations are used to generate AutoCMR images reformatted into traditional 2DCMR views, as shown for a representative patient (m, 4-chamber, n, 2-chamber, o, Short-axis). p, Vector norms of automatically generated views for AutoCMR data for 100 patients scanned at an academic hospital are compared to vector norms of the corresponding 2DCMR images acquired in the same scanning session and planned by an expert CMR technologist with > 20 years of experience. The angular difference in vector norms is generally lowest in the short-axis view (8.5±3.9°), with vector norms for all views less than 14° on average.

Figure S6: AutoCMR and 2DCMR agree in additional volume and function metrics.

a-g, In the pre-clinical model, volumetric (a, LVESV, b, LV Mass (LVM), c, RVESV, d, LAV in LVED, e, RAV in LVED) and functional (f, LVCO, g, RVCO) metrics demonstrate no significant difference (p>0.05) in paired t-tests comparing thin-slice 2DCMR cine to AutoCMR cine and low bias in Bland-Altman analysis. h-t, In 100 patients scanned at an Academic Hospital, AutoCMR cine volumetric (h, LVESV, i, LVM, j, RVESV, k, Minimum LA Area (LAA min) calculated from single-slice 4-chamber images, l, Minimum RA Area (RAA min) calculated from single-slice 4-chamber images, m, maximum LA Volume (LAV Max), n, minimum LA Volume (LAV Min), o, maximum RA Volume (RAV Max), p, Minimum RA Volume (RAV Min)) and functional (q, LVCO, r, RVCO, s, LA EF, t, RA EF) metrics are correlated with same-day 2DCMR (ICC≥0.56). LVM calculated using AutoCMR is overestimated (bias=-34.2g) in part due to difficulty in distinguishing the myocardium from epicardial fat in AutoCMR cine images. RVESV and atrial volumes and calculated functional metrics, although correlated, are biased by the calculation of 3D volumes from 2D images in 2DCMR versus true 3D volume calculation in AutoCMR data. u, Mid-ascending aorta diameter (MAAD) measurements from automated aorta segmentations at manually selected locations on AutoCMR cine data in LVED are minimally biased (0.125) and highly correlated (ICC=0.98) with clinically acquired MAAD measurements from CTA and MRA in 16 total patients (CTA: 9 patients; MRA: 7 patients).  v-ah, In 25 patients scanned at a Community Health Center, AutoCMR cine volumetric (v, LVESV, w, LVM, x, RVESV, y, LAA min, z, RAA min, aa, LAV Max, ab, LAV Min, ac, RAV Max, ad, RAV Min) metrics are correlated with 2DCMR data previously acquired at an academic hospital (ICC≥0.72). Reduced correlation of functional metrics (ae, LVCO, af, RVCO, ag, LA EF, ah, RA EF) and increased bias compared to corresponding results at an Academic Hospital are expected due to elapsed time (up to 1 year) between CHC AutoCMR and 2DCMR exams.

Table S7: AutoCMR Patient Population.

Patient demographics and clinical characteristics are reported for 52 patients in the auto-segmentation model training and validation set, 100 patients scanned with AutoCMR at an academic hospital, a subset of 30 patients scanned at an academic hospital and included in the diagnostic and image quality review, 25 patients scanned with AutoCMR at a community health center (CHC), and a subset of 8 patients scanned at a CHC and included in the diagnostic and image quality review.

|  | Training / Validation (N=52) | Academic Hospital (N=100) | Academic Hospital Subset (N=30) | CHC (N=25) | CHC Subset (N=8) |
| --- | --- | --- | --- | --- | --- |
| Age, mean (SD), years | 57.0 (17.7) | 53.9 (15.6) | 50.9 (17.6) | 61.5 (13.3) | 63.9 (12.7) |
| Sex, no. male (%) | 21 (40%) | 48 (48%) | 13 (43%) | 13 (52%) | 5 (63%) |
| Height, mean (SD), cm | 168 (14) | 170 (12) | 167 (17) | 169 (8.6) | 168 (5) |
| Weight, mean (SD), kg | 72.0 (14.8) | 84.1 (19.0) | 75.4 (17.0) | 82.9 (19.1) | 74.3 (9.6) |
| BMI, mean (SD), kg/m2 | 25.4 (11.7) | 29.0 (9.1) | 28.2 (13.1) | 28.8 (5.7) | 26.2 (3.3) |
| BSA, mean (SD), m2 | 1.75 (0.40) | 1.96 (0.33) | 1.86 (0.24) | 1.96 (0.25) | 1.86 (0.13) |
| LGE positive, no. (%) | 22 (42%) | 46 (46%) | 16 (53%) | 8 (32%) | 3 (38%) |
| ICM, no (%) | 1 (2%) | 11 (11%) | 5 (17%) | 3 (12%) | 1 (13%0 |
| NICM, no (%) | 21 (40%) | 34 (34%) | 14 (47%) | 7 (28%) | 3 (38%) |
| DCM, no (%) | 3 (6%) | 5 (5%) | 0 (0%) | 1 (4%) | 1 (13%) |
| HCM, no (%) | 9 (17%) | 20 (20%) | 5 (17%) | 6 (24%) | 1 (13%) |
| Valve Dysfunction, no (%) | 20 (38%) | 29 (29%) | 15 (50%) | 8 (32%) | 1 (13%) |
| Pericarditis, no (%) | 7 (13%) | 4 (4%) | 1 (3%) | 0 (0%) | 0 (0%) |
| Dilated Aorta, no (%) | 11 (21%) | 10 (10%) | 4 (13%) | 1 (4%) | 0 (0%) |
| Cardiac Mass Thrombus, no (%) | 0 (0%) | 1 (1%) | 1 (3%) | 0 (0%) | 0 (0%) |
| No Significant Findings, no (%) | 12 (23%) | 40 (40%) | 10 (33%) | 10 (40%) | 5 (63%) |
| Other, no (%) | 3 (6%) | 3 (3%) | 0 (0%) | 3 (12%) | 0 (0%) |

Video S8: AutoCMR data integrates with extended reality platforms for advanced data visualization and interaction.

AutoCMR cine images and segmentations from a patient with aortic valve regurgitation were imported to the extended reality platform CWRUXR and visualized with an Android-compatible phone.

Figure S9: AutoCMR cine segmentations provide high temporal resolution, reducing computational burden for digital twins.

The high-quality LV volume curve across a single cardiac cycle is depicted with blue data points, representing 10 selected LV geometries from end-diastole (ED) to the end of the cardiac cycle. Two methods for determining the stress-free (unloaded) LV geometry are illustrated. Method I: Utilizing the Klotz curve62,63 with the assumed ED pressure (EDP∼8 mmHg), an empirical estimated stress-free LV volume is calculated. This volume guides selection of the closest intermediate geometry from the 30-frame AutoCMR dataset, representing a near-zero-pressure state occurring prior to ED. The Method I unloaded geometry is calculated in seconds. Method II: Starting from the ED geometry, (1) the LV cavity volume is incrementally increased and a nonlinear finite-element problem is solved to compute the resulting internal pressure (∼ 8 mmHg), producing a trial loaded geometry. The resulting pressure, cavity volume, and shape are then compared with the measured ED state (2). The volume increment is adjusted and the loading step repeated until good agreement is achieved between the computed pressure, volume, and geometry and the measured state. The converged displacement field is inverted and applied to estimate a zero-pressure geometry (3). This geometry is reloaded to the target pressure (4) and the cycle is repeated over steps 2-4 until the simulation reproduces the measured ED pressure–volume relationship63.  On a 16-core workstation a typical run of four to five unload–reload cycles converges in approximately one hour; runtime decreases almost linearly with additional cores. While offering greater than 1000x reduction in computational time, the directly selected AutoCMR geometry (Method I) closely matches the inverse displacement geometry (Method II), showing minor deviations primarily in myocardial thickness, which is an expected outcome given the fixed boundary conditions and material properties.

Figure S10: AutoCMR cine and delayed enhancement imaging is feasible with metal implants, high BMI, arrythmia, and diffuse LGE.

a, Metal implants and sternal wires are common in cardiac MRI patients. Because AutoCMR cine and DE acquisitions employ spoiled gradient echo acquisitions, susceptibility artifacts surrounding metal are localized and minimized compared to often employed bSSFP acquisitions. i. AutoCMR cine and DE images in a patient with a patent foramen ovale closure device are observed to have minimal artifact. ii-iii. AutoCMR cine (ii) and DE (iii) imaging in a patient with a mitral valve ring and sternotomy wires showing localized susceptibility artifacts. b, The efficacy of regional-optimized virtual coil combination (ROVir28) in reducing wrap artifacts is demonstrated in a representative high BMI patient (33.1 kg/m2). c, AutoCMR cine acquisitions are observed to be robust to cardiac arrythmia through rejection of data collected during heartbeats with outlier RR-intervals. Through pulse-oximeter and ECG monitoring at the time of the exam, an example patient was observed to have a regular heart rate during the first five minutes of the cine acquisition, and frequent arrythmia during the remaining five minutes. This trend was similarly observed in the cardiac self-gating signal (i) and the corresponding RR-interval distribution (ii) calculated with peak detection applied to the cardiac signal (blue dots). Despite arrythmia beginning mid-scan, the resulting AutoCMR cine image quality is preserved (iii).  d, AutoCMR DE acquisitions detect enhancement in ischemic and non-ischemic patterns. In an example patient, mild enhancement in the inferoseptal wall and near RV insertion points is observed in both clinical 2DCMR and AutoCMR DE.