Supplementary data

Standard Preparation For heavy metal analysis on ICP-OES

We Calibrate the following elements to be analyzed on our ICP-OES Manganese (Mn), Lead (Pb), Chromium (Cr) and Cadmium (Cd). We group the above elements into two as follow, and  prepared the following concentrations by pouring from the 1000ppm stock solutions and will prepare eight standard series for each group below table.

Standard A                                            Standard B

Pb    0.35ml        1.4ppm                    Mn   1 ml             4 ppm

Cr     0.35 ml      1.4 ppm

Cd    0.35 ml     1.4 ppm

Table 1: Standard series and intensity for heavy metal analysis on ICP-OES

| Standard A |  |  |  |  |  |  |  | Standard B |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Standard Series | Volume ml | Pb ppm | Intensity Cps Pb | Cr ppm | Intensity Cps Cr | Cd ppm | Intensity Cps Cd | Mn ppm | Intensity Cps Mn |
| STD O | 0 | 0 | 2249.46 | 0 | 3565.85 | 0 | 2159.54 | 0 | 4000.92 |
| STD 1 | 1 | 0.028 | 2550.17 | 0.028 | 4677.13 | 0.028 | 4574.07 | 0.08 | 21765.9 |
| STD 2 | 2 | 0.056 | 2760.63 | 0.056 | 5890.4 | 0.056 | 7117.68 | 0.16 | 40111.2 |
| STD 3 | 3 | 0.084 | 2934.16 | 0.084 | 6922.76 | 0.084 | 9511.17 | 0.24 | 59487.8 |
| STD 4 | 10 | 0.28 | 4248.72 | 0.28 | 14116.4 | 0.28 | 25730.4 | 0.8 | 190843.0 |
| STD 5 | 20 | 0.56 | 6156.55 | 0.56 | 25439.0 | 0.56 | 49826.5 | 1.6 | 375443 |
| STD 6 | 30 | 0.84 | 7943.14 | 0.84 | 36111.4 | 0.84 | 73729.1 | 2.4 | 556720 |
| STD 7 | 40 | 1.12 | 9849.38 | 1.12 | 46790.4 | 1.12 | 97566.8 | 3.2 | 754182 |
| STD 8 | 50 | 1.4 | 11687.3 | 1.4 | 59530.8 | 1.4 | 120513 | 4.0 | 957068 |

N.b

All Stadared stock solution Are 1000ppm in 3% nitric acid which are imported from inorganic Ventures.  By dilutions with 0.5 % Nitric acid we prepare the calibration scheme As follows.

C1   Mn   1000mg/l              C2  Mn  4 mg/l

V1      ?                                      V2         250 ML

V1=C2V2/C1

=   4 mg/l Mn*250 ml/1000 ml

=    1 ml

To prepare the working standards we prepare as follow from the above intermediate solution.

C1   std1   4mg/l              C2 std 1   0.08 mg/l

V1    =?                                V2 sid1    50ml

V1 = 0.08 mg/l*50ml/4mg/l

V1   =   1ml

Table 2: Iinstrument Detection Limit and wave length

| Metals | IDL (mg/L) | Wave length (nm) |
| --- | --- | --- |
| Pb | 0.0007 | 220.353 |
| Cd | 0.0001 | 231.604 |
| Cr | 0.001 | 267.716 |
| Mn | 0.0002 | 324.754 |

Table 3: Standard concentration of the metals to be analyzed by ICP-OES

| Metals | Concentration of standards (mg/l) |
| --- | --- |
| Cd | 0.028, 0.056, 0.084, 0.28, 0.56, 0.84, 1.12, 1.4 |
| Pb | 0.028, 0.056, 0.084, 0.28, 0.56, 0.84, 1.12, 1.4 |
| Cr | 0.028, 0.056, 0.084, 0.28, 0.56, 0.84, 1.12, 1.4 |
| Mn | 0.08, 0.16, 0.24, 0.8, 1.6, 2.4, 3.2, 4.0 |

Table 4: The values of physico-chemical, heavy metal parameters and WHO values

| Sampling station | Parameters |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Temp(°c) | PH | EC(µS/cm) | TDS(ppm) | Cd(mg/L) | Pb(mg/L) | Cr(mg/L) | Mn(mg/L) |
| sample 1 | 19.1 | 7.08 | 187.2 | 131.4 | 0.00557 | 0.00761 | 0.02642 | 0.07307 |
| sample 2 | 18.9 | 7.11 | 174.5 | 138.7 | 0.00672 | 0.00826 | 0.02957 | 0.07042 |
| sample 3 | 19.2 | 6.72 | 185.3 | 137.5 | 0.00678 | 0.00778 | 0.04353 | 0.08513 |
| sample 4 | 19.0 | 6.54 | 206 | 140.1 | 0.00712 | 0.00854 | 0.03528 | 0.06126 |
| sample 5 | 19.2 | 6.65 | 203.7 | 139.6 | 0.00701 | 0.0083 | 0.04241 | 0.09053 |
| Mean | 19.08 | 6.86 | 191.34 | 137.46 | 0.00664 | 0.00810 | 0.03544 | 0.076082 |
| WHO | 25 | 6.5-8.5 | <400 | <500 | 0.003 | 0.01 | 0.05 | 0.1 |