Supplementary Tables

Supp. Table 1. List of 176 targets included in mNGS assay and their associated category (see Excel file).

Supp. Table 2. Reference and discrepancy testing of clinical samples by DHM and Microba.

|  | Seegene PCR Assay |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | GI-Bacteria (I) | GI-Bacteria (II) | H. pylori | GI-Helminth | GI-Parasite | GI-Virus | MCS | MALDI-TOF | Antigen |
| DHM | 480 | 0 | 0 | 0 | 0 | 162 | 468 | 468 | 136 |
| Microba | 476 | 468 | 51 | 47 | 470 | 55 | 0 | 0 | 0 |

Supp. Table 3. mNGS species or target names comprising comparison reference targets along with Seegene PCR assays and Ct thresholds.

| Reference Target | Seegene PCR Assay | Ct threshold | mNGS Assay Equivalent Targets |
| --- | --- | --- | --- |
| Aeromonas spp. | GI-Bacteria (I) | ≤45 | A. caviae, A. dhakensis, A. hydrophila, A. veronii |
| Campylobacter spp. | GI-Bacteria (I) | ≤45 | C. coli, C. jejuni |
| Edwardsiella tarda | - | - | E. tarda |
| Helicobacter pylori | H. pylori & ClariR | ≤50 | H. pylori |
| Plesiomonas shigelloides | - | - | P. shigelloides |
| Salmonella spp. | GI-Bacteria (I) | ≤45 | S. bongori, S. enterica |
| Vibrio spp. | GI-Bacteria (I) | ≤45 | V. cholerae, V. parahaemolyticus, V. vulnificus |
| Yersinia enterocolitica | GI-Bacteria (I) | ≤45 | Y. enterocolitica |
| Cryptosporidium spp. | GI-Parasite | ≤43 | C. felis, C. meleagridis, C. muris, C. ubiquitum, C. viatorum, C. hominis/parvum/cuniculus/tyzzeri complex |
| Cyclospora cayetanensis | GI-Parasite | ≤43 | C. cayetanensis |
| Entamoeba histolytica | GI-Parasite | ≤43 | E. histolytica |
| Enterobius vermicularis | GI-Helminth (I) | ≤45 | E. vermicularis |
| Giardia lamblia | GI-Parasite | ≤43 | G. intestinalis |
| Strongyloides stercoralis | GI-Helminth (I) | ≤45 | S. stercoralis |
| Taenia spp. | GI-Helminth (I) | ≤45 | T. saginata, T. solium |
| C. difficile Toxin B | GI-Bacteria (I) | ≤45 | C. difficile toxin A/B |
| EAEC virulence factors | GI-Bacteria (II) | ≤43 | EAEC virulence factors |
| Shigella spp. / EIEC virulence factors | GI-Bacteria (I) | ≤45 | EIEC invasion plasmid antigen |
| EPEC virulence factors | GI-Bacteria (II) | ≤43 | EPEC virulence factors |
| ETEC heat labile/stable toxins | GI-Bacteria (II) | ≤43 | ETEC heat labile and heat stable toxins |
| STEC Shiga toxin | GI-Bacteria (II) | ≤43 | STEC Shiga toxin |
| Adenovirus F | GI-Virus | ≤40 | Adenovirus F (serotype 40/41) |

Supp. Table 4. Number of samples with conventional test results for each clinical target.

| Target Name | Target Type | No. Samples |
| --- | --- | --- |
| Aeromonas spp. | Bacterial | 497 |
| Campylobacter spp. | Bacterial | 497 |
| Cryptosporidium spp. | Eukaryotic | 490 |
| Cyclospora cayetanensis | Eukaryotic | 496 |
| Edwardsiella tarda | Bacterial | 468 |
| Entamoeba histolytica | Eukaryotic | 497 |
| Enterobius vermicularis | Eukaryotic | 468 |
| Giardia intestinalis | Eukaryotic | 497 |
| Helicobacter pylori | Bacterial | 61 |
| Plesiomonas shigelloides | Bacterial | 464 |
| Salmonella spp. | Bacterial | 497 |
| Strongyloides stercoralis | Eukaryotic | 468 |
| Taenia spp. | Eukaryotic | 468 |
| Vibrio spp. | Bacterial | 497 |
| Yersinia enterocolitica | Bacterial | 497 |
| Adenovirus F (serotype 40/41) | Viral | 193 |
| C. difficile Toxin B | Virulence Factor | 497 |
| EAEC virulence factors | Virulence Factor | 468 |
| EIEC virulence factors | Virulence Factor | 497 |
| EPEC virulence factors | Virulence Factor | 468 |
| ETEC heat labile/stable toxins | Virulence Factor | 468 |
| STEC Shiga toxin | Virulence Factor | 468 |

Supp. Table 5. Simulated faecal spike-in samples were generated for increasing numbers of read pairs. Pathogen targets were generated at a fixed number of read pairs whereas read pairs for gene targets were generated to achieve a fixed depth of coverage. Estimated biomass is given for a faecal sample with 6.23 x 1010 equ. orgs/g faeces assuming 3.9 Mb organisms and 1000 bp genes. Numbers in parentheses indicate the percentage of reads from a target in a standard mNGS assay sample of 16 million read pairs and considering length for the gene targets.

| Depth | Pathogen Target | Pathogen Biomass (equ. orgs/g) | Gene Target | Gene Biomass (equ. orgs/g) |
| --- | --- | --- | --- | --- |
| Ultra-low | 16 (0.0001) | 6.23×104 | 0.5 (6.25×10-6 to 8.75×10-5; mean = 1.16×10-5) | 1.52×107 to 2.13×108 |
| Low | 160 (0.001) | 6.23×105 | 1.0 (6.25×10-6 to 1.69×10-4; mean = 2.11×10-5) | 1.52×107 to 4.11×108 |
| Medium | 1,600 (0.01) | 6.23×106 | 2.5 (1.25×10-5 to 4.25×10-4; mean = 5.75×10-5) | 3.04×107 to 1.03×109 |
| High | 16,000 (0.1) | 6.23×107 | 5.0 (2.50×10-5 to 8.50×10-4; mean = 1.16×10-4) | 6.07×107 to 2.07×109 |

Supp. Table 6. Targets identified by mNGS assay in faecal samples used to generate simulated communities with spiked-in targets.

| Sample ID | Targets Identified by mNGS Assay |
| --- | --- |
| BBW5020 | Campylobacter concisus |
| BBW5021 | Clostridium perfringens |
| BBW5022 | blaACT, blaCMY_Group2, blaSHV, Salmonella enterica, Campylobacter concisus |

Supp. Table 7. List of mNGS assay targets in each of the 22 groups used to generate in silico faecal spike-in samples (see Excel file).

Supp. Table 8. Diagnostic test results for each of the 510 DHM clinical samples (see Excel file).

Supp. Table 9. Outcome of mNGS assay for each target in the DHM data set.

| Bacteria | TP | FP | TN | FN | Prevalence (%) |
| --- | --- | --- | --- | --- | --- |
| Aeromonas spp. | 30 | 2 | 450 | 15 | 9.05 |
| Campylobacter spp. | 41 | 1 | 451 | 4 | 9.05 |
| Edwardsiella tarda | 3 | 0 | 465 | 0 | 0.64 |
| Helicobacter pylori | 1 | 0 | 40 | 20 | 34.43 |
| Plesiomonas shigelloides | 23 | 0 | 441 | 0 | 4.96 |
| Salmonella spp. | 31 | 1 | 456 | 9 | 8.05 |
| Vibrio spp. | 16 | 2 | 474 | 5 | 4.23 |
| Yersinia enterocolitica | 30 | 1 | 463 | 3 | 6.64 |
| Total | 175 | 7 | 3240 | 56 |  |

| Eukaryotes | TP | FP | TN | FN | Prevalence (%) |
| --- | --- | --- | --- | --- | --- |
| Cryptosporidium spp. | 36 | 2 | 452 | 0 | 7.35 |
| Cyclospora cayetanensis | 16 | 0 | 480 | 0 | 3.23 |
| Entamoeba histolytica | 2 | 0 | 492 | 3 | 1.01 |
| Enterobius vermicularis | 9 | 0 | 459 | 0 | 1.92 |
| Giardia intestinalis | 26 | 0 | 466 | 5 | 6.24 |
| Strongyloides stercoralis | 1 | 0 | 467 | 0 | 0.21 |
| Taenia spp. | 1 | 0 | 467 | 0 | 0.21 |
| Total | 91 | 2 | 3283 | 8 |  |

| Viruses | TP | FP | TN | FN | Prevalence (%) |
| --- | --- | --- | --- | --- | --- |
| Adenovirus F (serotype 40/41) | 41 | 4 | 120 | 28 | 35.75 |

| Virulence factors | TP | FP | TN | FN | Prevalence (%) |
| --- | --- | --- | --- | --- | --- |
| C. difficile toxin A/B | 34 | 0 | 444 | 19 | 10.66 |
| EAEC virulence factors | 56 | 1 | 393 | 18 | 15.81 |
| EIEC virulence factors | 38 | 0 | 434 | 25 | 12.68 |
| EPEC virulence factors | 29 | 0 | 372 | 67 | 20.51 |
| ETEC heat labile/stable toxins | 16 | 0 | 440 | 12 | 5.98 |
| STEC Shiga toxin | 1 | 0 | 466 | 1 | 0.43 |
| Total | 174 | 1 | 2549 | 142 |  |

| All targets | TP | FP | TN | FN |
| --- | --- | --- | --- | --- |
| Total | 481 | 14 | 9192 | 234 |

Supp. Table 10. Additional targets identified by the mNGS assay compared to current standard-of-care diagnostic testing performed by DHM (see Excel file).

Supp. Table 11. Performance of the mNGS assay compared to PCR reference testing (see Excel file).

Supp. Table 12. Sensitivity of mNGS assay on samples with Seegene PCR results stratified by Ct values greater or less than 35 (see Excel file).

Supp. Table 13. Reproducibility of mNGS targets across biological replicates (see Excel file).

Supp. Table 14. Performance of each target in the mNGS assay on the in silico faecal spike-in samples (see Excel file).

Supp. Table 15. Exploration of FP predictions for the in silico faecal samples (see Excel file).

Supplementary Figures

Supp. Figure 1. Decision trees for resolving discordant conventional and mNGS assay test results. Samples with a single conventional test result are either i) concordant with the mNGS assay and classified as a TP or TN, ii) discordant with the mNGS assay and subjected to an additional independent PCR test (Supp. Figure 1A). A majority vote (>50% consensus) was used to establish the expected result for samples provided with multiple conventional test results or where a single discordant test result prompted application of an additional independent PCR test (B). Samples where conventional tests had a majority result were used to evaluate the mNGS assay, whereas samples without a majority result (i.e. equal number of conventional tests indicating the absence or presence of a target) were subjected to additional PCR testing in order to establish a consensus point of truth.

Supp. Figure 2. Number of identified reads for mNGS targets with increasing concentration in a standard faecal sample for pathogen (A) and gene (B) targets. For pathogen targets, the red dotted lines indicate the ultra-low (16 read pairs), low (160 read pairs), and medium (1,600 read pairs) read depths used for the in silico samples. Identification of pathogen targets required 105 to 106 equ. orgs/g faeces where the number of identified reads fall between the ultra-low and low read depths.  For gene targets, the red dotted lines indicate the ultra-low (0.5x coverage; 1.6 read pairs), low (1x coverage; 3.2 read pairs), medium (2.5x coverage; 8.3 read pairs), and high (5x coverage; 16.6 read pairs) read depths used for the in silico samples assuming an average gene length of 1,000 bp. Robust identification of gene targets required ≥108 equ. orgs/g faeces where the majority of targets are at or above the low in silico read depth.

Supplementary Notes

Supp. Note 1. Simulated samples were used to determine the limit of detection of the mNGS assay by spiking in isolate cells at decreasing equ. orgs/g faeces into a homologous faecal matrix (see Analytical Performance on Simulated Samples). The faecal matrix was determined to have a biomass of 4.19 to 8.67 x 1010 equ. orgs/g in a faecal sample. Sensitivity reached 100% for all mNGS assay target pathogens at 106 equ. orgs/g (Table 2). This corresponds to a relative abundance of 0.00115% (100% x 106 / 8.67 x 1010) to 0.00239% (100% x 106 / 4.19 x 1010). For an mNGS assay operating at 16 million read pairs this corresponds to 184 to 382 read pairs from each species, assuming a genome size equal to the mean genome size of all organisms in the sample. The low depth in silico faecal samples were set at 160 read pairs (0.001% relative read abundance) to correspond with the expected limit of detection of a target requiring approximately 106 equ. orgs/g (Supp. Table 5). Similarly, ultra-low depth was set at 16 read pairs (0.0001% relative read abundance) to simulate targets detectable at 105 equ. orgs/g.

AMR and VF genes are ~1,000 bp and thus expected to have an mNGS assay limit of detection that is 3900 time lower than for a bacterial pathogen (3.9 Mb / 1000 bp). This is 3 orders of magnitude smaller which is in good agreement with the analytical results where bacterial pathogens have 100% sensitivity at 106 equ. orgs/g and only a subset of AMR and VF genes reach 100% sensitivity at 108 equ. orgs/g (i.e. indicating 109 equ. orgs/g or greater is required before all gene targets reach 100% sensitivity). Read pairs from AMR and VF gene targets were simulated to achieve a fixed depth of coverage between 0.5x and 2.0x as this corresponds to an expected biomass between 1.52 x 107 to 2.07 x 109 equ. orgs/g (Supp. Table 5) which is around the limit of detection established for AMR and VF genes using simulated samples.

Supp. Note 2. The mNGS assay demonstrates excellent specificity and PPV across all targets as a result of few FP predictions. The limited number of FPs that occurred were investigated to establish factors resulting in incorrect predictions (126 FPs compared to 8,558 TPs and 126,635 TNs; Supp. Table 14). Notably, FP predictions only occurred for 9 of 176 (5.1%) targets. The only two pathogens to result in FPs were Helicobacter cinaedi and H. fennelliae and these were the results of the in silico samples containing the other Helicobacter species (i.e. Helicobacter cinaedi being identified in samples with H. fennelliae, and vice versa). There were 3 AMR, 1 VF, and 3 VF with host targets resulting in FPs. Similar, to the Helicobacter FPs, the FP for CTX-M-G8 is the result of the in silico sample contain the closely related genes CTX-M-G1 and CTX-M-G9. The cause of the other FP AMR and VF gene predictions is less clear though may be the result of interaction with DNA in the faecal background or other targets spiked into the sample (e.g. identification of the Shiga toxin in a sample that contains stx1A as a spike in and identified to contain a Shigella/E. coli species in the faecal background).