Table 1. Specifications of EBT3 Radiochromic film

| Name of Film GAFChromic EBT3 Film |  |  |  |
| --- | --- | --- | --- |
| Structure of the Film | 28 μm active layer sandwiched between 125 μm surface matte polyester layers |  | 125 μm’lik yüzeydeki mat polyester katmanlar arasında sandviç şeklindeki 28 μm’lik aktif katman |
| Size | 8"x10" (~ 20x25 cm) |  | 8"x10" (~ 20x25 cm) |
| Dynamic Dose Range | 0.1 Gy – 20 Gy |  | 0,1 Gy – 20 Gy |
| Dependence of Energy | The difference in optical density is less than 5% when irradiated between 100 keV and 18 MeV |  | 100 keV ile 18 MeV arasında ışınlandığında optik yoğunluktaki fark %5’den küçük |
| Dependence of dose rate | The difference in optical density when 10 Gy is irradiated between dose rates of 3.4 Gy/min and 0.034 Gy/min is less than 5% |  | 3,4 Gy/dak ile 0,034 Gy/dak doz hızları arasında 10 Gy ışınlandığındaki optik yoğunluk farkı %5’den küçük |
| Stability Under Light | The optical density change under 1000 lx (lux) daily illumination intensity is less than 5x10-3 |  | Günlük 1000 lx(lüks) aydınlanma şiddeti altındaki optik yoğunluk değişimi 5x10-3’ den küçük |

Table 2. Specification of Anthropomorphic Paediatric Phantom and Dimensions

| Parts of phantom | Includes | Weight (kg) | Dimensions |
| --- | --- | --- | --- |
| Head | Skull, Sinus Cavities, Ear Canals, Mandibula and Maxilla, Representative Teeth | 2 kg | H: 17 cm W:14 cm D: 17 cm |
| Trunk | Chest, Rib Bones, Spinal Discs, Trachea, Scapula, Clavicle, Humerus Head, Pelvis, Femur Head, Lungs | 2 kg “This section is left blank to allow different organs to be studied with different materials.” | H: 34 cm W:21 cm |
| Arms | Humerus, Ulna and Radius Bones Hands and arms are designed in a natural hand position. | Weight of each arm:0.67 kg | H: 34 cm W:8 cm D: 5 cm |
| Right and Left leg | Femur, Tibia, Fibula and Foot (Talus, Metatarsal, Cuboid) The feet are designed in a dorsiflexion foot position. | Weight of each leg :2.3 kg | H: 42 cm W:9 cm D: 18 cm |

Table 3. CTDIvol and DLP values ​​of the paediatric phantom in the specified protocols

| Exam | Tuve Voltage (kVp) | Tube Current (mA) | Pitch | CTDIvol (mGy) | DLP (mGy.cm) | DE (mSv) |
| --- | --- | --- | --- | --- | --- | --- |
| Head | 120 | 100 | 1.01 | 38.9 | 455.53 | 1.82 |
|  | 100 | 100 | 1.01 | 33.4 | 393.69 | 1.57 |
| Abdomen and Pelvis | 120 | 100 | 1.22 | 5.5 | 219.5 | 4.39 |
|  | 100 | 100 | 1.22 | 5.1 | 115.78 | 2.32 |
| Thorax | 120 | 100 | 1.22 | 4.4 | 106.71 | 1.92 |
|  | 100 | 100 | 1.22 | 4.2 | 86.9 | 1.74 |

*DE: Dose Estimation

Table 4. Dose conversion factors for adult and paediatric patients per DLP

| Body region | k (mSv. mGy-1cm-1) |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | 0 age | 1 age | 5 age | 10 age | Adult |
| Head and Neck | 0.013 | 0.0085 | 0.0057 | 0.0042 | 0.0031 |
| Head | 0.011 | 0.0067 | 0.004 | 0.0032 | 0.0021 |
| Neck | 0.017 | 0.012 | 0.011 | 0.0079 | 0.0059 |
| Chest | 0.039 | 0.026 | 0.018 | 0.013 | 0.014 |
| Abdomen and Pelvis | 0.049 | 0.03 | 0.02 | 0.015 | 0.015 |
| Body | 0.044 | 0.028 | 0.019 | 0.014 | 0.015 |

Table 5. Comparison of paediatric phantom dose estimates with literature values

|  |  | CTDIvol |  |  |  | DLP |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CT Protokol | Age | This study | Galanski et al. (2005) | Shrimpton et al. (2003) | Gao et al. (2018) | This study | Galanski et al. (2005) | Shrimpton et al. (2003) | Gao et al. (2018) |
| Head | 1˗5 | 36 | 36 | 34 | 34 | 424.6 | 452 | 383 | 594 |
|  | 6˗10 |  | 44 | 44 | 36 |  | 582 | 508 | 599 |
| Abdomen and Pelvis | 1˗5 | 5.3 | 4 | 4 | 4 | 167.6 | 162 | NA | 162 |
|  | 6˗10 |  | 6 | 6 | 6 |  | 247 | NA | 247 |
| Thorax | 1˗5 | 4.3 | 3 | 11 | 3 | 96.8 | 61 | 198 | 89 |
|  | 6˗10 |  | 5 | 14 | 4 |  | 102 | 303 | 124 |