Table 1 Protocols for prostate mp-MRI and bp-MRI with recall system

| Sequence | 2D Axial T2 | 2D Coronal T2 | 2D Sagittal T2 | 2D Axial DWI | 3D Axial T1 DCE |
| --- | --- | --- | --- | --- | --- |
|  | FRFSE | FRFSE | FRFSE | FOCUS DW-EPI | DISCO FLEX |
| mp-MRI | ü | ü | ü | ü | ü |
| bp-MRI | ü | ü | ü | ü | X |
| Thickness (Gap) (mm) | 3.0 (0) | 3.0 (0) | 3.0 (0.5) | 3.0 (0) | 3.0 |
| Resolution (mm) | 0.40 x 0.67 (DL reconstruction 0.39 x 0.39) | 0.625 x 0.625 (DL reconstruction 0.39 x 0.39) | 0.750 x 0.833 (interpolated to 0.47 x 0.47) | 1.5 x 1.5 | 1.34 x 1.56 |
| FOV (cm) | 20.0 | 20.0 | 24.0 | 12.0 x 24.0 | 30.0 |
| NEX | 2.5 | 3 | 2 | 2 for b=50 s/mm2 and 16 for b=1400 s/mm2 | N/A |
| ETL | 23 | 25 | 21 | N/A | N/A |
| TR (ms) | >4000.0 | >3000.0 | >3000.0 | >3000.0 | <4.0 |
| TE (ms) | 105.0-115.0 | 100.0-110.0 | 115.0-125.0 | 60.0-65.0 | <2.0 |
| Receive Bandwidth (kHz) | ±41.7 | ±41.7 | ±35.7 | ±250.0 | ±142.9 |
| Acquisition Time | 4 m 56 s | 5 m 17 s | 3 m 37 s | 4 m 8 s | 3 m 36 s |
| Comments |  |  |  | b-values 50 and 1400 s/mm2 | 60 phases (temporal resolution 3-4 s) |

Position: Supine, arms down.

Contrast Agent: Gadobenate dimeglumine (MultiHance, Bracco, Milan, Italy) 0.06 mL/kg by body weight.

Power Injector Rate: 2.0 mL/s flow rate with 25.0 mL saline chaser.

All sequences acquired with 32 channel cardiac array coil. All two-dimensional (2D) sequences utilized Deep learn (DL) reconstruction (AIR Recon DL, GE Healthcare, Chicago, IL, USA).

Legend: Dynamic Contrast Enhanced (DCE). Fast Recovery Fast Spin Echo (FRFSE). Field-of-View Optimized and Constrained Undistorted Single-shot Diffusion-Weighted Echo Planar Imaging (FOCUS DW-EPI). Differential Sub-sampling with Cartesian Ordering with Flexible Temporal Resolution (DISCO FLEX). Field of View (FOV). Number of Excitations (NEX). Echo Train Length (ETL). Repetition Time (TR). Echo Time (TE).