Supplementary Table S1 The volumes measured by the GF220R during ventilation using a 1050 mL syringe in each gas mixture at 25 ˚C under atmospheric pressure and dry condition.

| Gas | Gas composition |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| N2 (%) O2 (%) N2O (%) S (%) D (%) CO2 (%) | 74 21 5 | 45 50 5 | 95 5 | 93 2 5 | 50 45 5 | 91 4 5 | 90 5 5 | 30 65 5 | 31 65 4 5 | 85 10 5 | 30 60 5 5 | 30 55 10 5 |
| Density (kg/m3) | 1.216 | 1.260 | 1.334 | 1.472 | 1.561 | 1.609 | 1.612 | 1.661 | 1.820 | 1.891 | 1.914 | 2.167 |
| Volume | 1024 ± 3.85 | 1024 ± 7.10 | 1029 ± 2.60 | 1023 ± 2.93 | 1022 ± 4.57 | 1023 ± 3.07 | 1017 ± 4.83 | 1009 ± 2.87 | 1021 ± 5.58 | 1005 ± 1.25 | 1008 ± 6.33 | 1000 ± 4.82 |

S = sevoflurane, D = desflurane

Supplementary equation S2.

The derivation of the equation (3) from the equation (2)

As described in the discussion section, total pressure integrated and controlled by the ventilator (Pvent) is divided into two parts as follows: the pressure inflating the test lung (Plung) and the pressure used for flowing (ΔE). Based on the Bernoulli's principle, this state is expressed as follows:

Pvent ＝ ΔE + Plung　··· equation (2)

Here, the Darcy–Weisbach equation expresses the pressure (∆E) using the flow volume, length, and diameter of the circuit, density of the fluid, and flow coefficient. This equation can also be expressed using the mean flow velocity of each section of the breathing circuit, such as the inspiratory limb, Y piece and tracheal tube Ui, the density of the gas mixture (ρ), and dimensionless loss coefficient of each section (ci') as follows:

∆E=i=1Nci'12ρUi2　··· equation (2-1)

The loss coefficient (ci') is composed of several factors, including loss of friction and length and diameter of each section. According to equation (2-1), the density of the gas mixture affects energy loss and circuit characteristics.

Because the volume flow rate (Q) is the product of the mean flow velocity (U) and the cross-sectional area (A) of the breathing circuit (Q = UA), the ΔE during the inspiratory phase of t seconds can be expressed using the equation (2-1) as follows:

ΔE during the inspiratory phase of t seconds≒ct12ρQin2··· equation (2-2)

The mean inspiratory volume flow rate is indicated by Qin and the mean dimensionless loss coefficient during the inspiratory phase of t seconds is indicated by ct. Moreover, the pressure inflating the test lung (Plung) during the inspiratory phase of t seconds is also expressed using the elastance coefficient of a test lung (cp') and inspiratory volume flow rate (Qin) as follows:

Plung=cp'0tQin(t')dt'··· equation (2-3)

The above equation is integrated over the inspiratory phase of 1 s and expressed using the elastance coefficient of the test lung during the inspiratory phase (cp) and mean inspiratory volume flow rate (Qin) as follows:

Plung=cpQin··· equation (2-4)

Therefore, the formula (2) in the inspiratory phase can be transcribed as follows:

Pvent=ct12ρQin2+cpQin ··· equation (3)

Supplementary equation S3

This section explains how pressure gradient, density, and flow rate relate in differential pressure flowmeters. There are roughly two types of differential pressure flowmeters used in clinical practice, namely fixed or variable orifice differential pressure flowmeter.

Pressure gradient (ΔP) between the upstream and downstream sides of the orifice is expressed by the following equation:

ΔP=12ρQA22-12ρQA12=121A22-1A12ρQ2 or Q2=2A12A22A12-A22ΔPρ … equation (6-1)

where A1 is the cross-sectional area of the upstream side of the orifice and A2 is that of the orifice. In the fixed type, the cross-sectional area of the orifice (A2) is a fixed value, which indicates that the pressure gradient is proportional to density and square of flow rate.

Meanwhile, for a variable orifice type, A2 is a function of ΔP (= AΔP). This relationship can be expressed by the following equation:

Q2=2A12AΔP2A12-AΔP2ΔPρ … equation (6-2)

Additionally, an equilibrium for the elastic restoring force of the orifice flap (Fx) at the pressure gradient of ΔP can be expressed by the following approximate equation:

Fx= ΔP*A1-AΔP or A1-AΔP= FxΔP … equation (6-3)

Where A(P), and (A1-AΔP) are the cross-sectional area of the variable orifice and effective surface area of the flap plate at the pressure gradient of ΔP, respectively.

Inserting equation (6-3) to equation (6-2) can be described as follows:

Q2=2A12AΔP2A12-AΔP2ΔPρ=2A12AΔP2A1 + AΔPA1 - AΔPΔPρ= 2A12AΔP2A1+AΔP*FxΔPΔPρ=2A12AΔP2A1+AΔPFxΔP2ρ

putting C(ΔP)=2A12AΔP2A1+AΔPFx

Q2= CΔPΔP2ρ, Q=CΔPρ*ΔP or ΔP=ρCΔP*Q … equation (7)

Equation (7) indicates that the pressure gradient is proportional to the flow rate and the square root of the density in a variable orifice type. Additionally, ρ is essential to estimate Q, as well as fixed type.

In any type of differential pressure flowmeter, volume flow rate (Q) can be expressed as the following general equation:

ΔP=12ρQAΔP2-12ρQA12=12ζ(ΔP)ρQAΔP2 or Q=2AΔP2ΔPζ(ΔP)ρ  ··· equation (7-1)

The compensation coefficient is indicated by ζ(ΔP), which is a function of ΔP. ρ is essential to estimate Q. Here, anesthesia machines without gas module, including this study’s Avance®, cannot use the respiratory gas density (ρdlv) to correct the measured values, and instead, use the preset density (ρpst). Thus, during ventilation with the gas density of ρdlv, the relationship between delivered flow (Qdlv) and the flow measured by the anesthesia machine (QA) can be expressed by the following equation:

Qdlv=2AΔP2ΔPζ(ΔP)ρdlv=2AΔP2ζ(P)ρpst2AΔP2QA2ζ(ΔP)ρdlv=ρpstρdlv*QA… equation (7-2)

This equation represents that the actual flow rate can be derived from the measured flow rate by any type of differential pressure flowmeter without a density compensation system.

Supplementary Table S4 Densities of gases at 0 ˚C and 25 ˚C under atmospheric pressure and dry condition.

|  | O2 | N2 | N2O | CO2 | Desflurane | Sevoflurane |
| --- | --- | --- | --- | --- | --- | --- |
| 0 ℃ (kg/m3) | 1.429 | 1.25 | 1.978 | 1.977 | 7.502 | 8.929 |
| 25 ℃ (kg/m3) | 1.309 | 1.145 | 1.812 | 1.811 | 6.873 | 8.180 |

Densities of known gases at 0 ˚C were obtained from National Astronomical Observatory of Japan Chronological scientific tables, 2019 Ed, Maruzen publishing. Densities of desflurane and sevoflurane at 0 ˚C were calculated from the molecular weight. Densities at 25 ˚C were calculated by the following formula. [Density at 25 ˚C in atmospheric pressure] = [density at 0 ˚C in atmospheric pressure] × 273.15/298.15.