[bookmark: _Hlk189821674]Supplementary Table 1 : 
A. [bookmark: _Hlk189821682]Comparison of the 7 lots of CSCM-NXC® media used to test disposable devices and their associations.
	
	CSCM-NXC® Lot (10)0000015060
n=111
	CSCM-NXC® Lot (10)0000015372
n=45
	CSCM-NXC® Lot (10)0000015420
n=30
	CSCM-NXC® Lot (10)0000018512
n=15
	CSCM-NXC® Lot (10)0000018536
n=30
	CSCM-NXC® Lot (10)0000018761
n=15
	CSCM-NXC® Lot (10)0000018848
n=15
	P-Value
	P-Value*
	Pool of 6 CSCM-NXC® Lots except Lot (10)0000018848
n=246

	Embryological rates
	
	
	
	
	
	
	
	
	
	

	   E1.5 Cleavage rate (n, %)
	111 (100.0)
	45 (100.0)
	30 (100.0)
	15 (100.0)
	30 (100.0)
	15 (100.0)
	15 (100.0)
	1.0000a
	1.0000 a
	246 (100.0)

	   E4.0 Blastulation rate (n, %)
	103 (92.8)
	43 (95.6)
	26 (86.7)
	13 (86.7)
	30 (100.0)
	15 (100.0)
	14 (93.3)
	0.4107b
	0.3018b
	230 (93.5)

	   E4.0 GQ Blastocysts rate (n, %)
	90 (87.4)
	37 (86.0)
	24 (92.3)
	12 (92.3)
	28 (93.3)
	15 (100.0)
	5 (35.7)
	0.0007b
	0.7474b
	206 (89.6)

	   E5.0 Blastulation rate (n, %)
	108 (97.3)
	43 (95.6)
	28 (93.3)
	15 (100.0)
	30 (100.0)
	15 (100.0)
	15 (100.0)
	0.9012b
	0.8264b
	239 (97.2)

	   E5.0 GQ Blastocysts rate (n, %)
	75 (69.4)
	34 (79.1)
	22 (78.6)
	12 (80.0)
	28 (93.3)
	13 (86.7)
	13 (86.7)
	0.1504b
	0.1154b
	184 (77.0)

	   E5.0 Hatched Blastocysts rate (n, %)
	63 (56.8)
	22 (48.9)
	15 (50.0)
	5 (33.3)
	18 (60.0)
	6 (40.0)
	5 (33.3)
	0.3430b
	0.4422b
	129 (52.4)

	Cell number per blastocyst
	
	
	
	
	
	
	
	
	
	

	E5.0 stained blastocysts (n)
	103
	39
	14
	15
	29
	15
	14
	0.1803c
	0.1803c
	215

	   Mean (SD)
	167 (40)
	164 (36)
	161 (20)
	168 (46)
	168 (27)
	187 (20)
	122 (34)
	
	
	167.6 (36.2)

	   Min; Max
	[25;250]
	[63;223]
	[120;189]
	[41;229]
	[94;221]
	[160;230]
	[36;163]
	
	
	[25;250]


This table presents the comparison of the embryological parameters among the 7 lots of CSCM-NXC (Irvine Scientific, Irvine, California, USA) used as controls for the MEAs. 
[bookmark: _Hlk189821722]Comparisons were performed using Fisher’s exact test a, the Likelihood Ratio b, and the Kruskal-Wallis test c
*P-value calculated excluding Lot (10)0000018848, MEA : Mouse Embryo Assay; GQ: Good quality
B. Comparison of the 5 lots of EmbryoMax® KSOM media used to test disposable devices and their associations
	
	EmbryoMax® KSOM
Lot 3824971
n=15
	EmbryoMax® KSOM
Lot 3925106
n=27
	EmbryoMax® KSOM
Lot 3925107
n=45
	EmbryoMax® KSOM
Lot 3989108
Pool
n=30
	EmbryoMax® KSOM
Lot 3925108
Control C
n=15
	EmbryoMax® KSOM
Lot 3925108
Control Q
n=15
	P-Value
	P-Value*
	Pool of 4 lots of EmbryoMax® KSOM except controls C and Q
n=117

	Embryological rates
	
	
	
	
	
	
	
	
	

	   E1.5 Cleavage rate (n, %)
	15 (100.0)
	27 (100.0)
	45 (100.0)
	29 (96.7)
	15 (100.0)
	15 (100.0)
	0.9990a
	0.9932 a
	116 (99.1)

	   E4.0 Blastulation rate (n, %)
	15 (100.0)
	27 (100.0)
	44 (97.8)
	26 (86.7)
	15 (100.0)
	15 (100.0)
	0.6619 a
	0.5474 a
	112 (95.7)

	   E4.0 GQ Blastocysts rate (n, %)
	14 (93.3)
	26 (96.3)
	39 (88.6)
	21 (80.8)
	13 (86.7)
	10 (66.7)
	0.4147 a
	0.6165 a
	100 (89.3)

	   E5.0 Blastulation rate (n, %)
	15 (100.0)
	27 (100.0)
	44 (97.8)
	26 (86.7)
	15 (100.0)
	15 (100.0)
	0.6619 a
	0.5474 a
	112 (95.7)

	   E5.0 GQ Blastocysts rate (n, %)
	11 (73.3)
	24 (88.9)
	39 (88.6)
	22 (84.6)
	14 (93.3)
	13 (86.7)
	0.9642 a
	0.9367 a
	96 (85.7)

	   E5.0 Hatched Blastocysts rate (n, %)
	6 (40.0)
	12 (44.4)
	23 (51.1)
	15 (50.0)
	4 (26.7)
	10 (66.7)
	0.7000 a
	0.9725 a
	56 (47.9)

	Cell number per blastocyst
	
	
	
	
	
	
	
	
	

	E5.0 stained blastocysts (n)
	15
	27
	44
	12
	15
	15
	<.0001b
	0.1570b
	98

	   Mean (SD)
	165 (29)
	175 (26)
	181 (38)
	152 (43)
	119 (14)
	145 (38)
	
	
	174 (35)

	   Min;Max
	[95;199]
	[106;231]
	[70;257]
	[66;228]
	[93;137]
	[73;186]
	
	
	[66;257]


This table presents a comparison of the embryological parameters among the different lots of EmbryoMax® KSOM (Millipore, Billerica, Massachusetts, USA) used as controls for the MEAs. 
Comparisons were performed using the Likelihood Ratioa and the ANOVAb.
*P-value calculated excluding controls C and Q, MEA : Mouse Embryo Assay; GQ: Good quality
