Additional material

Definition of exertion using accelerometer data

Exertion was quantified using data from an accelerometer-integrated pulse oximeter. To capture body movement intensity, we calculated the "jerk," defined as the rate of change of acceleration over time. Jerk reflects how quickly acceleration changes along the three orthogonal axes (x, y, z), and thus serves as a sensitive marker of movement. The absolute value of jerk, representing the rate of change in three-axis acceleration within a ±4g range, was quantized into 256 levels (0 to 255). The jerk values were averaged over each 60-second period to obtain exertional acceleration data. An exertional period was defined as the interval beginning when the average jerk value exceeded a threshold of 40 and ending one minute after the value fell below this threshold.

Methodology for processing the modified SpO₂ waveforms

The modified SpO₂ waveform was derived from the original signal through a three-step process. In the first step, desaturation dips were identified using predefined criteria: desaturation ≥ 3%, a duration between 4 and 120 seconds, and recovery to baseline within 20 seconds (Additional Figure 2A). In the second step, the precise onset and offset of each desaturation dip were determined. As illustrated in Additional Figure 2A, a representative dip spans from Point A (onset) to Point B (offset). In the final step, the segment of the waveform between Point A and Point B was linearly interpolated to replace the desaturation dip with a straight line. The resulting modified waveform is shown in orange in Additional Figure 2B. When analysis was performed using the modified SpO₂ waveforms, the sawtooth wave pattern seen in transient sleep apnea or artifact was flattened (Additional Figures 2B and 2C).

Diagnostic accuracy of possible eNH using modified SpO2 waveform on an accelerometer-integrated pulse oximetry in patient-specific analysis

When analysis was performed using modified SpO₂ waveforms, 13 out of 36 (36.1%) patients were identified as positive for possible eNH. Out of these 13 patients, 10 were diagnosed as true positive cases of eNH, while 3 did not meet the diagnostic criteria for eNH. Two out of the 3 patients were judged to have subclinical eNH (Additional Figure 1). The sensitivity and specificity of accelerometer-integrated pulse oximetry for detecting eNH were found to be 100% (95%CI: 72.2%–100%) and 88.5% (95%CI: 71.0%–96.0%), respectively.

Additional_Figure_1.docx

Subclinical episodic nocturnal hypercapnia and desaturation event due to exertion

Notes: These are judged by overnight accelerometer-integrated pulse oximetry and transcutaneous carbon dioxide pressure monitoring. Grey arrows indicate episodic desaturation events. White arrowheads indicate subclinical episodic nocturnal hypercapnia, representing episodic desaturation events and episodic increases in transcutaneous carbon dioxide pressure (PtcCO₂, 3 mmHg ≤ PtcCO₂ < 5 mmHg) without acceleration information. When an episodic desaturation event is accompanied by acceleration information, the episodic desaturation event is judged to be due to exertion.

Abbreviations: eNH, episodic nocturnal hypercapnia PtcCO2, transcutaneous carbon dioxide pressure SpO2, oxygen saturation

Additional_Figure_2.docx

Additional_Figure_2A: Definition of the desaturation dip

According to the criteria shown, desaturation dips were defined and detected. Point A marks the beginning (onset), and Point B marks the end (offset) of the desaturation dip shown in the waveform.

Abbreviations: SpO2, oxygen saturation

Additional_Figure_2B: Comparison of SpO₂ Waveforms Derived from Three Distinct Methodologies

This figure presents a comparison of SpO₂ waveforms derived using three distinct methodologies. The X-axis represents time, while the Y-axis shows SpO₂ levels. The orange line represents the modified SpO₂ waveform using our novel approach in this study. The blue line displays the original SpO₂ waveform using a 3-second moving average. The yellow line shows the SpO₂ waveform using a 2-minute moving average. The modified SpO₂ waveform (orange line) shows a flattened waveform profile compared to the other waveforms, effectively minimizing the effects of sleep apnea and artifacts, even when present.

Abbreviations: SpO2, oxygen saturation.

Additional_Figure_2C: Comparison of SpO₂ Waveforms Derived from Three Distinct Methodologies in a case with hypoventilation

This figure presents a comparison of SpO₂ waveforms derived using three distinct methodologies. The X-axis represents time, while the Y-axis shows SpO₂ levels. The orange line represents the modified SpO₂ waveform using our novel approach in this study. The blue line displays the original SpO₂ waveform using a 3-second moving average. The yellow line shows the SpO₂ waveform using a 2-minute moving average. The modified SpO₂ waveform effectively reduces the influence of sleep apnea and provides a clinically relevant representation of SpO₂ changes, especially in cases of hypoventilation.

Additional_Figure_3.docx

Additional Figure 3A: Proportion of exertional events and non-exertional events in the desaturation event-specific analysis.

Abbreviations: eNH: episodic nocturnal hypercapnia

Additional_Figure_3B: Proportion of exertional events and non-exertional events in the patient-specific analysis.

Abbreviations: eNH: episodic nocturnal hypercapnia

Additional_Figure_4: A case where original SpO2 waveform suggests possible eNH, while modified SpO2 excludes the diagnosis. (A) Using the original SpO2 waveform, the criteria for possible eNH are met, but the criteria for eNH are not met. (B)Using the modified SpO2 waveform, neither the criteria for possible eNH nor the criteria for eNH are met. Abbreviations: eNH, episodic nocturnal hypercapnia PtcCO2, transcutaneous carbon dioxide pressure SpO2, oxygen saturation