# Supplementary Information

Fast and noninvasive electronic nose for sniffing out COVID-19 based on exhaled breath-print recognition

Dian Kesumapramudya Nurputra1,2,*, Ahmad Kusumaatmaja3, Mohamad Saifudin Hakim4, Shidiq Nur Hidayat3,5, Trisna Julian3,5, Budi Sumanto3, Yodi Mahendradhata6,7, Antonia Morita Iswari Saktiawati7,8, Hutomo Suryo Wasisto5,*, Kuwat Triyana3,*

1 	Department of Child Health, Faculty of Medicine, Public Health and Nursing, Universitas Gadjah Mada, Jl. Farmako Sekip Utara, Yogyakarta 55281, Indonesia

2	Postgraduate Program in Clinical Medicine Science, Faculty of Medicine, Public Health and Nursing, Universitas Gadjah Mada, Jl. Farmako Sekip Utara, Yogyakarta 55281, Indonesia

3	Department of Physics, Faculty of Mathematics and Natural Sciences, Universitas Gadjah Mada, Sekip Utara PO Box BLS 21, Yogyakarta 55281, Indonesia

4	Department of Microbiology, Faculty of Medicine, Public Health and Nursing, Universitas Gadjah Mada, Jl. Farmako Sekip Utara, Yogyakarta 55281, Indonesia

5	PT Nanosense Instrument Indonesia, Umbulharjo, Yogyakarta 55167, Indonesia

6	Department of Health Policy and Management, Faculty of Medicine, Public Health and Nursing, Universitas Gadjah Mada, Yogyakarta, Indonesia

7	Center for Tropical Medicine, Faculty of Medicine, Public Health and Nursing, Universitas Gadjah Mada, Jl. Farmako Sekip Utara, Yogyakarta 55281, Indonesia

8 	Department of Internal Medicine, Faculty of Medicine, Public Health and Nursing, Universitas Gadjah Mada, Jl. Farmako Sekip Utara, Yogyakarta 55281, Indonesia

*	Corresponding authors.

E-mails: dian.k.nurputra@ugm.ac.id (D.K.N.); h.wasisto@nanosense-id.com (H.S.W.); triyana@ugm.ac.id (K.T.)

Contents

1.	Pre-conditioning of GeNose C19 to ambient air	S-2

2.	GeNose C19 assessment toward acetone vapors	S-3

3.	Effects of ambient temperature and humidity on sensor array	S-4

4.	Principal component analysis (PCA)	S-6

5.	Exhaled breath data processing steps	S-7

6.	Non-parametric ANOVA test	S-8

7.	Gas chromatography-mass spectroscopy (GC-MS)	S-9

# Pre-conditioning of GeNose C19 to ambient air

Figure S1 Pre-conditioning of GeNose C19. a Real-time sensor array responses during preconditioning process where all the sensors are exposed to ambient air. b Measured temperature and humidity levels inside the gas chamber. All the sensors reach their stable conditions after being pre-heated for ~20 min. Each sensor possesses an internal heater that enables an activation temperature at a few hundred degrees Celsius.

# GeNose C19 assessment toward acetone vapors

Figure S2 Exposure assessment of GeNose C19 to acetone vapors. a Gas measurement setup for GeNose C19 assessment toward specific volatile organic compound (VOC). Acetone vapors were used as a VOC model, in which they were injected to the gas chamber using a 1 μL microsyringe. b Responses of ten different metal oxide semiconductor sensors (S1-S10) integrated in GeNose 19 system during their exposure to acetone vapors with a concentration of 0.10 µL. c Comparison of sensor signal readout between GeNose C19 and calibrated digital voltmeter. The extracted response features are defined as F1 – F10. d Sensor responses during exposure to acetone vapors with different concentrations (0.04 – 0.10 µL). Each vapor concentration was tested ten times to acquire quantitative values. e Classification of four different acetone vapor concentrations employing linear discriminant analysis (LDA).

# Effects of ambient temperature and humidity on sensor array

Figure S3 Cross-temperature test of GeNose C19. a-j Responses of ten different gas sensors in GeNose C19 (S1 – S10) during cross-temperature tests. The temperature was modified in real time, while the humidity was kept at relative stable values. k Sensor array sensor signals during the first temperature test. l Temperature and humidity measured during cross-sensitivity assessment. From three repeated measurements, the sensors yield similar responses indicating their high sensing reproducibility.

Figure S4 Cross-humidity test of GeNose C19. a-j Responses of ten different gas sensors in GeNose C19 (S1 – S10) during cross-humidity tests. The humidity was modified in real time, while the temperature was kept at relative stable values. k Sensor array sensor signals during the first humidity test. l Temperature and humidity measured during cross-sensitivity assessment. From three repeated measurements, the sensors yielded similar responses indicating their reliable sensing results.

# Principal component analysis (PCA)

Figure S5 Principal component analysis (PCA) score plots of RT-qPCR-confirmed positive (red) and negative (blue) COVID-19. a PC1 vs PC2, b PC2 vs PC3, c PC2 vs PC3, and d eigenvalues of 15 components.

# Exhaled breath data processing steps

Figure S6 Block diagram of machine learning-based data processing steps used in GeNose C19 system. Four different models (i.e., linear discriminant analysis (LDA), support vector machine (SVM), multilayer perceptron (MLP), and deep neural network (DNN)) were investigated to differentiate and classify the exhaled breath data.

# Non-parametric ANOVA test

Table S1 Non-parametric ANOVA test (Kruskal-Wallis test) of sensor responses towards exhaled breaths of positive and negative COVID-19 patients. Four features (i.e., maximum, median, standard deviation, and variance) were extracted from responses of ten sensors (S1 – S10).

| No. | Feature | Sensor | Positive |  | Negative |  | Statistic value | p-value | H0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Mean | Std | Mean | Std |  |  |  |
| 1 | Maximum | S1* | 343.89 | 316.86 | 219.69 | 226.95 | 11.213 | 0.001 | Reject |
| 2 |  | S2 | 427.29 | 371.71 | 327.35 | 267.41 | 3.743 | 0.053 | Not-Reject |
| 3 |  | S3 | 146.14 | 204.64 | 142.49 | 145.08 | 2.874 | 0.09 | Not-Reject |
| 4 |  | S4 | 323.13 | 322.51 | 342.42 | 266.55 | 2.588 | 0.108 | Not-Reject |
| 5 |  | S5* | 304.14 | 300.57 | 402.78 | 289.35 | 21.32 | 0 | Reject |
| 6 |  | S6* | 129.91 | 156.77 | 72.23 | 93.91 | 20.304 | 0 | Reject |
| 7 |  | S7 | 265.48 | 248.51 | 224.64 | 220.64 | 1.341 | 0.247 | Not-Reject |
| 8 |  | S8 | 389.18 | 422.31 | 297.10 | 298.56 | 1.95 | 0.163 | Not-Reject |
| 9 |  | S9* | 517.61 | 473.33 | 205.67 | 241.83 | 41.745 | 0 | Reject |
| 10 |  | S10* | 215.68 | 332.79 | 53.34 | 70.28 | 35.381 | 0 | Reject |
| 11 | Median | S1* | 200.17 | 336.24 | 107.27 | 272.65 | 6.53 | 0.011 | Reject |
| 12 |  | S2 | 338.18 | 381.69 | 261.21 | 299.94 | 2.086 | 0.149 | Not-Reject |
| 13 |  | S3* | 38.24 | 195.17 | 56.38 | 194.44 | 8.202 | 0.004 | Reject |
| 14 |  | S4 | 243.63 | 267.12 | 261.70 | 226.76 | 2.494 | 0.114 | Not-Reject |
| 15 |  | S5* | 210.63 | 325.26 | 320.94 | 341.65 | 25.197 | 0 | Reject |
| 16 |  | S6 | 47.26 | 168.42 | 21.93 | 111.48 | 3.745 | 0.053 | Not-Reject |
| 17 |  | S7 | 143.54 | 250.47 | 107.76 | 287.07 | 0.225 | 0.635 | Not-Reject |
| 18 |  | S8 | 274.07 | 405.06 | 150.57 | 373.71 | 3.399 | 0.065 | Not-Reject |
| 19 |  | S9* | 367.83 | 480.54 | 106.08 | 286.46 | 23.252 | 0 | Reject |
| 20 |  | S10* | 147.77 | 293.67 | 24.54 | 77.82 | 15.207 | 0 | Reject |
| 21 | Standard deviation | S1* | 118.77 | 80.99 | 94.00 | 79.20 | 12.953 | 0 | Reject |
| 22 |  | S2 | 134.68 | 100.11 | 108.91 | 73.92 | 3.347 | 0.067 | Not-Reject |
| 23 |  | S3* | 47.46 | 50.34 | 51.87 | 50.46 | 7.259 | 0.007 | Reject |
| 24 |  | S4 | 93.36 | 84.73 | 107.23 | 86.28 | 2.88 | 0.09 | Not-Reject |
| 25 |  | S5* | 88.00 | 75.43 | 123.26 | 78.72 | 34.309 | 0 | Reject |
| 26 |  | S6* | 43.43 | 41.58 | 31.95 | 35.14 | 10.969 | 0.001 | Reject |
| 27 |  | S7 | 85.00 | 62.98 | 87.46 | 70.23 | 0.051 | 0.821 | Not-Reject |
| 28 |  | S8 | 111.67 | 101.97 | 114.84 | 81.48 | 3.68 | 0.055 | Not-Reject |
| 29 |  | S9* | 150.36 | 116.12 | 75.42 | 70.88 | 52.675 | 0 | Reject |
| 30 |  | S10* | 59.10 | 81.15 | 19.00 | 17.06 | 39.02 | 0 | Reject |
| 31 | Variance | S1* | 20665.89 | 26613.63 | 15108.07 | 28209.00 | 12.953 | 0 | Reject |
| 32 |  | S2 | 28160.64 | 34178.64 | 17325.57 | 21990.55 | 3.347 | 0.067 | Not-Reject |
| 33 |  | S3* | 4785.78 | 10405.56 | 5236.81 | 15701.41 | 7.259 | 0.007 | Reject |
| 34 |  | S4 | 15895.43 | 23890.76 | 18942.55 | 26418.83 | 2.88 | 0.09 | Not-Reject |
| 35 |  | S5* | 13433.46 | 23012.07 | 21389.87 | 25004.13 | 34.309 | 0 | Reject |
| 36 |  | S6* | 3615.11 | 8594.24 | 2255.90 | 9220.30 | 10.969 | 0.001 | Reject |
| 37 |  | S7 | 11190.39 | 15020.13 | 12581.33 | 22580.10 | 0.051 | 0.821 | Not-Reject |
| 38 |  | S8 | 22867.84 | 38770.38 | 19826.55 | 28451.64 | 3.68 | 0.055 | Not-Reject |
| 39 |  | S9* | 36092.04 | 46977.83 | 10712.47 | 23274.98 | 52.675 | 0 | Reject |
| 40 |  | S10* | 10077.74 | 30541.28 | 651.94 | 1479.80 | 39.02 | 0 | Reject |

# Gas chromatography-mass spectroscopy (GC-MS)

Table S2 Volatile organic compounds (VOCs) in the exhaled breaths of positive and negative COVID-19 patients measured by gas chromatography-mass spectroscopy (GC-MS). Two sampled positive COVID-19 patients are indicated with N and S. Four investigated negative COVID-19 patients are written as R, K, NS, and SK.

|  | Positive COVID-19 patients |  |  |  |  |  | Negative COVID-19 patients |  |  |  |  |  |  |  |  |  | Molecular formu-la | Volatile Catego-ry | Detec-ted by Sen-sor |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| VOC Substances | N1 | N2 | N3 | S1 | S2 | S3 | R1 | R2 | R3 | K1 | K2 | K3 | NS1 | NS2 | SK1 | SK2 |  |  |  |
| Hydrogen cyanide |  |  | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | HCN | acid | S2, S6, S9 |
| Nitrogen | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | N2 |  |  |
| Carbon monoxide | √ | √ | √ | √ | √ |  | √ | √ |  | √ |  | √ |  |  | √ |  | CO |  | S1, S3, S4, S5, S6, S8 |
| Nickel tetracarbonyl | √ |  |  | √ |  | √ |  |  | √ |  |  | √ |  | √ | √ | √ | C4NiO4 | carbon | S1, S3, S4, S5, S5, S6, S8 |
| butanedioic acid, 2-cyano-2,3-dimethyl-, diethyl ester |  | √ |  |  |  |  | √ |  |  |  |  |  |  |  |  |  | C11H17NO4 | Ester | S2, S9 |
| Manganese, acetylpentacarbonyl-, (OC-6-21)- | √ |  |  |  |  |  |  |  |  | √ |  |  |  |  |  |  | C7H3MnO6- | Furan |  |
| Ethylene |  |  |  | √ | √ |  |  |  |  |  | √ |  |  |  | √ | √ | C2H4 | Hydrocarbon | S10 |
| Carbon monoxoide | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | CO |  | S1, S3, S4, S5, S6, S8 |
| Nitrogen | √ | √ | √ |  |  |  |  | √ |  |  |  |  |  |  |  |  | N2 |  |  |
| Ethylene | √ | √ | √ |  |  |  |  |  |  |  |  | √ |  |  |  |  | C2H4 | Hydrocarbon | S10 |
| butanedioic acid, 2-cyano-2,3-dimethyl-, diethyl ester |  |  |  |  |  |  |  |  | √ | √ |  |  |  |  |  |  | C11H17NO4 | Ester | S2, S9 |
| Acetic acid, ethoxyhydroxy-, ethyl ester |  |  | √ |  |  |  |  |  |  | √ |  | √ |  |  |  |  | C10H18O3 | Ester | S2, S9 |
| Ethanol | √ | √ | √ |  |  |  | √ | √ | √ | √ | √ | √ |  |  |  |  | C2H6O | Alcohol | S1-S10 |
| Butyl aldoxime, 3-methyl-, syn- |  |  |  |  |  |  |  |  |  |  |  |  | √ | √ | √ | √ | C5H11NO | Hydrocarbon | S10 |
| Acetic acid, dimethoxy-, methyl ester |  |  |  |  |  |  |  |  |  |  | √ |  |  |  |  |  | C5H10O4 | Ester | S2, S9 |
| Cyclohexan-1,4,5-triol-3-one-1-carboxylic acid |  |  |  |  |  |  |  |  |  |  |  |  |  | √ |  |  | C7H10O6 | acid | S2, S9 |
| Acetic acid, anhydride with formic acid |  |  |  |  |  |  |  |  |  |  |  |  |  |  | √ |  | C10H18O3 | Ester | S2, S9 |
| Pyridine, 2-nitro- |  |  |  |  |  |  |  |  |  |  |  |  |  | √ |  |  | C5H4N2O2 | organic | S8,S9,S10 |
| Tetrahydropyrrole-3-ol-5-carboxylic acid, 1-acetyl-, methyl ester |  |  |  |  |  |  |  |  |  |  |  |  | √ |  |  |  | C10H14O3 | Ester | S2,S9 |
| cis-3-Butyl-4-vinyl-cyclopentene |  |  |  |  |  |  |  |  |  |  |  |  |  | √ |  |  | C11H18 | Hydrocarbon | S10 |
| Bicyclo[2.2.1]heptane, 2,2-dimethyl-3-methylene-, (1S)- |  |  |  |  |  |  | √ | √ | √ | √ |  | √ | √ | √ | √ | √ | C10H16 | Hydrocarbon | S10 |
| 2,4-Heptadien-6-ynal, (E,E)- |  |  |  |  |  |  |  |  |  |  |  |  |  | √ |  |  | C7H6O | Hydrocarbon | S10 |
| Decane, 2,6,8-trimethyl- |  |  |  |  |  |  |  |  |  |  |  |  | √ | √ | √ | √ | C13H28 | Hydrocarbon | S10 |
| Heptane, 2,2-dimethyl- |  |  |  |  |  |  | √ |  |  |  |  |  |  |  |  |  | C10H16 | Hydrocarbon | S10 |
| Oxalic acid, bis(isobutyl) ester |  |  |  |  |  |  |  |  | √ |  |  |  |  |  |  |  | C10H18O4 | Ester | S2,S9 |
| Cyclotetrasiloxane, octamethyl- |  |  |  |  |  |  |  |  |  | √ |  |  |  |  |  |  | C8H24O4Si4 | Hydrocarbon | S10 |
| Benzene, [(1-methyl-2-propenyl)oxy]- |  |  |  |  |  |  |  |  |  |  |  |  | √ | √ | √ | √ | C14H18O2 | Hydrocarbon | S10 |
| 1-Cyclopentene-1-propanoic acid, 2-[3,7-bis[(trimethylsilyl)oxy]-1-octenyl]-5-oxo-, methyl ester | √ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | C23H42O5Si2 | Ester | S2,S9 |
| Benzoic acid, hydrazide |  |  |  |  |  |  |  |  |  |  |  |  | √ | √ | √ | √ | C7H8N2O | Ester | S2,S9 |
| 3-Hydroxymandelic acid, ethyl ester, di-TMS |  |  |  |  |  |  |  |  |  |  |  |  | √ | √ |  |  | C16H28O4Si2 | Ester | S2,S9 |
| 2,5-Dihydroxybenzaldehyde, 2TMS derivative | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ |  |  |  |  | C13H22O3Si2 | Aldehyde | S2 |
| Isobornyl thiocyanoacetate | √ |  |  |  |  |  |  |  |  |  |  |  |  |  | √ |  | C13H19NO2S | Hydrocarbon | S10 |
| 1-Nonadecanol, TMS derivative | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ |  |  |  |  | C22H48OSi | Alcohol | S1-S10 |
| 3,7,11,15-tetramethylhexadecan-1,2,3-triol silylated |  |  |  |  |  |  |  |  |  |  |  |  | √ |  |  |  | C20H40O | Alcohol | S1-S10 |
| Silane, dimethyl |  |  | √ |  |  | √ |  |  |  | √ |  |  |  |  |  |  | C22H34O3Si2 | Hydrocarbon | S10 |
| Bicyclo[2.2.1]heptane, 2-chloro-2,3,3-trimethyl- |  |  |  |  |  |  |  |  |  |  |  |  |  | √ |  |  | C10H17Cl | Hydrocarbon | S10 |
| Cyclohexanol, 2-methyl-3-(1-methylethenyl)-, acetate, (1a,2a,3a)- |  |  |  |  |  |  |  |  |  |  |  |  |  | √ | √ |  | C12H20O2 | acid | S2,S9 |
| 3-Isopropoxy-1,1,1,7,7,7-hexamethyl-3,5,5-tris(trimethylsiloxy)tetrasiloxane | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | C18H52O7Si7 | Hydrocarbon | S10 |
| Molybdenum, tricarbonyl-?-6-spiro[1,3-dioxolane-2,1'-benzocyclobutene) | √ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | C17H28O2 | Hydrocarbon | S10 |
| 3-Butoxy-1,1,1,7,7,7-hexamethyl-3,5,5-tris(trimethylsiloxy)tetrasiloxane |  |  |  |  |  |  | √ | √ | √ | √ | √ | √ |  |  | √ | √ | C17H50O7Si7 | Hydrocarbon | S10 |
| 1-[2,4-Bis(trimethylsiloxy)phenyl]-2-[(4-trimethylsiloxy)phenyl]propan-1-one |  |  |  |  |  |  |  |  |  | √ | √ | √ |  |  |  |  | C24H38O4Si3 | Hydrocarbon | S10 |
| Phenol, 2-benzyloxy-3,6-difluoro- |  |  |  |  |  |  | √ |  |  | √ |  |  |  |  |  |  | C13H12O2 | Alcohol | S1-S10 |