Tabel 1. Results of Condom Catheter Testing

| Aspects of testing | Type of Condoms |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Brand A | Brand B | Brand C | Brand D | Brand E |
| Specifications Material Length (mm) Wide (mm) Texture Color | Natural latex 178 52 Nontextured Noncolored | Natural latex 190 50 Nontextured Noncolored | Natural latex 178 49 Nontextured Noncolored | Natural latex 178 53 Ribbed Colored | Natural latex 178 52 Dotted Noncolored |
| Shape |  |  |  |  |  |
| (a) Tying at 3 cm from the tip (standard knot site) 500 mL saline water | Rounded at the base | Rounded at the base | Rounded at the base | Rounded at the base | Rounded at the base |
| 1000 mL saline water |  |  |  |  |  |
| 3000 mL saline water | Elongated | Elongated | Elongated | Elongated | Elongated |
| (b) Tying in the middle of the condom (proposed knot site) 500 mL saline water | Rounded at almost all part | Rounded at almost all part | Rounded at almost all part | Rounded at almost all part | Rounded at almost all part |
| 1000 mL saline water |  |  |  |  |  |
| 3000 mL saline water | Widened | Widened | Widened | Widened | Widened |
| Maximum capacity For both knot site type | Up to 5000 mL | Up to 5000 mL | Up to 5000 mL | Up to 5000 mL | Up to 5000 mL |
| Leakage (In 2 × 24 h) For all types of knot site and volume | None | None | None | None | None |

Table 2. Flow Velocity based on Type of Catheter

| Type number and diameter (mm) | Mean of flow velocity (min) | P-value* (95%CI) | P-value** (95%CI) |
| --- | --- | --- | --- |
| 14F (4.7) | 5.40 | 0.001 (3.60-4.07) |  |
| 16F (5.3) | 4.36 |  |  |
| 18F (6.0) | 3.57 |  | 0.111(3.31-3.52) |
| 20F (6.7) | 3.56 |  |  |
| 22F (7.3) | 3.46 |  |  |
| 24F (8.0) | 3.20 |  |  |
| No catheter | 3.30 |  |  |

*ANOVA test for all catheter types and no catheter

**ANOVA test for catheter nos 18F, 20F, 22F, and 2F4 and no catheter