Table 2. Advantages and disadvantages of methods of decontamination

| Method | Advantages | Disadvantages | Notes |
| --- | --- | --- | --- |
| Gaseous |  |  |  |
| HPGP | Good biocidal efficacy against MS2, Bacteriophage Phi6, influenza virus H3N, and V. stomatitis Did not affect the filtration performance of masks N95, SN95, models 1860, 1860S, 8210, 9210. | Strongly affected the particle penetration (%) of N95 and SN93 masks (55 minutes) Caused leakage (5 cycles) | No experiments showing the physical appearance of masks after the treatment. |
| VHP | Good biocidal efficacy against Geobacillus stearothermophius, T1 Bacteriophage, T7 Bacteriophage, Pseudomonas phage phi-6, V. stomatitis, MHV, A. niger, Geobaccilus spores and SARS-CoV-2 Did not affect fit factors (AO Safety N9504C and models 1860, 1860S, Aura 1870+, 8511) Did not affect the filtration performance of masks N95 and SN95. | Leading the stiffness of straps, less shine and tarnish of nosebands. Exceeding 20 cycles leading to the stiffer straps Still remained insignificant oxidant residuals after treatment. possibly high cost |  |
| EtO (725 - 883 mg/L, 1 – 3 cycles, 1 hour) | Good biocidal efficacy against V. stomatitis Did not affect the filtration performance, physical appearance of masks N95 and SN95. No chemical residuals. fit test passed after 1 - 3 cycles | Traces of contaminants | EtO might be harmful to the human health |
| Ozone gas | Good biocidal efficacy against HCoV-22E, H1N1, S. aureus Did not affect the filtration performance of fabric materials, physical appearance, and the mechanical integrity |  |  |
| Vapor ethanol | Did not affect the filtration performance of model 1860S and 8210 |  | Degraded the filtration performance of model 9210 |
| Heat and humidity |  |  |  |
| MHI | Good biocidal efficacy against H1N1, H5N1, MS2, phi6, H3N2, MHV, and S. aureus Did not affect the filtration performance of fabrics and masks N95 and SN95, Chen Heng V9501 KN95, HKYQ Fit factors ≥ 100 for most models | Failed to kill G. stearothermophilus High temperature (125oC) damaged the filtration performance A detachment of the inner foam nose cushion on varied models. Stronger odour for N95 model 1860 | Fit factors of Chen Heng V9501 KN95, and HKYQ N95 were strongly decreased |
| Autoclave | Good biocidal efficacy against B. subtilis, V. stomatitis and SARS-CoV-2 No remarkable visual changes were observed in the P100 respirators, though the respirator media itself felt softer. | Strongly decreased the filtration performance of most models Deformed models Safe Worker 1016, 1860S, 8210 Fit test failed (N95 models 1860 and 8210) Degraded appearances of N95 masks | The filtration performance of masks N95 were varied and inconsistent across studies. Aura 1862, 8210, 1860, SH-2950, 9210 still had > 95% of filtration efficacy. Still effectively retained particles 0.5 – 1.0 µm. |
| Dry heat (70 - 100oC, 15 - 60 min) | Slow clearance of SARS-CoV-2 strain Not effective against MS2 and MRSA Did not affect the filtration performance of fabrics and masks N95, P100. No visible changes in N95, surgical masks, cotton masks, quilted masks (at 100oC, 15 minutes) Preserve fit factors of varied models | Poor efficacy against MS2, phi6, MRSA, S. aureus and G. stearothermophilus | Reduce fit factor of AOSafety N9504C respirators |
| MGS (2 - 75 min) | Good effect against H5N1, H1N1 Good effect against MS2 if treatment duration > 45 min. Did not affect the filtration performance of masks N95, SN95 and P100. Fit test passed for all tested masks. No visible physical characteristics on N95 model 1860 | A detachment of the inner foam nose cushion on model 1870, sometimes caused melting. | All components of masks should completely be exposed the steam |
| Microwave steam bag | Good effect against MS2 Not significantly change the filtration performance. | Water retention |  |
| Rice cooker generated steam | Good effect against MRSA, MS2 Did not cause physical changes |  |  |
| Boiling water generated steam | Good effect against H120 | Degraded the filtration performances of fabric and Venus 4420 mask | Recharge after treating could recover the filtration efficacy |
| Rice cooker (149-164˚C, 3 min) | Good effect against B. subtilis. Did not affect the filtration performance of masks N95. | Strongly reduced the filtration performance of masks Gauze, Spunlace. | No experiments showing the physical appearances after treatment. |
| Liquid/chemicals |  |  |  |
| Hydrogen peroxide | Aerosolized hydrogen peroxide inactivated HSV-1, CVB3 and phi6 Liquid hydrogen peroxide did not affect the filtration performance of masks N95, SN95 and P100. Fit test passed (model 8511) | Still remained oxidant residuals after treatment. | No experiments showing the biocidal efficacy after treatment. (liquid hydrogen peroxide) |
| Dimethyl dioxirane |  | Still remained oxidant residuals after treatment. Metal oxidized and strong odour | No experiments showing the biocidal efficacy and filtration performance after treatment. |
| Ethanol (70%, 10 min) | Quite good effect against B. subtilis. Quick reductions of SARS-CoV-2 Less changes in filter air flow resistance | Strongly reduced the filtration performance of masks N95 and Gauze Decreased fit factor scores | Quick reductions of microorganisms but the maximal reduction was lower than other methods No experiments showing the physical appearances after treatment. Recharge or dried in vacuum condition after treating could recover the filtration efficacy |
| Isopropanol (100%, 10 - 30 min) |  | Strongly reduced the filtration performance of masks N95, GAUZE, Spunlace. Ink faded | No experiments showing the biocidal efficacy after treatment. |
| Mixed oxidants |  | Still remained oxidant residuals after treatment. Metal oxidized and strong odour | No experiments showing the biocidal efficacy and filtration performance after treatment. |
| Laundering | Soap and water showed no affection of physical appearance | Treatment of soap and water strongly reduced the filtration performance of masks N95and P 100 Treatment of water and detergent reduced the filtration efficacy of fabrics | No experiments showing the biocidal efficacy after treatment. Recharge after treating could recover the filtration efficacy |
| Sodium hypochlorite (0.006 – 10%, 10 -30 min) | Good biocidal effect from the concentration of 0.25% against H1N1, MS2, B. subtilis. Did not affect the filtration performance of masks P100. | Significantly reduced the filtration performance of masks N95, SN95. Strongly degraded physical appearance of N95 and SN95 Strongly damage the structure of Gauze | The filtration performance of N95 after treatment were varied and inconsistent across studies. |
| Energy |  |  |  |
| UVGI | Good biocidal effect against MS2 from the doses of 4.32 J/cm2. Good biocidal effect against H1N1, H5N1 from the doses of 1 J/cm2. Did not affect the filtration performance, physical appearances of masks N95, SN95 and P100. Fit test passed for all models | Dose of 30.32 J/cm2 reduced fit factor scores and filtration performance of fabrics Dose over 120 J/cm2 reduced the strength of masks Could cause phytochemical damages | Repeated the contamination/decontamination reduce the efficacy Low dose showed poor biocidal efficacy Soil load affected the decontamination of UVGI |
| Microwave oven irradiation (1100 W, 2 min, 1 cycle) | Did not affect the filtration performance, physical appearances of masks N95 and P100. | Some masks SN95 and P100-I melted during treatment. | No experiments showing the biocidal efficacy after treatment. |
| Gamma irradiation | Fit test passed | Reduced the filtration efficacy |  |
| Cleaning Wipes |  |  |  |
| Wipes containing disinfecting agents | Good effect against S. aureus and H1N1 by different agents (Benzalkonium chloride, sodium hypochlorite, QAC) Mostly did not affect the filtration performance, physical appearances of masks N95. | Some masks cleaning by wipes containing benzalkonium chloride had particle penetration > 5% |  |

EtO: Ethylene oxide, HPGP: Hydrogen peroxide gas plasma, HPV: Hydrogen Peroxide Vapor, MGS: Microwave-generated steam, MHI: Moist heat incubation, QAC: 2–2-pdiisobutylphenoxyethoxyethyldimethyl ammonium chloride (a quaternary ammonium chloride).