Delpuech et al., Bulk Heat Inactivation Supplementary Figures.

Supplementary Figure S1: Sample racking and trolley system used for bulk HI using in the Electrolux Skyline Combi Oven. Lower left image shows samples in UN3373 packaging placed in a tray (upper left) ready for loading onto the trolley (right).

Supplementary Figure S2: Bulk HI cycle optimisation using the Industrial Steam oven, Experiments 2 and 3. Experiment 1 is shown in Fig. 1c. Twenty thermocouple containing swab vials were placed throughout a full oven load of mock samples and temperate monitored during execution of a HI cycle. All thermocouple vials achieve the intended viral inactivation conditions of 65°C for 10 min and no vial is heated beyond approximately 77°C.

Experiment 2:

Experiment 3:

Supplementary Figure S3: Validation of the HI cycle using different vial types.

Typical OP/NP swab vials in use and the UN3373 outer bag used for containment:

Temperature data from thermocouple containing swab vials placed throughout a full oven of mock samples during execution of a HI cycle. The Red Vacutainer type showed the potential for a significantly different heating profile to other vials in the same run, although this was within the heating range of vial on the whole seen in other experiments. Notably however, several of this vial type were seen to warp after the HI cycle so were excluded from the bulk HI process using this HI cycle.

Supplementary Figure S4: Validation of the HI cycle using different OP/NP swab vials containing multiple swabs. Thermocouple-containing vials with single or 5 swabs (multi) were placed throughout a full oven of mock samples and temperate monitored during execution of a HI cycle. The number of swabs does not alter the heating of samples outside the typical range of that seen in single swab vials.

Supplementary Figure S5: Variability of SARS-CoV-2 detection. Clinical OP/NP swab samples (23) were tested over 3 independent days (E1,E2,E3). Dots represent replicate (n=2, n=3) measurements of each sample on the same day. Using an analysis of variance to estimate the pooled within-sample variability we calculated a standard deviation of 0.33 Cq. Temperature T65 is 65⁰C in a thermocycler, 65C is 65⁰C the standard oven.

Supplementary Figure S6: SARS-CoV-2 detection by RT-qPCR following incubation at 65C for 20 min,  60 min or 90 min compared to 10 min.

SARS-CoV-2 Cq (10 minutes)

SARS-CoV-2 Cq (20 minutes)

SARS-CoV-2 Cq (10 minutes)

SARS-CoV-2 Cq (60 minutes)

minutes

SARS-CoV-2 Cq (10 minutes)

SARS-CoV-2 Cq (90 minutes)

Supplementary Figure S7: Cq values from concordance studies of clinical OP/NP swab samples (n=100) using the bulk HI protocol compared to the standard approved process.

All positive clinical samples by Cq value. bulk HI protocol with sample processing withing 1 to 4 hours.

SARS-CoV-2 Cq (Standard process)

SARS-CoV-2 Cq (Bulk HI)

RT-qPCR data for selected samples using each protocol. Green circles show the standard approved process, orange starts show the bulk HI protocol with sample processing withing 1 to 4 hours, pink triangles show the bulk HI protocol with sample processing after 10 hours.

Supplementary Figure 8: Concordance data of clinical OP/NP swab samples using the bulk HI protocol compared to the standard approved process.

Test result concordance with a delay of up to 4h or 10h following bulk HI. Samples are categorised as ‘Pos to Pos’ (Green; Positive in clinical test and positive in bulk HI test) or ‘Pos to Neg’ (Red; Positive in clinical test and negative in bulk HI test), and binned according to Cq value.

Correlation in Cq values with a delay of up to 4h or 10h following bulk HI.

Supplementary Figure S9: Test result concordance with a delay of up 24h following bulk HI. Samples are categorised as ‘Pos to Pos’ (Green; Positive in clinical test and positive in bulk HI test), ‘Pos to Neg’ (Red; Positive in clinical test and negative in bulk HI test), or ‘Neg to Pos’ (Blue; Negative in clinical test and positive in bulk HI test) and binned according to Cq value.

Supplementary Figure S10: Test result concordance following the above fall-back process of employing a second heat step after sample lysis.

Samples are categorised as ‘Pos to Pos’ (Green; Positive in clinical test and positive in bulk HI test), ‘Pos to Neg’ (Red; Positive in clinical test and negative in bulk HI test), or ‘Neg to Pos’ (Blue; Negative in clinical test and positive in bulk HI test) and binned according to Cq value.

Supplementary Figure S11: Sequencing coverage was not affected by the implementation of bulk HI at the CCTC (2_after) compared to before implementation (1_before).

Supplementary Table1: The optimised oven program for bulk HI. The oven is preheated prior to sample loading, then ramping temperatures are used to rapidly heat the samples while minimising overheating. A dry step is included at the end to reduce moisture on the out bags and make them easier to handle in the lab.

| Step | Temp °C | Humidity (%) | Time (min) |
| --- | --- | --- | --- |
| Preheat | 85 | n/a |  |
| 1 | 85 | 99 | 5 |
| 2 | 75 | 99 | 13 |
| 3 | 70 | 99 | 10 |
| 4 | 70 | 0 | 2 |

Supplementary Table 2: Effect of temperature on SARS-Cov-2 detection by RT-qPCR. The mean and standard deviation (SD) of the difference in Cq value obtained using the higher temperature versus 65°C is shown.

|  | Change in Cq at 75°C vs 65°C Cq (75°C – 65°C) | Change in Cq at 80°C vs 65°C Cq (80°C – 65°C) |
| --- | --- | --- |
| Mean | 0.39 | 0.69 |
| SD | 0.4 | 0.55 |
| N | 43 | 49 |

Supplementary Table 3: Effect of incubation time at 65°C on SARS-Cov-2 detection by RT-qPCR. The mean and standard deviation (SD) of the difference in Cq value obtained using the longer time versus 10 min is shown.

|  | Change in Cq at 20 min Cq (20min-10min) | Change in Cq at 60 min Cq (60min-10min) | Change in Cq at 90 min Cq (90min-10min) |
| --- | --- | --- | --- |
| Mean | 0.01 | 0.79 | 1.10 |
| SD | 0.50 | 0.66 | 0.75 |
| N | 128 | 127 | 129 |

Supplementary Table 4: Effect of bulk HI protocols on SARS-Cov-2 detection by RT-qPCR concordance studies. The mean and standard deviation (SD) of the difference in Cq value obtained using the bulk HI protocol versus the standard approved process is shown.

|  | 4h Cq (4h-Std) | 10h Cq (10h-Std) | 24h Cq (24h-Std) | Fall back HI process Cq (Fall back HI-Std) |
| --- | --- | --- | --- | --- |
| Mean | 0.81 | 0.19 | -0.60 | 0.19 |
| SD | 1.08 | 0.84 | 0.85 | 0.85 |
| N | 100 | 54 | 19 | 54 |

Supplementary Table 5: Bulk HI concordance data summary.