## Common
TEST METHOD VALIDATION AND VERIFICATION - NONWAIVED TESTS
* This is a summary of common CAP requirements for the Verification/Validation of new Method/Tests * Verification/Validation of new laboratory test requires identifying the intended use of the test and establishing demonstrated evidence that provides a high degree of assurance that a test will consistently perform as expected
CAP Requiremenet | Analytical Verification/Validation | Test | Eligibility | Criteria of Verification/ Validation | Definition | Method | Qualitative | Quantitative
COM.40300 | Analytical Verification | FDA Approved | FDA Approved | Analytical accuracy | Closeness of the agreement between the results of a single measurement and the true value of the analyte | 1- Accuracy is verified by comparing results to a definitive or reference method, or an established comparative method. 2- Use of matrix-appropriate reference materials, patient specimens (altered or unaltered), or other commutable materials with known concentrations or activities may be used to verify accuracy. 3- The Use of routine quality control materials or calibrators used to calibrate The method is not appropriate. 4- Ideally, testing should be spread over a minimum of 5 days. | (TP+TN)/[TP+FP+TN+FN]*100 | Slope Y intercepts Correlation Coefficient ( r )
|  |  |  | Analytical precision | Repeatability and reproducibility of results | Precision is verified by repeat measurement of samples at varying concentrations/activities within run and between run over a period of time. | TP /(TP + FP)*100 | %CV Overall % CV
|  |  |  | Repeatability | * The closeness of agreement between results of tests performed on the same test items, by the same analyst, on the same instrument, under the same conditions in the same location and repeated over a short period of time. * It represents ‘within-run precision’. | At least six (6) measurements of the same control sample, divided into aliquots or a pooled sample, divided into aliquots. The measurements should cover the entire duration of the daily operation of the laboratory. A 15 min interval at least, is required between the six measurements. | Comparable results | Mean SD %CV Overall Mean Overall SD Overall %CV
|  |  |  | Rreproducibility | * The closeness of agreement between results of tests carried out on the same test items, taking into account the broadest range of variables encountered in real laboratory conditions, including different laboratories. * It represents ‘inter-laboratory precision’.2 | At least six (6) measurements of the same control sample under different conditions, e.g. different days and using at least two different batches of the analyte reagents. The results of the internal quality control charts (QC) may also be used. | Comparable results | CVR%
|  |  |  | Reportable range [measuring interval, analytical measurement range (AMR), linear range] | The range of test result values over which the laboratory has verified accuracy of the instrument or test system measurement response. The boundaries of the reportable range are the lowest and highest analyte concentrations that generate results that are reliably produced by a test method without dilution of the specimen. The lower limit of linearity is frequently referred to as the lower limit of quantification (LLOQ) and the upper limit of linearity as the upper limit of quantification (ULOQ). | By testing a series of samples of known concentrations or a series of known dilutions of a highly elevated patient specimen or standard with concentrations across the anticipated measuring range. Equally spaced intermediate concentrations are recommended but not strictly required (8).The measured values are compared to the assigned values, typically by plotting the measured values on the y axis and the assigned values on the x axis. | NA | Regression Coefficient, R²
COM.40605 |  |  |  | Reference Intervals | Reference intervals are important to allow a clinician to assess patient results against an appropriate population. The reference intervals must be established or verified for each analyte and specimen source (eg, blood, urine, cerebrospinal fluid), when appropriate
COM.40325 | Analytical Verification | Internationally recognized regulatory authority | Eg. the European Union's Conformité Européenne (CE) Marking | Analytical accuracy | Same as above
|  |  |  | Analytical precision | Same as above
|  |  |  | Reportable range | Same as above
COM.40605 |  |  |  | Renference Intervals | Same as above
|  |  |  | Any other performance characteristic required to ensure analytical test performance | Examples; specimen stability, reagent stability, linearity, carryover, and cross-contamination, as appropriate and applicable. | [Sheet 2] | [Sheet 2] | [Sheet 2]
COM.40350 MOL.30785 MOL.31130 | Analytical Validation | Modified FDA-cleared/approved Tests and LDTs | Also Applies to : 1- Tests that are not approved by an internationally recognized regulatory authority. 2- Approved tests that have been modified by the laboratory. | Analytical accuracy | Same as above | 1- For quantitative tests, a minimum of 20 samples with analyte concentrations distributed across the analytical measurement range should be used. Proportionate mixtures of samples may be used to supplement the study population. 2- For qualitative tests, a minimum of 20 samples, including positive, negative, and lowpositive samples with concentrations near the lower level of detection should be used; equivocal samples should not be used. 3- For certain methods that test multiple analytes (eg, next-generation sequencing, HPLC, GC-MS, MALDI-TOF, etc.), analytic accuracy may be established for each method (not necessarily each analyte), as appropriate.
|  |  |  | Analytical precision | Same as above
MOL.31130 |  |  |  | Reportable range | Reportable range encompasses the full range of reported values: ● For qualitative tests that would include all reportable outcomes (eg, homozygous wild type, heterozygous or homozygous variant). ● For quantitative tests, the laboratory must define the analytical measurement range (AMR) as described in the Quantitative Assays; Calibration and Standards section of the checklist. | 7 to 11 concentrations across the anticipated measuring range. To ascertain the widest possible linear range, additional samples at concentrations 20 to 30% wider than the anticipated measuring range can be used. Samples should be tested in replicates of two to our, depending on the anticipated imprecision of the assay (8). Ideally, all results for a single analyte will be obtained on the same day. | NA
COM.40605 |  |  |  | Renference Intervals | Same as above
|  |  |  | Analytic sensitivity | * The lowest concentration or amount of the analyte or substance that can be measured or distinguished from a blank (lower limit of detection-LOD) * The proportion of positive results correctly identified by the test |  | TP/[TP+FN]*100 | Mean SD
COM.40500 |  |  |  | Analytic specificity | * The ability of a test or procedure to correctly identify or quantify an entity in the presence of interfering or cross-reactive substances that might be expected to be present (interferences) * The proportion of negative results correctly identified by the test. | Data on interferences may be obtained from manufacturers or published literature, as applicable | TN/[TN+FP]*100
|  |  |  | Any other performance characteristic required to ensure analytical test performance | Examples; specimen stability, reagent stability, linearity, carryover, and cross-contamination, as appropriate and applicable. [Sheet 2] | [Sheet 2] | [Sheet 2] | [Sheet 2]
Important:
* If multiple identical instruments or devices are in use, there must be records (data and written assessment) showing that the method performance specifications have been separately validated for each test and instrument or device.
* The written assessment must include an evaluation of each component of the validation study, including the acceptability of the data. If data include discordant results, there must be a record of the discordance and investigation of any impact on the approval of the test for clinical use.
* If a method is verified by someone other than the laboratory's personnel (eg, manufacturer's representative), the laboratory must have records to show that the verification correlates with in-house test performance by showing confirmation of performance specifications by the laboratory personnel testing known specimens.

## Other characteristic
TEST METHOD VALIDATION AND VERIFICATION - NONWAIVED TESTS
CAP Requiremenet | Criteria of Verification/ Validation | Definition | Method | Qualitative | Quantitative
MOL.31251 | Cut-Off | The discrete amount of analyte carried by measuring system from one sample reaction into subsequent sample reactions, thereby erroneously affecting the apparent amounts in subsequent samples. | The threshold value that distinguishes a positive from a negative result
| Truness/Bais | An estimate of a systematic measurement error. | Assessment of trueness is achieved by using control samples provided commercially, reference materials, results from interlaboratory schemes and recovery. | At least six (6) negative and six (6) positive samples with a known value should be analysed . Trueness is calculated from the ratio [true positive + true negative] / [total number of samples]. In case the kit contains both a negative and a positive control sample, these could be analyzed in duplicate for six (6) days. | is calculated by recovery experiments (mixing with control sample or calibrator with different concentrations) and measuring samples of known concentration six (6) times at least, within the same day. The %recovery (R%) and the %error (bias) are calculated, respectively. It is supplemented and monitored by participation in proficiency testing schemes.
| Diagnostic/Clinical accuracy | Comparison with known clinical diagnosis
| Method comparison |  | * Daily for minimum 10 days * Minimum 10 samples per category (available result) * For 2 categories minimum total of 30 samples | Agreement between data ( coefficient)
| Uncertinity |  |  | NA
Carryover/Cross-contamination

## New Test-Method Info
TEST METHOD VALIDATION AND VERIFICATION - NONWAIVED TESTS
Section: | MDI
New Test/Method Name: | CMV_REAL STAR
New Test/Method Category: | □ FDA Approved □ Internationally recognized regulatory authority □ FDA-Modified □ LDT
New Test/Method Category: | □ Qualitative □ Quantitative
Equipment ID: | BIORAD (1)
Objective of the verification/validation: | checking whether the specification of a new kit captures the customer's requirements,
Expected Results:
Reference Method: | Previous validated method (qiagen Artus kit)
Other Data:
Prepared By: | Zeena Shaheen, Senior MT.

## Specimens ID-Expected result
[5 empty] | CUT OFF:456 Copies/ml | CUT OFF: 265 IU/ML |  |  | CUT OFF : 153 Copies/ml | CUT OFF : Copies/ml
NO. | SAMPLE ID | DATE | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit Copies/ml _BIORAD | Results Real Star kit IU/ml _BIORAD | Results Real Star kit Copies/ml _ROTOR GENE | QIAGEN KIT IU/ml_ROTOR GENE REFERENCE METHOD | QIAGEN KIT Copies/ml_ROTOR GENE REFERENCE METHOD | Results Real Star kit Copies/ml _ROTOR GENE
1 | V0524099 | 06/05/2024 | PATIENT | ACCURACY | 7800 | 4534.8837 | 1050 | 2947.6744 | 5070 | 3460
2 | V0524283 | 13/05/2024 | PATIENT | ACCURACY/PRECISION | 187000 | 108720.9302 | 137000 | 76162.7907 | 131000 | 162000
3 | V0524348 | 14/05/2024 | PATIENT | ACCURACY | 21500 | 12500 | 11700 | 11918.6047 | 20500 | 16100
4 | V0524350 | 14/05/2024 | PATIENT | ACCURACY/PRECISION | 2060 | 1197.6744 | 1380 | 4994.186 | 8590 | 2646
5 | V0524514 | 20/05/2024 | PATIENT | ACCURACY | 2630 | 1529.0698 | 3350 | 2779.0698 | 4780 | 4200
6 | V0524522 | 20/05/2024 | PATIENT | ACCURACY/PRECISION | 33000 | 19186.0465 | 35600 | 26395.3488 | 45400 | 34900
7 | V0524523 | 20/05/2024 | PATIENT | ACCURACY | 4020 | 2337.2093 | 6240 | 2819.7674 | 4850 | 7750
8 | V0524524 | 20/05/2024 | PATIENT | ACCURACY | 29400 | 17093.0233 | 26500 | 19127.907 | 32900 | 29100
9 | V0524594 | 22/05/2024 | PATIENT | ACCURACY/SENSITIVITY/LOD | 56100 | 32616.2791 | 44100 | 54651.1628 | 94000 | 57500
10 | V0524604 | 22/05/2024 | PATIENT | ACCURACY | 5890 | 3424.4186 | 5450 | 4796.5116 | 8250 | 4200
11 | V0227220 | 28/05/2024 | PATIENT | ACCURACY | 2310 | 1343.0233 | 9140 | 5354.6512 | 9210 | 4861
12 |  | 26/01/2023 | CAP | ACCURACY | 3410 | 1982.5581 | 2210 | 4465.1163 | 7680 | 4480
13 |  | 31/01/2022 | CAP | ACCURACY/PRECISION | 43600 | 25348.8372 | 45800 | 14593.0233 | 25100 | 43600
14 | V0624092 | 04/06/2024 | PATIENT | ACCURACY | 3040 | 1767.4419 | 5095 | 2796.5116 | 4810 | 7230
15 | V0624108 | 04/06/2024 | PATIENT | ACCURACY/PRECISION | 5910 | 3436.0465 | 2460 | 1889.5349 | 3250 | 3509
16 | V0624153 | 05/06/2024 | PATIENT | ACCURACY | 1210 | 703.4884 | 1680 | 674.4186 | 1160 | 833
17 | V0624352 | 12/06/2024 | PATIENT | ACCURACY | 96900 | 56337.2093 | 26500 | 45406.9767 | 78100 | 95000
18 | V0624333 | 12/06/2024 | PATIENT | ACCURACY | 6020 | 3500 | 2790 | 2034.8837 | 3500 | 2050
19 | V0624330 | 12/06/2024 | PATIENT | ACCURACY | 1450 | 843.0233 | 980 | 598.8372 | 1030 | 3190
20 | V0624297 | 11/06/2024 | PATIENT | ACCURACY | 549 | 319.186 | 235 | 379.0698 | 652 | 472
21 | V0624278 | 11/06/2024 | PATIENT | ACCURACY | 5110 | 2970.9302 | 9800 | 1244.186 | 2140 | 1400
22 | V0624240 | 10/06/2024 | PATIENT | ACCURACY | 4590 | 2668.6047 | 7060 | 5005.814 | 8610 | 2240
23 | V0227301 | 30/05/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | BK POSITIVE
24 | V0227193 | 27/05/20204 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | HBV POSITIVE
25 | V0524654 | 22/05/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | HBV POSITIVE
26 | V0227175 | 27/05/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
27 | V0624174 | 05/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
28 | V0624167 | 05/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
29 | V0624127 | 03/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | HBV POSITIVE
30 | V0624105 | 04/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | ENTERO POSITIVE
31 | V0624089 | 04/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
32 |  |  | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | LIPEMIC
33 |  |  | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | ICTERIC
34 | V0624071 | 03/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
35 | V0624054 | 03/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE

## Specimens ID-Expected resul (2
[6 empty] | CUT OFF:456 Copies/ml | CUT OFF: 265 IU/ML |  |  | CUT OFF : 153 Copies/ml | CUT OFF : Copies/ml
NO. | SAMPLE NAME | DATE | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit Copies/ml _BIORAD | Results Real Star kit IU/ml _BIORAD | Results Real Star kit Copies/ml _ROTOR GENE | QIAGEN KIT IU/ml_ROTOR GENE REFERENCE METHOD | QIAGEN KIT Copies/ml_ROTOR GENE REFERENCE METHOD | Results Real Star kit Copies/ml _ROTOR GENE
1 | SAMPLE_01 | V0524099 | 06/05/2024 | PATIENT | ACCURACY | 7800 | 4534.8837 | 1050 | 2947.6744 | 5070 | 3460
2 | SAMPLE_02 | V0524283 | 13/05/2024 | PATIENT | ACCURACY/PRECISION | 187000 | 108720.9302 | 137000 | 76162.7907 | 131000 | 162000
3 | SAMPLE_03 | V0524348 | 14/05/2024 | PATIENT | ACCURACY | 21500 | 12500 | 11700 | 11918.6047 | 20500 | 16100
4 | SAMPLE_04 | V0524350 | 14/05/2024 | PATIENT | ACCURACY/PRECISION | 2060 | 1197.6744 | 1380 | 4994.186 | 8590 | 2646
5 | SAMPLE_05 | V0524514 | 20/05/2024 | PATIENT | ACCURACY | 2630 | 1529.0698 | 3350 | 2779.0698 | 4780 | 4200
6 | SAMPLE_06 | V0524522 | 20/05/2024 | PATIENT | ACCURACY/PRECISION | 33000 | 19186.0465 | 35600 | 26395.3488 | 45400 | 34900
7 | SAMPLE_07 | V0524523 | 20/05/2024 | PATIENT | ACCURACY | 4020 | 2337.2093 | 6240 | 2819.7674 | 4850 | 7750
8 | SAMPLE_08 | V0524524 | 20/05/2024 | PATIENT | ACCURACY | 29400 | 17093.0233 | 26500 | 19127.907 | 32900 | 29100
9 | SAMPLE_09 | V0524594 | 22/05/2024 | PATIENT | ACCURACY/SENSITIVITY/LOD | 56100 | 32616.2791 | 44100 | 54651.1628 | 94000 | 57500
10 | SAMPLE_10 | V0524604 | 22/05/2024 | PATIENT | ACCURACY | 5890 | 3424.4186 | 5450 | 4796.5116 | 8250 | 4200
11 | SAMPLE_11 | V0227220 | 28/05/2024 | PATIENT | ACCURACY | 2310 | 1343.0233 | 9140 | 5354.6512 | 9210 | 4861
12 | SAMPLE_12 |  | 26/01/2023 | CAP | ACCURACY | 3410 | 1982.5581 | 2210 | 4465.1163 | 7680 | 4480
13 | SAMPLE_13 |  | 31/01/2022 | CAP | ACCURACY/PRECISION | 43600 | 25348.8372 | 45800 | 14593.0233 | 25100 | 43600
14 | SAMPLE_14 | V0624092 | 04/06/2024 | PATIENT | ACCURACY | 3040 | 1767.4419 | 5095 | 2796.5116 | 4810 | 7230
15 | SAMPLE_15 | V0624108 | 04/06/2024 | PATIENT | ACCURACY/PRECISION | 5910 | 3436.0465 | 2460 | 1889.5349 | 3250 | 3509
16 | SAMPLE_16 | V0624153 | 05/06/2024 | PATIENT | ACCURACY | 1210 | 703.4884 | 1680 | 674.4186 | 1160 | 833
17 | SAMPLE_17 | V0624352 | 12/06/2024 | PATIENT | ACCURACY | 96900 | 56337.2093 | 26500 | 45406.9767 | 78100 | 95000
18 | SAMPLE_18 | V0624333 | 12/06/2024 | PATIENT | ACCURACY | 6020 | 3500 | 2790 | 2034.8837 | 3500 | 2050
19 | SAMPLE_19 | V0624330 | 12/06/2024 | PATIENT | ACCURACY | 1450 | 843.0233 | 980 | 598.8372 | 1030 | 3190
20 | SAMPLE_20 | V0624297 | 11/06/2024 | PATIENT | ACCURACY | 549 | 319.186 | 235 | 379.0698 | 652 | 472
21 | SAMPLE_21 | V0624278 | 11/06/2024 | PATIENT | ACCURACY | 5110 | 2970.9302 | 9800 | 1244.186 | 2140 | 1400
22 | SAMPLE_22 | V0624240 | 10/06/2024 | PATIENT | ACCURACY | 4590 | 2668.6047 | 7060 | 5005.814 | 8610 | 2240
23 | SAMPLE_23 | V0227301 | 30/05/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | BK POSITIVE
24 | SAMPLE_24 | V0227193 | 27/05/20204 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | HBV POSITIVE
25 | SAMPLE_25 | V0524654 | 22/05/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | HBV POSITIVE
26 | SAMPLE_26 | V0227175 | 27/05/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
27 | SAMPLE_27 | V0624174 | 05/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
28 | SAMPLE_28 | V0624167 | 05/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
29 | SAMPLE_29 | V0624127 | 03/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | HBV POSITIVE
30 | SAMPLE_30 | V0624105 | 04/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | ENTERO POSITIVE
31 | SAMPLE_31 | V0624089 | 04/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
32 | SAMPLE_32 |  |  | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | LIPEMIC
33 | SAMPLE_33 |  |  | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | ICTERIC
34 | SAMPLE_34 | V0624071 | 03/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE
35 | SAMPLE_35 | V0624054 | 03/06/2024 | PATIENT | SPECIFICITY/REFERENCE INTERVAL | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | NEGATIVE | EBV POSITIVE

## Accuracy
ACCURACY
Item | Data | SAMPLE NAME | SAMPLE ID | DATE | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit Copies/ml _BIORAD | Results Real Star kit IU/ml _BIORAD | QIAGEN KIT Copies/ml_ROTOR GENE REFERENCE METHOD | QIAGEN KIT IU/ml_ROTOR GENE REFERENCE METHOD
Number/Info of samples | 22 positive Whole Blood Samples were used for accuracy validation 14 negative Whole Blood | SAMPLE_01 | V0524099 | 06/05/2024 | PATIENT | ACCURACY | 7800 | 4534.8837 | 5070 | 2947.6744
Analytical Accuracy Study Design | A comparison of methods study By using proficiency testing samples and patients samples with known results that performed with previous validated method. | SAMPLE_02 | V0524283 | 13/05/2024 | PATIENT | ACCURACY | 187000 | 108720.9302 | 131000 | 76162.7907
Data analysis | The Realstar CMV PCR Kit was compared to the Qiagen Artus CMV. DNA was isolated using the QIAamp DSP Virus Kit, and analysis was carried out on the Biorad CFX 96 (1) Instrument. For comparative testing the results obtained by using the artus CMV RG PCR Kit were compared to those of theRealStar CMV Assay. In comparison to the results collected with the Qiagen Artus CMV Assay as a reference assay, There was very good correlation between quantitative results obtained with the RealStar workflow and the CMV Qiagen Artus (correlation coefficient R = 0.93). | SAMPLE_03 | V0524348 | 14/05/2024 | PATIENT | ACCURACY | 21500 | 12500 | 20500 | 11918.6047
Criteria for acceptance | Rule of Thumb for Interpreting the Size of a Correlation Coefficient refer to NIH https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3576830/ Size of Correlation Interpretation 0.90 to 1.00 : Very high positive (negative) correlation 0.70 to 0.90 : High positive (negative) correlation 0.50 to 0.70 :Moderate positive (negative) correlation 0.30 to 0.50 : Low positive (negative) correlation 0.00 to 0.30 :negligible correlation Our data gave Correlation Coefficient ( r ) = 0.994 as calculated here. which is very high positive correlation. | SAMPLE_04 | V0524350 | 14/05/2024 | PATIENT | ACCURACY | 2060 | 1197.6744 | 8590 | 4994.186
Other Notes |  | SAMPLE_05 | V0524514 | 20/05/2024 | PATIENT | ACCURACY | 2630 | 1529.0698 | 4780 | 2779.0698
References | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3576830/ | SAMPLE_06 | V0524522 | 20/05/2024 | PATIENT | ACCURACY | 33000 | 19186.0465 | 45400 | 26395.3488
|  | SAMPLE_07 | V0524523 | 20/05/2024 | PATIENT | ACCURACY | 4020 | 2337.2093 | 4850 | 2819.7674
|  | SAMPLE_08 | V0524524 | 20/05/2024 | PATIENT | ACCURACY | 29400 | 17093.0233 | 32900 | 19127.907
|  | SAMPLE_09 | V0524594 | 22/05/2024 | PATIENT | ACCURACY | 56100 | 32616.2791 | 94000 | 54651.1628
|  | SAMPLE_10 | V0524604 | 22/05/2024 | PATIENT | ACCURACY | 5890 | 3424.4186 | 8250 | 4796.5116
|  | SAMPLE_11 | V0227220 | 28/05/2024 | PATIENT | ACCURACY | 2310 | 1343.0233 | 9210 | 5354.6512
|  | SAMPLE_12 |  | 26/01/2023 | CAP | ACCURACY | 2410 | 1401.1628 | 5992.5 | 3484.0116
|  | SAMPLE_13 |  | 31/01/2022 | CAP | ACCURACY | 43600 | 25348.8372 | 60350 | 35087.2093
|  | SAMPLE_14 | V0624092 | 04/06/2024 | PATIENT | ACCURACY | 3040 | 1767.4419 | 4810 | 2796.5116
|  | SAMPLE_15 | V0624108 | 04/06/2024 | PATIENT | ACCURACY | 5910 | 3436.0465 | 3250 | 1889.5349
|  | SAMPLE_16 | V0624153 | 05/06/2024 | PATIENT | ACCURACY | 1210 | 703.4884 | 1160 | 674.4186
|  | SAMPLE_17 | V0624352 | 12/06/2024 | PATIENT | ACCURACY | 96900 | 56337.2093 | 78100 | 45406.9767
|  | SAMPLE_18 | V0624333 | 12/06/2024 | PATIENT | ACCURACY | 6020 | 3500 | 3500 | 2034.8837
|  | SAMPLE_19 | V0624330 | 12/06/2024 | PATIENT | ACCURACY | 1450 | 843.0233 | 1030 | 598.8372
|  | SAMPLE_21 | V0624278 | 11/06/2024 | PATIENT | ACCURACY | 5110 | 2970.9302 | 2140 | 1244.186
|  | SAMPLE_22 | V0624240 | 10/06/2024 | PATIENT | ACCURACY | 4590 | 2668.6047 | 8610 | 5005.814
correlation coefficient R
0.93374
|  | SAMPLE_23 | V0227301 | 30/05/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_24 | V0227193 | 27/05/20204 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_25 | V0524654 | 22/05/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_26 | V0227175 | 27/05/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_27 | V0624174 | 05/06/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_28 | V0624167 | 05/06/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_29 | V0624127 | 03/06/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_30 | V0624105 | 04/06/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_31 | V0624089 | 04/06/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_32 |  |  | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_33 |  |  | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_34 | V0624071 | 03/06/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE
|  | SAMPLE_35 | V0624054 | 03/06/2024 | PATIENT | ACCURACY | NEGATIVE | NEGATIVE |  | NEGATIVE

## Precision
PRECISION
Item | Data | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | FIRST RUN COPIES/ML | FIRST RUN IU/ML | SECOND RUN COPIES/ML | SECOND RUN IU/ML | THIRD RUN COPIES/ML | THIRD RUN IU/ML | Expected Result QIAGEN KIT Copies/ml | Expected Result QIAGEN KIT IU/ml
Number/Info of samples | 8 Whole Blood Samples were used for repeatability validation. 20 Whole Blood Samples were used for reproducibility validation. | SAMPLE_13 | CAP | REPRODUCIBILITY | 43600 | 25348.8372 | 36400 | 21162.7907 | 24100 | 14011.6279 | 25100 | 14593.0233
Analytical Precision Study Design | Repeat measurement of samples at varying concentrations within run and between runs over a period of time. Multiple estimations in replicates under particular conditions which could be similar or changed . These conditions could be the same method, by the same operator, on the same instrument, within the same laboratory, using the same reagent materials within a short period(the same day) which represents the process entitled repeatability . Whereas, the previous process options but with different operators, different reagents, different days and the same test method was considered as reproducibility | SAMPLE_02 | V0524283 | REPRODUCIBILITY | 187000 | 108720.9302 | 193000 | 112209.3023 | 130000 | 75581.3953 | 131000 | 76162.7907 | Repeat measurement of samples at varying concentrations between three runs by different operators.
Data analysis | The Realstar CMV PCR Kit was compared to the Qiagen Artus CMV. DNA was isolated using the QIAamp DSP Virus Kit, and analysis was carried out on theBiorad cfx 96 (1) Instrument. For comparative testing, the results obtained by using the artus CMV RG PCR Kit were compared to those of the RealStar CMV Assay. In comparison to the results collected with the Qiagen Artus CMV Assay as a reference assay, There was very good correlation between quantitative results obtained with the RealStar workflow and the CMV Qiagen Artus (CV <5.0). | SAMPLE_06 | V0524522 | REPRODUCIBILITY | 33000 | 19186.0465 | 62600 | 36395.3488 | 34800 | 20232.5581 | 45400 | 26395.3488
Criteria for acceptance | CV MUST BE <5.0 Refer to : https://www.westgard.com/lessons/z-stats-basic-statistics/73-lesson34.html#:~:text=CVs%20of%205%25%20or%20less,the%20CV%20may%20be%20low. | SAMPLE_04 | V0524350 | REPRODUCIBILITY | 5060 | 2941.8605 | 8190 | 4761.6279 | 6200 | 3604.6512 | 8590 | 4994.186
Other Notes |  | SAMPLE_04 1/2 | V0524350 | REPRODUCIBILITY |  |  | 1020 | 593.0233 | 1190 | 691.8605 | 880 | 511.6279
References | https://www.westgard.com/lessons/z-stats-basic-statistics/73- | SAMPLE_04 1/2 | V0524350 | REPRODUCIBILITY |  |  | 1100 | 639.5349 | 1080 | 627.907 | 880 | 511.6279 | MEAN | 3450
|  | SAMPLE_04 1/2 | V0524350 | REPRODUCIBILITY |  |  | 789 | 458.7209 | 985 | 572.6744 | 880 | 511.6279 | STDDEV | 4451.3706
|  | SAMPLE_21 | V0624278 | REPRODUCIBILITY | 5910 | 3436.0465 | 4895 | 2845.9302 | 4120 | 2395.3488 | 3250 | 1889.5349 | CV | 1.2903
[6 empty] | 1 read |  | 2nd read |  | at the same run
|  | SAMPLE_13 | CAP | REPEATABILITY | 43600 | 25348.8372 | 17500 | 10174.4186 |  |  | 25100 | 14593.0233 | Repeat measurement of samples at varying concentrations within the same run by same techonologist
|  | SAMPLE_02 | V0524283 | REPEATABILITY | 187000 | 108720.9302 | 145000 | 84302.3256 |  |  | 131000 | 76162.7907
|  | SAMPLE_06 | V0524522 | REPEATABILITY | 33000 | 19186.0465 | 33500 | 19476.7442 |  |  | 45400 | 26395.3488
|  | SAMPLE_04 | V0524350 | REPEATABILITY | 2060 | 1197.6744 | 7190 | 4180.2326 |  |  | 8590 | 4994.186
|  | SAMPLE_05 | V0524514 | REPEATABILITY | 3350 | 1947.6744 | 5264 | 3060.4651 |  |  | 4780 | 2779.0698
|  | SAMPLE_04 | V0524350 | REPEATABILITY | 2060 | 1197.6744 | 3280 | 1906.9767 |  |  | 8590 | 4994.186
|  | SAMPLE_04 1/2 | V0524350 | REPEATABILITY | 1120 | 651.1628 | 690 | 401.1628 |  |  | 880 | 511.6279
|  | SAMPLE_04 1/2 | V0524350 | REPEATABILITY | 958 | 556.9767 | 913 | 530.814 |  |  | 880 | 511.6279
|  | SAMPLE_04 1/2 | V0524350 | REPEATABILITY | 989 | 575 | 741 | 430.814 |  |  | 880 | 511.6279
|  | SAMPLE_21 | V0624278 | REPEATABILITY | 4210 | 2447.6744 | 3589 | 2086.6279 |  |  | 3250 | 1889.5349
|  | SAMPLE_09 DILUTED | V0524594 | REPEATABILITY | 2430 | 1412.7907 | 2510 | 1459.3023 |  |  | 1800 | 1046.5116
|  | SAMPLE_09 DILUTED | V0524594 | REPEATABILITY | 2150 | 1250 | 1300 | 755.814 |  |  | 1800 | 1046.5116
|  | SAMPLE_09 DILUTED | V0524594 | REPEATABILITY | 2110 | 1226.7442 | 2310 | 1343.0233 |  |  | 1800 | 1046.5116
|  | SAMPLE_09 DILUTED | V0524594 | REPEATABILITY | 1110 | 645.3488 | 1080 | 627.907 |  |  | 1800 | 1046.5116
|  | SAMPLE_09 DILUTED | V0524594 | REPEATABILITY | 1290 | 750 | 1620 | 941.8605 |  |  | 1800 | 1046.5116
[13 empty] | MEAN | 15890
[13 empty] | STDDEV | 34084.2431
[13 empty] | CV | 2.145

## Reportable range
REPORTABLE RANGE
Item | Data | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | CALCULATED RESULTS Real Star kit Copies/ml | CALCULATED RESULTS Real Star kit IU/ml | Expected Result COPIES/ml | Expected Result IU/ml
Number/Info of samples | 19 reads of standards were used for Reportable range validation | QS 1 | STANDARD CALIBRATORS | REPORTABLE RANGE | 17200000 | 10000000 | 17200000 | 10000000
Reportable Range Study Design | By testing a series of standards and do serial dilution for the lowest std to reach a concentrations across the anticipated measuring range. The measured values are compared to the assigned values, typically by plotting the measured values on the y axis and the assigned values on the x axis. | QS 2 | STANDARD CALIBRATORS | REPORTABLE RANGE | 1720000 | 1000000 | 1720000 | 1000000
Data analysis | The measured values are compared to the assigned values, typically by plotting the measured values on the y axis and the assigned values on the x axis. Calculate Regression Coefficient, R². R²: 0.999 (ACCEPTED VALUE) | QS 3 | STANDARD CALIBRATORS | REPORTABLE RANGE | 172000 | 100000 | 172000 | 100000
Criteria for acceptance | An R2 of 1.0 indicates that the data perfectly fit the linear model REFER TO https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4678365/ | QS 4 | STANDARD CALIBRATORS | REPORTABLE RANGE | 17200 | 10000 | 17200 | 10000
Other Notes |  | CMV S4 1/2 _1 | STANDARD CALIBRATORS | REPORTABLE RANGE | 6480 | 3767.4419 | 8600 | 5000
References | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4678365/ | CMV S4 1/2 _1 | STANDARD CALIBRATORS | REPORTABLE RANGE | 5090 | 2959.3023 | 8600 | 5000
|  | CMV S4 1/2 _1 | STANDARD CALIBRATORS | REPORTABLE RANGE | 6050 | 3517.4419 | 8600 | 5000
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | REPORTABLE RANGE | 2140 | 1244.186 | 4300 | 2500
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | REPORTABLE RANGE | 3130 | 1819.7674 | 4300 | 2500
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | REPORTABLE RANGE | 3350 | 1947.6744 | 4300 | 2500
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | REPORTABLE RANGE | 1190 | 691.8605 | 2150 | 1250
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | REPORTABLE RANGE | 1009 | 586.6279 | 2150 | 1250
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | REPORTABLE RANGE | 1160 | 674.4186 | 2150 | 1250
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | REPORTABLE RANGE | 1023 | 594.7674 | 1075 | 625
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | REPORTABLE RANGE | 1170 | 680.2326 | 1075 | 625
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | REPORTABLE RANGE | 901 | 523.8372 | 1075 | 625
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | REPORTABLE RANGE | 465 | 270.3488 | 456 | 265.1163
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | REPORTABLE RANGE | 666 | 387.2093 | 456 | 265.1163
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | REPORTABLE RANGE | 406 | 236.0465 | 456 | 265.1163
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | REPORTABLE RANGE | 0 | 0 | 268.9 | 156.3372
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | REPORTABLE RANGE | 0 | 0 | 268.9 | 156.3372
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | REPORTABLE RANGE | 0 | 0 | 268.9 | 156.3372 | Regression Coefficient, R²
1

## Reference Intervals
REFERENCE INTERVAL
Item | Data | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit IU/ml | Expected Result QIAGEN KIT IU/ml
Number/Info of samples | 13 Whole Blood Samples were used for Reference Range validation | SAMPLE_23 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
Reference Interval Study Design | Perform assay on normal controls, from individuals who don’t have the disease. | SAMPLE_24 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
Data analysis | All samples results were negative as expected. | SAMPLE_25 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
Criteria for acceptance | Gave results as normal individuals. | SAMPLE_26 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
Other Notes |  | SAMPLE_27 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
References | https://digital.seracare.com/assay-ver-wp | SAMPLE_28 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
|  | SAMPLE_29 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
|  | SAMPLE_30 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
|  | SAMPLE_31 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
|  | SAMPLE_32 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
|  | SAMPLE_33 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
|  | SAMPLE_34 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE
|  | SAMPLE_35 | PATIENT | REFERENCE INTERVAL | NEGATIVE | NEGATIVE

## Sensitivity
SENSITIVITY
Item | Data | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit Copies/ml | Expected Result
Number/Info of samples | 125 reads from Whole Blood Sample and the lowest satandard were used for Sensitivity validation | QS 4 | STANDARD CALIBRATORS | SENSITIVITY | 17200 | 17200
Analytic sensitivity Study Design | By testing a serial dilution of sample with a known concentration of the target substance and the lowest standard until reach the LOD . Results should be compared with those of a gold-standard reference . | CMV S4 1/2 _1 | STANDARD CALIBRATORS | SENSITIVITY | 6480 | 8600
Data analysis | Serial dilution for Patient' sample and the lowest standard with known concentration were done and performed using realstar kit . After analysis an expected values from each dilution were given . | CMV S4 1/2 _1 | STANDARD CALIBRATORS | SENSITIVITY | 5090 | 8600
Criteria for acceptance | results of samples after dilution gave the results from the highest value to the lowest value the can be detected by a kit. lowest value that can be detected with confidence 80% is 456 copies/ml | CMV S4 1/2 _1 | STANDARD CALIBRATORS | SENSITIVITY | 6050 | 8600
Other Notes |  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | SENSITIVITY | 3140 | 4300
References |  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | SENSITIVITY | 3130 | 4300
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | SENSITIVITY | 3350 | 4300
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | SENSITIVITY | 2190 | 2150
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | SENSITIVITY | 2100 | 2150
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | SENSITIVITY | 1860 | 2150
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | SENSITIVITY | 1023 | 1075
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | SENSITIVITY | 1170 | 1075
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | SENSITIVITY | 901 | 1075
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | SENSITIVITY | 465 | 465
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | SENSITIVITY | 666 | 465
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | SENSITIVITY | 406 | 465
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | SENSITIVITY | 0 | 268.9
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | SENSITIVITY | 0 | 268.9
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | SENSITIVITY | 0 | 268.9
|  | SAMPLE_09 | PATIENT SAMPLE | SENSITIVITY | 5600
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2430 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2150 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2110 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1110 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1290 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2010 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1300 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1310 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1080 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1620 | 1046.5116
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1130 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1210 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1520 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1010 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1260 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1170 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1070 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1542 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 717 | 523
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 776 | 523
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 419 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 979 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 600 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 841 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 857 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 402 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 745 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 460 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 645 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 969 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1170 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 941 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1020 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 687 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2200 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 317 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1100 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 629 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 662 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1050 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 609 | 419
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1150 | 419
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 640 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 352 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 315 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 525 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 291 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 512 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 760 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 452 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 315 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 241 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 380 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 761 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 763 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 502 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 574 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 451 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 315 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1610 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 643 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1090 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 570 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 554 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 594 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 269 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 573 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 980 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 182 | 291
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 366 | 291
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 740 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 323 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 574 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 765 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 362 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 588 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 768 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 751 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 784 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 242 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 387 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 426 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 305 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 315 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 289 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 324 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 354 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 242 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 387 | 265
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 426 | 265
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 207
|  |  | BIORAD |  | % | 95% CI.
|  |  | SAMPLE 1800 | 10 FROM 10 | 100%
|  |  | SAMPLE 900 | 9 FROM 10 | 90% | 779-2286
|  |  | SAMPLE 720 | 23 FROM 25 | 92% | 601-1017
|  |  | SAMPLE 500 | 20 FROM 25 | 80% | 306-699
|  |  | SAMPLE 456 | 20 FROM 25 | 80% | 301 - 546
|  |  | SAMPLE 365 | 0 FROM 10 | 0%

## Sensitivity (2)
SENSITIVITY
Item | Data | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit Copies/ml | Results Real Star kit IU/ml | Expected Result Copies/ml | Expected Result IU/ml
Number/Info of samples | 112 reads from Whole Blood Sample and the lowest satandard were used for Sensitivity validation | QS 4 | STANDARD CALIBRATORS | SENSITIVITY | 17200 | 10000 | 17200 | 10000
Analytic sensitivity Study Design | By testing a serial dilution of sample with a known concentration of the target substance and the lowest standard until reach the LOD . Results should be compared with those of a gold-standard reference . | CMV S4 1/2 _1 | STANDARD CALIBRATORS | SENSITIVITY | 6480 | 3767.4419 | 8600 | 5000
Data analysis | Serial dilution for Patient' sample and the lowest standard with known concentration were done and performed using realstar kit . After analysis an expected values from each dilution were given . | CMV S4 1/2 _1 | STANDARD CALIBRATORS | SENSITIVITY | 5090 | 2959.3023 | 8600 | 5000
Criteria for acceptance | results of samples after dilution gave the results from the highest value to the lowest value the can be detected by a kit. lowest value that can be detected with | CMV S4 1/2 _1 | STANDARD CALIBRATORS | SENSITIVITY | 6050 | 3517.4419 | 8600 | 5000
Other Notes |  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | SENSITIVITY | 3140 | 1825.5814 | 4300 | 2500
References |  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | SENSITIVITY | 3130 | 1819.7674 | 4300 | 2500
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | SENSITIVITY | 3350 | 1947.6744 | 4300 | 2500
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | SENSITIVITY | 2190 | 1273.2558 | 2150 | 1250
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | SENSITIVITY | 2100 | 1220.9302 | 2150 | 1250
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | SENSITIVITY | 1860 | 1081.3953 | 2150 | 1250
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | SENSITIVITY | 1023 | 594.7674 | 1075 | 625
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | SENSITIVITY | 1170 | 680.2326 | 1075 | 625
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | SENSITIVITY | 901 | 523.8372 | 1075 | 625
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | SENSITIVITY | 465 | 270.3488 | 465 | 265
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | SENSITIVITY | 666 | 387.2093 | 465 | 265
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | SENSITIVITY | 406 | 236.0465 | 465 | 265
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | SENSITIVITY | 0 | 0 | 268.9 | 156.3372
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | SENSITIVITY | 0 | 0 | 268.9 | 156.3372
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | SENSITIVITY | 0 | 0 | 268.9 | 156.3372
|  | SAMPLE_09 | PATIENT SAMPLE | SENSITIVITY | 56000 | 32558.1395
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2430 | 1412.7907 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2150 | 1250 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2110 | 1226.7442 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1110 | 645.3488 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1290 | 750 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2010 | 1168.6047 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1300 | 755.814 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1310 | 761.6279 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1080 | 627.907 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 1047 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1620 | 941.8605 | 1800 | 1046.5116
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1130 | 656.9767 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1210 | 703.4884 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1520 | 883.7209 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1010 | 587.2093 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1260 | 732.5581 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1170 | 680.2326 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1070 | 622.093 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1542 | 896.5116 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 523 IU/ML | PATIENT SAMPLE | SENSITIVITY | 717 | 416.8605 | 900 | 523.2558
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 776 | 451.1628 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 419 | 243.6047 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 979 | 569.186 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 600 | 348.8372 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 841 | 488.9535 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 857 | 498.2558 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 402 | 233.7209 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 745 | 433.1395 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 460 | 267.4419 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 645 | 375 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 969 | 563.3721 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1170 | 680.2326 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 941 | 547.093 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1020 | 593.0233 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 687 | 399.4186 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 2200 | 568 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 317 | 295 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1100 | 639.5349 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 629 | 365.6977 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 662 | 384.8837 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1050 | 610.4651 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 609 | 354.0698 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 419 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1150 | 668.6047 | 720 | 418.6047
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 640 | 372.093 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 352 | 204.6512 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 315 | 283 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 525 | 305.2326 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 291 | 169.186 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 512 | 297.6744 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 760 | 441.8605 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 452 | 262.7907 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 315 | 183.1395 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 241 | 240 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 380 | 220.9302 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 761 | 442.4419 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 763 | 443.6047 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 502 | 291.8605 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 574 | 333.7209 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 451 | 262.2093 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 315 | 183.1395 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1610 | 489 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 643 | 373.8372 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 291 IU/ML | PATIENT SAMPLE | SENSITIVITY | 1090 | 633.7209 | 500 | 290.6977
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 570 | 331.3953 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 554 | 322.093 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 594 | 345.3488 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 269 | 156.3953 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 573 | 333.1395 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 980 | 569.7674 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 182 | 206 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 366 | 212.7907 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 740 | 430.2326 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 323 | 187.7907 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 574 | 333.7209 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 765 | 444.7674 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 362 | 210.4651 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 588 | 341.8605 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 768 | 446.5116 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 751 | 436.6279 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 784 | 455.814 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 242 | 140.6977 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 387 | 225 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 265 IU/ML | PATIENT SAMPLE | SENSITIVITY | 426 | 247.6744 | 456 | 265.1163
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  | SAMPLE_09_DILUTED TO 207 IU/ML | PATIENT SAMPLE | SENSITIVITY | 0 | 0 | 356 | 206.9767
|  |  | BIORAD |  | % | % | 95% CI. | 95% CI
|  |  | SAMPLE 1047 IU/ML | 10 FROM 10 | 100% | 100% |  | 749-1159
|  |  | SAMPLE 523 IU/ML | 9 FROM 9 | 90% | 100% | 779-2286 | 573-800
|  |  | SAMPLE 419 IU/ML | 23 FROM 23 | 92% | 100% | 601-1017 | 389_583
|  |  | SAMPLE 291 IU/ML | 20 FROM 20 | 80% | 100% | 306-699 | 246-421
|  |  | SAMPLE 265 IU/ML | 20 FROM 20 | 80% | 100% | 301 - 546 | 255-372
|  |  | SAMPLE 207 IU/ML | 0 FROM 10 | 0% | 0%
Serial dilution for Patient' sample with known concentration were done from concentration (1047IU/ML - 207 IU/ML) and performed using realstar kit to detect the lowest concentration with high confidence can be detected by kit. After analysis an expected values from each dilution were given . 95 % CONFEDINCE INTERVAL FOR (265 IU/ML) IS : 255-372. Our cutoff will be : 265 IU/ML.
Disclaimer when result is negative: A negative quqntitative CMV PCR result does not rule out the presence of CMV genomic DNA below the assay Limit of Detection (LOD 265 IU/ML; 95% CI = 255-372).
Methodology: A real-time PCR technology was performed using CMV PCR RealStar® Kit 1.0 using CFX-96 BioRad instrument, for the detection and quantification of cytomegalovirus (CMV) specific DNA in human EDTA whole blood
RECOMMENDATIONS
Our recommendation : Any sample with a positive result below the 265 IU/ML LOD, will be repeated for confirmation. If the the first and second results are discordant , a third atempt may be performed, using a newly extracted sample nucleic acid .
CAP RESULTS
|  |  | SAMPLE CAP | LOG 10 IU/ML |  | IU/ML |  | CAP RANGE LOG 10 IU/ML | CAP RANGE IU/ML
|  |  | VLS04_2024 | 4.62 |  | 41686 |  | 4.35-4.93 | 22387-85113
[8 empty] | MEAN : 4.60 | 39810
|  |  | VLS13_2023 | 3.96 |  | 9135 |  | 3.35-4.07 | 2238.7-11748
[8 empty] | MEAN : 3.77 | 5919
|  |  | VLS_03_2023 | 3.58 |  | 3801 |  | 3.33-3.76 | 2137-5754
[8 empty] | MEAN : 3.52 | 3311
|  |  | VLS_04_2023 | 4.04 |  | 10964 |  | 3.62-4.13 | 4168-13489
[8 empty] | MEAN : 3.91 | 8128
CAP RESULTS
|  |  | SAMPLE CAP | ROCHE COBAS LOG 10 IU/ML |  | IU/ML |  | CAP RANGE LOG 10 IU/ML | CAP RANGE IU/ML
|  |  | VLS04_2024 |  |  |  |  | 4.35-4.72 | 22387-82450
[8 empty] | MEAN : 4.57 | 37153

## Specificity
SPECIFICITY | [7 empty] | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY | SENSITIVITY
Item | Data | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit IU/ml | Expected Result QIAGEN KIT IU/ml
Number/Info of samples | 13 Whole Blood Samples were used for Specificity validation | SAMPLE_23 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | BK POSITIVE
Analytic specificity Study Design | Use samples with positive results for other viruses . Use icteric and lipemic samples. For checking if these factor give false positive results. | SAMPLE_24 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | HBV POSITIVE
Data analysis | After comparing the resuts with previous reference method, All samples gave negative results . | SAMPLE_25 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | HBV POSITIVE
Criteria for acceptance | TN/[TN+FP]*100: 13/(13+0)*100 = 100% | SAMPLE_26 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | EBV POSITIVE
Other Notes |  | SAMPLE_27 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | EBV POSITIVE
References | Instructions for use AltoStar® CMV PCR Kit 1.5 HandBook | SAMPLE_28 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | EBV POSITIVE
|  | SAMPLE_29 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | HBV POSITIVE
|  | SAMPLE_30 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | ENTERO POSITIVE
|  | SAMPLE_31 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | EBV POSITIVE
|  | SAMPLE_32 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | LIPEMIC
|  | SAMPLE_33 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | ICTERIC
|  | SAMPLE_34 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | EBV POSITIVE
|  | SAMPLE_35 | PATIENT | SPECIFICITY | NEGATIVE | NEGATIVE | EBV POSITIVE

## Other Criteria
OTHER CRITERIA
Item/Data | Number/Info of samples | Study Design | Data analysis | Criteria for acceptance | Other Notes | References
Reagent stability | 1 CAP sample repeated over 6 weeks | Relate to performance of a product within a specified time period. the apparent change in measured quantity of an analyte over time, referred to as measurand drift. This used to assess actual changes in concentration . Define the period of time in which the product remains suitable after being placed into use until expiry date, this include the duration that a control product may be used after being first opened . | compare results with the original quantity that results by the reference method and calculate the r^2 which must be aroud 1 . The result is :0.924 which is accepted. | calculate the r^2 which must be aroud 1 . The result are : 0.924 which accepted. | https://www.westgard.com/essays/guest-essay/523-clsiep25.html | https://www.westgard.com/essays/guest-essay/523-clsiep25.html
Cut-Off | Covered by sensitivity validation
Truness/Bais | covered by accuracy validation
Linearity | 4 kit STDs | How accurate the measurements are through the expected range of measurements. By testing a series of samples of known concentrations and 4 standards and do serial dilution for the Highest std to reach a concentrations across the anticipated measuring range (AMR). The measured values are compared to the assigned values, typically by plotting the measured values on the y axis and the assigned values on the x axis. | Compare results with the original quantity that were expected and calculate the r^2 which must be around 1. The result are :0.997 which accepted. | The measured values are compared to the assigned values, typically by plotting the measured values on the y axis and the assigned values on the x axis. Calculate Regression Coefficient, R². R²: 0.997 (ACCEPTED VALUE) | https://www.westgard.com/lessons/basic-method-validation/54-lesson26.html | https://www.westgard.com/lessons/basic-method-validation/54-lesson26.html
Carryover/Cross-contamination |  | Testing CMV high positive and negative samples sequentialy in the same run | compare results with the original quantity that results by the reference method | the negative samples gave negative results .
Uncertinity | Covered by specificity and sensitivity
Method comparison
Diagnostic/Clinical accuracy
Limitations

## Reagent Stability
REAGENT STABILITY
SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | Results Real Star kit Copies/ml | Results Real Star kit IU/ml | Expected Result QIAGEN KIT Copies/ml | Expected Result QIAGEN KIT IU/ml
SAMPLE 04 | CAP SAMPLE |  | 7362 | 4280.2326 | 6035 | 3508.7209
SAMPLE 04 | CAP SAMPLE |  | 6408.3333 | 3725.7752 | 6035 | 3508.7209 | Run every weak
SAMPLE 04 | CAP SAMPLE |  | 4508.3333 | 2621.124 | 6035 | 3508.7209
SAMPLE 04 | CAP SAMPLE |  | 4708.3333 | 2737.4031 | 6035 | 3508.7209
SAMPLE 04 | CAP SAMPLE |  | 5178.3333 | 3010.6589 | 6035 | 3508.7209
SAMPLE 04 | CAP SAMPLE |  | 6541.3333 | 3803.1008 | 6035 | 3508.7209

## linearity
linearity
Item | Data | SAMPLE ID | DESCRIPTION (QC, PT, Patient sample, etc.) | PURPOSE | CALCULATED RESULTS Real Star kit Copies/ml | CALCULATED RESULTS Real Star kit IU/ml | Expected Result Copies/ml | Expected Result IU/ml
Number/Info of samples | 19 reads of standards were used for linearity validation | QS 1 | STANDARD CALIBRATORS | LINEARITY | 17200000 | 10000000 | 17200000 | 10000000
Reportable Range Study Design | By testing a series of standards and do serial dilution for the lowest std to detect an analyte range measurments from the highest to thelowest value that can be detected by kit. The measured values are compared to the assigned values, typically by plotting the measured values on the y axis and the assigned values on the x axis. | QS 2 | STANDARD CALIBRATORS | LINEARITY | 1720000 | 1000000 | 1720000 | 1000000
Data analysis | The measured values are compared to the assigned values, typically by plotting the measured values on the y axis and the assigned values on the x axis. Calculate Regression Coefficient, R². R²: 0.999 (ACCEPTED VALUE) | QS 3 | STANDARD CALIBRATORS | LINEARITY | 172000 | 100000 | 172000 | 100000
Criteria for acceptance | An R2 of 1.0 indicates that the data perfectly fit the linear model REFER TO https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4678365/ | QS 4 | STANDARD CALIBRATORS | LINEARITY | 17200 | 10000 | 17200 | 10000
Other Notes |  | CMV S4 1/2 _1 | STANDARD CALIBRATORS | LINEARITY | 6480 | 3767.4419 | 8600 | 5000
References | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4678365/ | CMV S4 1/2 _1 | STANDARD CALIBRATORS | LINEARITY | 5090 | 2959.3023 | 8600 | 5000
|  | CMV S4 1/2 _1 | STANDARD CALIBRATORS | LINEARITY | 6050 | 3517.4419 | 8600 | 5000
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | LINEARITY | 2140 | 1244.186 | 4300 | 2500
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | LINEARITY | 3130 | 1819.7674 | 4300 | 2500
|  | CMV S4 1/2 _2 | STANDARD CALIBRATORS | LINEARITY | 3350 | 1947.6744 | 4300 | 2500
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | LINEARITY | 1190 | 691.8605 | 2150 | 1250
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | LINEARITY | 1009 | 586.6279 | 2150 | 1250
|  | CMV S4 1/2 _3 | STANDARD CALIBRATORS | LINEARITY | 1160 | 674.4186 | 2150 | 1250
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | LINEARITY | 1023 | 594.7674 | 1075 | 625
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | LINEARITY | 1170 | 680.2326 | 1075 | 625
|  | CMV S4 1/2 _4 | STANDARD CALIBRATORS | LINEARITY | 901 | 523.8372 | 1075 | 625
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | LINEARITY | 465 | 270.3488 | 456 | 265.1163
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | LINEARITY | 666 | 387.2093 | 456 | 265.1163
|  | CMV S4 1/2.3 _5 | STANDARD CALIBRATORS | LINEARITY | 406 | 236.0465 | 456 | 265.1163
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | LINEARITY | 0 | 0 | 268.9 | 156.3372
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | LINEARITY | 0 | 0 | 268.9 | 156.3372
|  | CMV S4 1/2 _6 | STANDARD CALIBRATORS | LINEARITY | 0 | 0 | 268.9 | 156.3372 | Regression Coefficient, R²
1

## Section-Specific Criteria
Item/Data | Number/Info of samples | Study Design | Data analysis | Criteria for acceptance | Other Notes | References