Additional files

Additional file 1, Table 1,  Sensitivity analyses including all presumptive patients- Agreement between results from SS and MS using GeneXpert® MTB/RIF

- Assuming that all the missing morning specimens would give the same GeneXpert® MTB/RIF results as their counterpart spot specimens

| Morning Specimen | Spot Specimen# |  |  | Pa | Overall percentage of agreement (95% CI)b | Positive percent agreement (95% CI)b | Negative percent agreement (95% CI)b |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Positive | Negative | Total |  |  |  |  |
| Total | 183 | 703 | 886 | 0.99 | 99.3 (98.5-99.7) | 98.4 (95.3-99.4) | 99.6 (98.8-99.9) |
| Positive | 180 | 3 | 183 |  |  |  |  |
| Negative | 3 | 700 | 703 |  |  |  |  |
| Spot Specimen | Morning Specimen# |  | Total | Pa | Overall percentage of agreement (95% CI)b | Positive percent agreement (95% CI)b | Negative percent agreement (95% CI)b |
|  | Positive | Negative |  |  |  |  |  |
| Total | 155 | 582 | 886 | 0.99 | 99.3 (98.5-99.7) | 98.4 (95.3-99.4) | 99.6 (98.8-99.9) |
| Positive | 152 | 3 | 183 |  |  |  |  |
| Negative | 3 | 579 | 703 |  |  |  |  |

a McNemar test. b Wilson 95% confidence intervals. # Used as reference standard.

- Assuming that all the missing morning specimens would give the opposite GeneXpert® MTB/RIF results to their counterpart spot specimens

| Morning Specimen | Spot Specimen# |  |  | Pa | Overall percentage of agreement (95% CI)b | Positive percent agreement (95% CI)b | Negative percent agreement (95% CI)b |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Positive | Negative | Total |  |  |  |  |
| Total | 183 | 703 | 886 | <0.001 | 82.5 (79.9-84.9) | 83.1 (77.0-87.8) | 82.4 (79.4-85.0) |
| Positive | 152 | 124 | 276 |  |  |  |  |
| Negative | 31 | 579 | 610 |  |  |  |  |
| Spot Specimen | Morning Specimen# |  | Total | Pa | Overall percentage of agreement (95% CI)b | Positive percent agreement (95% CI)b | Negative percent agreement (95% CI)b |
|  | Positive | Negative |  |  |  |  |  |
| Total | 276 | 610 | 886 | <0.001 | 82.5 (79.9-84.9) | 55.1 (49.1-60.8) | 94.9 (92.9-96.4) |
| Positive | 152 | 31 | 183 |  |  |  |  |
| Negative | 124 | 579 | 703 |  |  |  |  |

a McNemar test. b Wilson 95% confidence intervals. # Used as reference standard.