Supplementary Table 3: Instrumentation and operational settings for U-Pb dating

| Laboratory & Sample Preparation |  |
| --- | --- |
| Laboratory name | Department of Earth Science, University of Bergen |
| Sample type/mineral | Zircon |
| Sample preparation | Conventional mineral separation, 2.5 cm resin mount, 0.025 m polish to finish |
| Imaging | CL, Zeiss Supra55 VP, 15 kV, 1 nA, 15mm working distance |
| Laser ablation system |  |
| Make, Model & type | RESOlution M-50 LR with a Coherent COMPexPRO® 110 193 nm ArF excimer laser |
| Ablation cell & volume | Two volume |
| Laser wavelength | 193 nm |
| Pulse duration | 20 ns |
| Fluence | 2.2 - 2.5 J.cm-2 |
| Repetition rate | 5 Hz |
| Ablation duration | 30 secs |
| Spot diameter | 26 µm |
| Sampling mode/pattern | Static spot ablation/circular |
| Carrier gas | He (0.75 l/min) with small amounts of N2 (0.004 ml/min) mixed in before entering the ICP-MS to increase sensitivity. |
| Signal smoothing device | “Squid” connected between the laser and the ICP-MS |
| ICP-MS Instrument |  |
| Make, Model & type | Nu Instruments, Nu Attom HR, SC-ICP-MS |
| Sample introduction | Ablation aerosol from laser ablation |
| RF power | 1300W |
| Cool gas | Ar 13 l/min |
| Aux gas | Ar 0.7 l/min |
| Make-up gas flow | Ar 0.49 l/min |
| Detection system | MasCom Electron Multiplier |
| Masses measured | 202Hg, 204(Hg + Pb), 206Pb, 207Pb, 208Pb, 232Th, 238U |
| Integration time per peak/dwell times | 200µs for 202Hg, 204(Hg + Pb), 206Pb, 208Pb, 232Th, 238U 800 µs for 207Pb and 235U |
| Number of sweeps per cycle | 10 |
| Total time per cycle | 0.0493 s |
| Analysis method | Deflector jump |
| IC Dead time | 14 ns |
| Detection Mode | Ion counting mode and ion-attenuated mode |
| Data Processing |  |
| Gas blank | 15 s |
| Calibration strategy | 91500 1065 Ma1 used as primary reference material, while Plesovice 337 Ma2, Mud Tank 731.8 Ma3, & GJ1 609 Ma4 are used as secondaries for validation. |
| Data processing package used / Correction for LIEF | Iolite4 (v. 4.4.5) for data normalization, uncertainty propagation, and age calculation. LIEF correction assumes reference material and samples behave identically. |
| Common-Pb correction | No common-Pb correction was applied to the data. |
| Data reduction | VizualAge UComPbine5 |
| Down-hole correction model | Exponential or Exponential + linear |
| Blank and instrumental bias correction | Spline_AutoSmooth |
| Uncertainty level & propagation | Ages are quoted at 2s absolute, propagation is by quadratic addition. |
| Quality control / Validation | Run1 – Mount1: Plesovice – Wtd ave 206Pb/238U age = 334 ± 1 (2s, MSWD = 7.39, n=4) GJ-1 – Wtd ave 206Pb/238U age = 600 ± 2 (2s, MSWD = 6.42, n=4) 91500 – Wtd ave 206Pb/238U age = 1062 ± 2 (2s, MSWD = 0.85, n=4) Run2 – Mount2: Plesovice – Wtd ave 206Pb/238U age = 353 ± 1 (2s, MSWD = 13.7, n=10) GJ-1 – Wtd ave 206Pb/238U age = 630 ± 1 (2s, MSWD = 7.22, n=12) 91500 – Wtd ave 206Pb/238U age = 1062 ± 1 (2s, MSWD = 0.30, n=12) Run3 – Mount3: Plesovice – Wtd ave 206Pb/238U age = 351 ± 1 (2s, MSWD = 2.48, n=11) GJ-1 – Wtd ave 206Pb/238U age = 618 ± 1 (2s, MSWD = 3.50, n=12) 91500 – Wtd ave 206Pb/238U age = 1063 ± 1 (2s, MSWD = 0.80, n=16) Mud tank - Wtd ave 206Pb/238U age = 728 ± 2 (2s, MSWD = 0.44, n=12) Run4 – Mount4: Plesovice – Wtd ave 206Pb/238U age = 345 ± 1 (2s, MSWD = 6.70, n=65) GJ-1 – Wtd ave 206Pb/238U age = 610 ± 1 (2s, MSWD = 5.98, n=65) 91500 – Wtd ave 206Pb/238U age = 1062 ± 1 (2s, MSWD = 0.92, n=63) |
| Data reporting | Reported in the format recommended by Horstwood 3 using a python script made by Joe Petrus (https://github.com/iolite-LA-ICP-MS/iolite4-python-examples/blob/master/export/PlasmAge.py). |
| Data visualization | Plotting of KDE using IsoplotR6. |

1.	Wiedenbeck M, et al. Three Natural Zircon Standards for U-Th-Pb, Lu-Hf, Trace Element and Ree Analyses. Geostandards and Geoanalytical Research, 19, 1-23 (1995)

2.	Sláma J, et al. Plešovice zircon — A new natural reference material for U–Pb and Hf isotopic microanalysis. Chemical Geology, 249, 1-35 (2008)

3.	Horstwood MSA, et al. Community-Derived Standards for LA-ICP-MS U-(Th-)Pb Geochronology – Uncertainty Propagation, Age Interpretation and Data Reporting. Geostandards and Geoanalytical Research, 40, 311-332 (2016)

4.	Jackson SE, Pearson NJ, Griffin WL, Belousova EA. The application of laser ablation-inductively coupled plasma-mass spectrometry to in situ U–Pb zircon geochronology. Chemical Geology, 211, 47-69 (2004)

5.	Chew DM, Petrus JA, Kamber BS. U–Pb LA–ICPMS dating using accessory mineral standards with variable common Pb. Chemical Geology, 363, 185-199 (2014)

6.	Vermeesch P. IsoplotR: A free and open toolbox for geochronology. Geoscience Frontiers, 9, 1479-1493 (2018)