|  | Pelvis | HN | Lung | Lung SBRT |  | Breast | Total |
| --- | --- | --- | --- | --- | --- | --- | --- |
| AXB (Dm,m) | -19% | -10% | -12.7% | -11.4% | -23.5% |  | -16.1% |
| Passed (>95%) | 0/10 | 1/10 | 0/5 | 0/5 | 0/5 |  | 1/40 |
| AXB (Dm,m)cal | -1.9% | -1.4% | -1.3% | -8.9% | -3.2% |  | -2.9% |
| Passed (>95%) | 10/10 | 9/10 | 4/5 | 0/5 | 7/10 |  | 30/40 |
| Improvement | 17.1% | 8.7% | 11.4% | 2.5% | 20.3% |  | 13.2% |

Table 1. Average percentage differences between AAA (reference) and AXB (Dm,m) comparisons with measurements in the ArcCHECK phantom.

Note: The passing rates, grouped by treatment site, are shown for both the standard dose-to-water calibration (AXB (Dm,m)) and the dose-to-medium calibration (AXB (Dm,m)cal).  The absolute difference (improvement) between AXB (Dm,m) and AXB (Dm,m)cal is presented. HN: head and neck, Pelvis: 5 cervix and 5 prostate plans.

Table 2. Absolute doses (cGy/100 MU) measured during the ArcCHECK dose-to-medium calibration outlined in Section 2B.

|  | 6 MV | 6 FFF | 10 MV | 10 FFF |
| --- | --- | --- | --- | --- |
| Dw,w | 120.1 | 118.8 | 123.7 | 123.0 |
| Dw,p | 120.0 | 118.0 | 123.0 | 120.7 |
| Dp,p | 116.4 | 114.6 | 119.5 | 117.2 |

Note: Dw,p is the measured dose using a Farmer chamber, in Perspex. Dp,p is the Spencer-Attix corrected dose in Perspex, calculated using Equation 7, i.e. Dp,p = Dw,pkp,w. Dw,w values from a previous dose-to-water calibration of the ArcCHECK (Section 2A).

| Comparison with CC04 | Pelvis | HN | Lung | Lung SBRT | Breast |
| --- | --- | --- | --- | --- | --- |
| AXB (Dm,m) (AXB Ref) | -0.1% | 0.3% | -1.0% | -0.8% | -0.7% |
| AXB (Dm,m) (AAA Ref) | 4.2% | 5.0% | 4.9% | 3.0% | 4.3% |
| AAA (Dw,w) (AAA Ref) | 0.1% | -0.4% | 0.5% | 0.8% | -0.4% |

Table 3. Average percentage differences between measured CC04 and AXB  (Dm,m) calculated point doses in the ArcCHECK phantom,

Note: The measurements are shown over the various retrospective patient plans: 10 Pelvis (5 prostate, 5 cervix), 10 head and neck (HN), 5 VMAT lung, 5 SBRT lung, and 10 breast plans. AXB  (Dm,m) calculated reference doses in the ArcCHECK phantom were used for relative calibration of the CC04 measurements. AAA and AXB  (Dm,m)  comparisons with CC04 measurements, using AAA reference doses for CC04 relative calibration are shown as reference. AXB and AAA Ref: CC04 relative doses were calibrated using AXB (Dm,m)  reference and AAA (Dw,w)  TPS calculations in the ArcCHECK phantom, respectively.