Supplementary information:

for

Polar microcystins or arginine methyl ester can serve as sensitive reference materials for system suitability tests in untargeted metabolomics using reversed-phase LC-MS: A case study	

Authors: Tommy Melzera, Georg Pohnerta, Nico Ueberschaarb*

a: Friedrich Schiller University Jena, Institute of Inorganic and Analytical Chemistry, Lessingstr. 8, 07743 Jena, Germany	
b: Friedrich Schiller University Jena, Mass Spectrometry Platform, Humboldtstr. 8, 07743 Jena, Germany

*Corresponding author email: nico.ueberschaar@uni-jena.de

Table S1: Retention times (Rt) and peak intensities of reference materials analysed on the deteriorated and a new Accucore C18 RP column

|  | Column intact |  | Column deteriorated |  | Deviation* [%] |  |
| --- | --- | --- | --- | --- | --- | --- |
| Reference material | Rt [min] | Intensity | Rt [min] | Intensity | Rt | Intensity |
| l-arginine | 0.60 | 1.1·107 | 0.63 | 8.6·106 | +5 | -22 |
| l-asparagine | 0.57 | 1.2·106 | 0.59 | 1.2·106 | +4 | 0 |
| l-aspartic acid | 0.62 | 1.1·105 | 0.65 | 1.1·105 | +5 | 0 |
| l-cysteine (neg) | 0.60 | 5.9·105 | 0.64 | 3.5·105 | +7 | -41 |
| l-glutamic acid | 0.62 | 1.5·106 | 0.65 | 9.0·105 | +5 | -40 |
| l-histidine | 0.60 | 8.7·106 | 0.63 | 5.6·106 | +5 | -36 |
| l-isoleucine | 0.86 | 1.7·106 | 0.95 | 1.6·106 | +10 | -6 |
| l-leucine | 0.94 | 2.1·106 | 1.05 | 2.1·106 | +12 | +0 |
| l-lysine | 0.56 | 4.3·105 | 0.60 | 5.0·105 | +7 | +16 |
| l-methionine | 0.76 | 4.7·106 | 0.79 | 5.2·106 | +4 | +11 |
| l-phenylalanine | 1.83 | 5.2·106 | 2.04 | 5.4·106 | +11 | +4 |
| l-proline | 0.62 | 1.1·107 | 0.65 | 6.7·106 | +5 | -39 |
| l-serine | 0.61 | 1.3·105 | 0.61 | 8.9·104 | 0 | -32 |
| l-threonine | 0.62 | 2.0·105 | 0.65 | 1.4·105 | +5 | -30 |
| l-tryptophan | 2.47 | 2.6·107 | 2.64 | 2.3·107 | +7 | -12 |
| l-tyrosine | 0.82 | 4.9·106 | 0.93 | 3.5·106 | +13 | -29 |
| l-valine | 0.67 | 2.7·107 | 0.69 | 2.6·107 | +3 | -4 |
| nicotinic acid | 0.75 | 7.1·107 | 0.77 | 8.4·107 | +3 | +18 |
| uridine | 0.75 | 3.4·105 | 0.79 | 5.0·105 | +5 | +47 |
| adenosine | 0.77 | 5.1·107 | 0.86 | 4.3·107 | +12 | -16 |
| inosine | 0.98 | 8.4·105 | 1.01 | 4.7·105 | +3 | -44 |
| creatine | 0.61 | 5.0·107 | 0.65 | 3.8·107 | +7 | -24 |
| creatinine | 0.59 | 9.5·107 | 0.65 | 6.4·107 | +10 | -33 |
| l-carnitine | 0.61 | 1.4·108 | 0.63 | 1.2·108 | +3 | -14 |
| acetyl-l-carnitine | 0.75 | 2.3·108 | 0.77 | 1.9·108 | +3 | -17 |
| l-pyroglutamic acid | 0.73 | 8.7·104 | 0.79 | 8.7·104 | +8 | 0 |
| cholic acid (neg) | 5.13 | 1.9·107 | 5.26 | 1.4·107 | +3 | -26 |
| clycocholic acid | 4.62 | 8.5·106 | 4.74 | 5.7·106 | +3 | -33 |
| d-raffinose (neg) | 0.62 | 2.7·105 | 0.64 | 1.0·105 | +3 | -63 |
| kynurenic acid | 2.66 | 1.8·106 | 2.69 | 2.3·106 | +1 | +28 |
| hippuric acid | 2.86 | 8.2·105 | 2.89 | 7.1·105 | +1 | -13 |
| 4-hydroxy hippuric acid | 2.34 | 4.7·105 | 2.42 | 3.9·105 | +3 | -17 |
| glycerophosphocholine | 0.61 | 6.9·105 | 0.63 | 4.0·105 | +3 | -42 |
| l-arginine methyl ester | 0.57 | 4.9·106 | 1-2.8 broad! | 3.1·105 | +233** | -94 |

* Deviations were calculated by the formula:		 ** Rt = 1.9 min in the medium of the broad peak was used

Deviation Rt % = Rt column deteriorated- Rt column intactRt column intactx 100

Deviation Intensity %= Intensity column deteriorated- Intensity column intactIntensity column intactx 100

Fig S2: Chromatogram of l-arginine methyl ester (top) and MC-RR (bottom); identical samples were analysed on the deteriorated column and a new C18 RP column

LC-MS instrument and used solvents:

The Dionex UltiMate 3000 UHPLC (Thermo Fisher Scientific) includes an WPS-3000 autosampler, an HPG-3400 rapid separation binary pump, and a TCC 3200 column compartment. The autosampler equipped with a 25 µL injection syringe and a 100 µL sample loop was maintained at 10°C. Chromatographic separation was performed at 25°C using a linear gradient of eluent A: 0.1% formic acid and 2% acetonitrile in water and eluent B: acetonitrile; with a constant flow of 0.4 mL/min and a total duration of 12 min. The injection volume for the reference standard mixture was 1 µL. Mass spectra were recorded with a Thermo QExactive plus orbitrap mass spectrometer coupled to a heated electrospray source (HESI). Column flow was switched at 0.5 min from waste to the MS and at 11.5 min again back to the waste, to prevent source contamination. For monitoring two full scan modes with switching polarities were selected with the following parameters. Polarity: positive; scan range: 100 to 1500 m/z; resolution: 70,000; AGC target: 3 × 106; maximum IT: 200 ms. General settings: sheath gas flow rate: 60; auxiliary gas flow rate 20; sweep gas flow rate: 5; spray voltage: 4.0 kV; capillary temperature: 320 °C; S-lens RF level: 90; auxiliary gas heater temperature: 400 °C; acquisition time frame: 0.5 - 11.5 min. For negative mode, all values were kept instead of the spray voltage which was set to 2.5 kV. LC-MS grade methanol was obtained from VWR International GmbH (Darmstadt, Germany), acetonitrile and water as eluents were purchased from Th. Geyer (Höxter, Germany), formic acid from Thermo Fisher (Dreieich, Germany).

Preparation of the in-house mixture used for SST:

The reference standards p-fluoro-l-phenylalanine, p-fluoro-benzoic acid and decanoic acid-D19 were purchased from Merck (Darmstadt, Germany). Each standard was dissolved with 1 mg/mL in methanol and an in-house mixture containing all three standards with 100 µg/mL was prepared. Aliquots with 10 µg/ml, ready for injection into the LC-MS system, were prepared from this mixture and stored at −20 °C.

Reference materials for determining retention times and peak intensities on both, the deteriorated column and a new C18 RP column:

The amino acid mixture AAS18 was purchased from Sigma Aldrich (Taufkirchen, Germany). The amino acid l-serin was analysed from a dilution with 25 µmol/L of this standard; the amino acids l-arginine, l-aspartic acid, l-cysteine, l-glutamic acid, l-histidine, l-isoleucine, l-leucine, l-lysine, l-methionine, l-phenylalanine, l-proline, l-serine, l-threonine and l-tyrosine were analysed from a dilution of this standard containing each amino acid with 0.2  µmol/L. The amino acids l-asparagine and l-tryptophan (Sigma Aldrich) were dissolved in MeOH and analysed with c= 10 mg/L; l-valine (Thermo Fisher) was analysed with 100 mg/L in MeOH.

The reference standards nicotinic acid, uridine, inosine, creatinine, l-pyroglutamic acid, cholic acid and l-arginine methyl ester were purchased from Sigma Aldrich. Adenosine, l-carnitine, d-raffinose, glycerophosphocholine, kynurenic acid, hippuric acid, 4-hydroxy hippuric acid were purchased from LGC Labor GmbH (Augsburg, Germany). Acetyl-l-carnitine and glycocholic acid were obtained from Biomol GmbH (Hamburg, Germany) and creatine from Thermo Fisher. All standards were dissolved in H2O or MeOH and analysed from dilutions with 100 µg/L and 10 µg/L.

All the described reference materials above were analysed from the same sample on both, the deteriorated and a new Accucore C18 RP column.

Weekly Orbitrap maintenance protocol

Routine for weekly maintenance check-up of the QExactive plus orbitrap

- Cool down the capillary to 50 °C

- Spray shield and inlet capillary: Clean 15 min (ultrasonic bath) with solvent mixture acetonitrile:isopropanol:methanol:tetrahydrofuran 1:1:1:1

  - If Spray shield is still dirty: remove fouling outside (polishing device) and repeat above procedure

- Spray capillary:

  - Remove fouling from outside if necessary

  - Flush with methanol (Hamilton syringe, should only give limited backpressure)

  - If necessary (e.g. blocked), disassemble and clean (ultrasonic bath with same solvent mixture described above on , inlet and outlet side)

- Refill

  - Eluent D			(80% acetonitrile / 20% water)

  - Syringe Wash			(20% methanol / 80% water)

  - Rear Seal Wash		(10% methanol / 90% water)

  - Eluents A and B if required 	(needs to be checked by users before run)

- Replace MeOH and MeOH/THF vials in autosampler B1 and B2

- Mass Calibration (positive and negative mode)

- Confirm that correct columns are installed

- If old samples in autosampler, remove

- Check UV/VIS flow cell for humidity leak

- Test backpressure of UHPLC with and without MS (100% ACN, no column, 0.4 mL/min)

- Document the check-up in the operator MS Word LOG file

- Run retention time standards on Accucore C18 RP column to validate retention times and signal intensities (from peak height). Note values & issues in operator MS Excel LOG file

Current calibration solution:

- Bruker, sodium formate, in 50% isopropanol

- Use tune file & customized calibration with 15 µL/min injection speed

- Tune File:	Tune mass calibration HCOONa.mstune (grey marked ions are used for calibration)

Mass list for calibration ions:

| Cluster Number | Negative | Positive |
| --- | --- | --- |
| 1 | 44.9928203 | 90.976645 |
| 2 | 112.985627 | 158.964069 |
| 3 | 180.973050 | 226.951493 |
| 4 | 248.960474 | 294.938917 |
| 5 | 316.947898 | 362.926341 |
| 6 | 384.935322 | 430.913765 |
| 7 | 452.922746 | 498.901189 |
| 8 | 520.910170 | 566.888613 |
| 9 | 588.897594 | 634.876036 |
| 10 | 656.885018 | 702.863460 |
| 11 | 724.872442 | 770.850884 |
| 12 | 792.859866 | 838.838308 |
| 13 | 860.847289 | 906.825732 |
| 14 | 928.834713 | 974.813156 |
| 15 | 996.822137 | 1042.800580 |

Settings for electrospray ionization for calibration:

| Parameter | Positive @ 40CID | Negative @ 25CID |
| --- | --- | --- |
| Sheath gas flow rate | 19 | 17 |
| Aux gas flow rate | 6 | 6 |
| Sweep gas flow rate | 0 | 0 |
| Spray voltage [kV] | 3.5 | 2.5 |
| Capillary temperature [°C] | 250 | 250 |
| S-lens RF level | 50 | 50 |
| Aux gas heater temp [°C] | 113 | 106 |

Run system suitability test using suggested modified solution containing p-fluoro-l-phenylalanine, p-fluoro-benzoic acid, decanoic acid-D19 and l-arginine methyl ester (10 µg/mL each in MeOH) using the pre-defined sequence and document the retention time and intensity for each analyte.