Sensitivity of red blood cell antibody screening systems for anti-D: A multicentre study using external quality assessment programmes

Supplementary Data

Supplement 1:	2

Table S1: Participant EQA providers and characteristics of their relevant cycles	2

Supplement 2:	3

Table S2: Total anti-D antibody detection rates of participants with individual EQA providers	3

Supplement 3:	4

Table S3: Detection rates of anti-D antibody in Samples 1 – 4	4

Supplement 4:	5

Table S4: Anti-D antibody detection rates of different test principles, test systems, and ways of reading reaction results	5

Supplement 5:	6

Table S5: Anti-D antibody detection rates of test systems with different EQA providers, automated vs. manual reading of results	6

Supplement 6:	9

Anti-D antibody detection rates of different test cell preparations	9

Table S6: Anti-D antibody detection rates of different test cell preparations	9

Figure S7: Detection rates of test cells produced by individual manufacturers	10

Supplement 7: PubILC	11

# Supplement 1:

### Table S1: Participant EQA providers and characteristics of their relevant cycles

| Country | EQA provider | no. participants1 | date(s) samples shipped to participant laboratories in 2024 (dd/mm) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | S1 | S2 | S3 | S4 |
| Australia | RCPAQAP | 125 | 03.06. | 01.07. | 02.09. | 28.10. |
| Austria | ÖQUASTA | 130 | 03.06. | 10.09. | 10.09. | 03.06. |
| Belgium | Sciensano | 144 | 17.06. | 17.06. | 17.06. | 17.06. |
| France | Biologie Prospective | 117 | 28.05. | 28.05. | 28.05. | 28.05. |
| Germany | INSTAND | 652 | 25.06. | 08.10. | 08.10. | 25.06. |
| Germany | RfB | 432 | 26.06. | 26.06. | 26.06. | 26.06. |
| Italy | AOU Careggi | 38 | 26.06. | 26.06. | 26.06. | 26.06. |
| Norway | Norwegian EQA immunohematology (Nasjonalkontrollen) | 75 | 03.06. | 03.06. | 03.06. | 03.06. |
| South Africa | SANBS | 303 | 07.08. | 09.10. | 06.11. | 04.09. |
| United Kingdom | UK NEQAS | 484 | 2 | 21.10. | 21.10. | 2 |

Footnotes: 1 mean, if different numbers in cycles; 2 samples 1 and 4 were not distributed by this EQA provider

# Supplement 2:

### Table S2: Total anti-D antibody detection rates of participants with individual EQA providers

|  | Sample 1 ∼0.1 IU/mL |  | Sample 2 ∼0.025 IU/mL |  | Sample 3 ∼0.01 IU/mL |  | Sample 4 ∼0.005 IU/mL |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| AOU-Careggi, Italy | 37/37 | 100% | 38/38 | 100% | 12/35 | 34.3% | 1/38 | 2.6% |
| Biologie Prospective, France | 117/117 | 100% | 113/113 | 100% | 74/114 | 64.9% | 41/115 | 35.7% |
| INSTAND, Gernamy | 579/586 | 98.8% | 555/586 | 94.7% | 186/575 | 32.3% | 44/572 | 7.7% |
| Öquasta, Austria | 130/130 | 100% | 122/126 | 96.8% | 43/126 | 34.1% | 3/130 | 2.3% |
| Nasjonalkontrollen, Norway | 71/71 | 100% | 69/71 | 97.2% | 23/72 | 31.9% | 0/71 | 0% |
| RCPAQAP, Australia | 125/125 | 100% | 102/127 | 80.3% | 33/126 | 26.2% | 1/126 | 0.8% |
| RfB, Germany | 431/432 | 99.8% | 413/430 | 96.0% | 114/398 | 28.6% | 22/432 | 4.9% |
| SANBS, South Africa | 234/313 | 74.8% | 149/341 | 43.7% | 0/303 | 0% | 1/303 | 0.3% |
| Sciensano, Belgium | 143/144 | 99.3% | 140/144 | 97.2% | 51/144 | 35.4% | 17/144 | 11.8% |
| UK NEQAS, United Kingdom | (1) |  | 483/484 | 99.8% | 305/482 | 63.3% | (1) |  |

Footnote: (1) sample was not used by the EQA provider

Note: The participating laboratories were not necessarily located exclusively in the country where the EQA provider is based.

# Supplement 3:

### Table S3: Detection rates of anti-D antibody in Samples 1 – 4

|  | Sample 1 ∼0.1 IU/mL |  | Sample 2 ∼0.025 IU/mL |  | Sample 3 ∼0.01 IU/mL |  | Sample 4 ∼0.005 IU/mL |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| All results | 1867/1955 | 95.5% | 2184/2460 | 88.8% | 841/2375 | 35.4% | 130/1931 | 6.7% |
| All automated read results | 507/509 | 99.6% | 922/961 | 95.9% | 429/947 | 45.3% | 44/507 | 8.7% |
| All manual read results | 277/353 | 78.4% | 224/413 | 54.2% | 77/386 | 19.9% | 5/348 | 1.4% |

# Supplement 4:

### Table S4: Anti-D antibody detection rates of different test principles, test systems, and ways of reading reaction results

|  | Sample 1 ∼0.1 IU/mL |  | Sample 2 ∼0.025 IU/mL |  | Sample 3 ∼0.01 IU/mL |  | Sample 4 ∼0.005 IU/mL |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Erythrocyte-magnetized technique (EMT) (Diagast), total | 38/39 | 97.4% | 39/40 | 97.4% | 37/39 | 94.9% | 38/39 | 97.4% |
| EMT (Diagast) - automated | 22/23 | 95.7% | 24/24 | 100% | 23/24 | 95.8% | 23/23 | 100% |
| EMT (Diagast) - manual | 2/2 | 100% | 2/2 | 100% | 2/2 | 100% | 2/2 | 100% |
| Solid phase RBC adherence (SPRCA) technology (Immucor), total | 75/75 | 100% | 113/115 | 98.3% | 99/110 | 90.0% | 30/66 | 45.5% |
| SPRCA (Immucor) - automated | 14/14 | 100% | 51/49 | 96.1% | 46/47 | 97.9% | 9/14 | 64.3% |
| SPRCA (Immucor) - manual | 3/0 | 100% | 1/3 | 33.3% | 0/3 | 0% | 0/2 | 0% |
| Column agglutination technology (CAT), all manufactures, both reading methods | 1527/1539 | 99.2% | 1850/1935 | 95.6% | 640/1886 | 33.9% | 60/1531 | 3.9% |
| Column agglutination technology (CAT), automated reading | 419/420 | 99.8% | 776/806 | 96.3% | 360/787 | 45.7% | 9/515 | 1.7% |
| BioRad | 274/274 | 100% | 398/422 | 94.3% | 182/414 | 44.0% | 9/274 | 3.3% |
| Grifols | 40/40 | 100% | 151/152 | 99.3% | 64/142 | 45.1% | 1/38 | 2.6% |
| OCD | 128/129 | 99.2% | 250/255 | 98.0% | 76/251 | 30.3% | 2/129 | 1.6% |
| Column agglutination technology (CAT), manual reading | 131/134 | 97.8% | 144/154 | 93.5% | 88/161 | 54.7% | 2/133 | 1.5% |
| BioRad | 112/112 | 100% | 119/123 | 96.7% | 46/115 | 40.0% | 2/110 | 1.8% |
| Grifols | 16/16 | 100% | 17/19 | 89.5% | 5/21 | 23.8% | 0/15 | 0% |
| OCD | 13/16 | 81.3% | 18/22 | 81.8% | 10/24 | 41.7% | 0/16 | 0% |
| Test tube technology (TTT) | 107/159 | 67.3% | 35/181 | 19.3% | 1/156 | 0.6% | 0/151 | 0% |

Legend: The difference between “total“ and the sum of “automatically” and “manually” read results is explained by a high number of results for which the type of reading was not disclosed.

# Supplement 5:

### Table S5: Anti-D antibody detection rates of test systems with different EQA providers, automated vs. manual reading of results

| Manufacturer EQA provider |  | Sample 1 ∼0.1 IU/mL |  | Sample 2 ∼0.025 IU/mL |  | Sample 3 ∼0.01 IU/mL |  | Sample 4 ∼0.005 IU/mL |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Diagast (EMT) |  |  |  |  |  |  |  |  |  |
| Biologie Prospective | Automated | 22/22 | 100% | 22/22 | 100% | 22/22 | 100% | 22/22 | 100% |
|  | Manual | 2/2 | 100% | 2/2 | 100% | 2/2 | 100% | 2/2 | 100% |
| SANBS | Automated | 0/1 | 0% | 2/2 | 100% | 0/1 | 0% | 1/1 | 100% |
|  | Manual | - |  | - |  | - |  | - |  |
| Immucor (SPRCA) |  |  |  |  |  |  |  |  |  |
| Biologie Prospective | Automated | 7/7 | 100% | 7/7 | 100% | 7/7 | 100% | 3/7 | 42.9% |
|  | Manual | - |  | - |  | - |  | - |  |
| RCPAQAP | Automated | - |  | - |  | - |  | - |  |
|  | Manual | 3/3 | 100% | 1/3 | 33.3% | 0/3 | 0% | 0/2 | 0% |
| ÖQUASTA | Automated | 1/1 | 100% | - |  | - |  | 1/1 | 100% |
|  | Manual | - |  | - |  | - |  | - |  |
| Sciensano | Automated | 6/6 | 100% | 6/6 | 100% | 6/6 | 100% | 5/6 | 83.3% |
|  | Manual | - |  | - |  | - |  | - |  |
| UK NEQAS | Automated | (1) |  | 35/35 | 100% | 33/34 | 97.1% | (1) |  |
|  | Manual |  |  | - |  | - |  |  |  |
| Bio-Rad (CAT) |  |  |  |  |  |  |  |  |  |
| ÖQUASTA | Automated | 32/32 | 100% | 33/34 | 97.1% | 8/35 | 22.9% | 0/32 | 0% |
|  | Manual | 66/66 | 100% | 55/58 | 94.8% | 18/57 | 31.6% | 2/66 | 3.0% |
| Nasjonalkontrollen | Automated | 18/18 | 100% | 17/18 | 94.4% | 10/18 | 55.6% | 0/18 | 0% |
|  | Manual | 13/13 | 100% | 13/13 | 100% | 4/14 | 28.6% | 0/13 | 0% |
| RCPAQAP | Automated | 25/25 | 100% | 23/25 | 92.0% | 6/25 | 24.0% | 0/25 | 0% |
|  | Manual | 11/11 | 100% | 10/11 | 90.9% | 5/11 | 45.5% | 0/11 | 0% |
| SANBS | Automated | 137/137 | 100% | 125/145 | 86.2% | 0/136 | 0% | 0/136 | 0% |
|  | Manual | 2/2 | 100% | 1/2 | 50.0% | 0/2 | 0% | 0/2 | 0% |
| Sciensano | Automated | 22/22 | 100% | 22/22 | 100% | 18/22 | 81.8% | 6/22 | 27.3% |
|  | Manual | 10/10 | 100% | 10/10 | 100% | 7/10 | 70.0% | 0/12 | 0% |
| UK NEQAS | Automated | (1) |  | 139/139 | 100% | 117/137 | 85.4% | (1) |  |
|  | Manual |  |  | 21/21 | 100% | 10/23 | 43.5% |  |  |
| Biologie Prospective | Automated | 22/22 | 100% | 22/22 | 100% | 13/22 | 59.1% | 0/22 | 0% |
|  | Manual | 6/6 | 100% | 6/6 | 100% | 1/6 | 17.7% | 0/6 | 0% |
| Grifols (CAT) |  |  |  |  |  |  |  |  |  |
| RCPAQAP | Automated | 8/8 | 100% | 7/8 | 87.5% | 3/8 | 37.5% | 0/7 | 0% |
|  | Manual | 14/14 | 100% | 12/14 | 85.7% | 4/13 | 30.8% | 0/13 | 0% |
| Sciensano | Automated | 27/27 | 100% | 27/27 | 100% | 1/27 | 3.7% | 1/27 | 3.7% |
|  | Manual | 2/2 | 100% | 2/2 | 100% | 1/2 | 50.0% | 0/2 | 0% |
| Biologie Prospective | Automated | 4/4 | 100% | 3/3 | 100% | 0/3 | 0% | 0/3 | 0% |
|  | Manual | - |  | - |  | 0/1 | 0% | 0/1 | 0% |
| ÖQUASTA | Automated | 1/1 | 100% | 1/1 | 100% | 0/1 | 0% | 0/1 | 0% |
|  | Manual | - |  | - |  | - |  | - |  |
| UK NEQAS | Automated | (1) |  | 113/113 | 100% | 60/107 | 56.1% | (1) |  |
|  | Manual |  |  | 3/3 | 100% | 0/5 | 0% |  |  |
| OCD (CAT) |  |  |  |  |  |  |  |  |  |
| Biologie Prospective | Automated | 10/10 | 100% | 10/10 | 100% | 5/10 | 50% | 0/10 | 0% |
|  | Manual | 1/1 | 100% | 1/1 | 100% | 0/1 | 0% | 0/1 | 0% |
| Öquasta | Automated | 26/26 | 100% | 29/29 | 100% | 15/29 | 51.7% | 0/26 | 0% |
|  | Manual | 4/4 | 100% | 4/4 | 100% | 2/4 | 50.0% | 0/4 | 0% |
| Nasjonalkontrollen | Automated | 29/29 | 100% | 29/29 | 100% | 5/27 | 18.5% | 0/29 | 0% |
|  | Manual | 1/1 | 100% | 1/1 | 100% | 1/3 | 33.3% | 0/1 | 0% |
| RCPAQAP | Automated | 9/9 | 100% | 7/9 | 77.8% | 2/9 | 22.2% | 0/5 | 0% |
|  | Manual | 5/5 | 100% | 3/5 | 60.0% | 1/5 | 20.0% | 0/5 | 0% |
| Sciensano | Automated | 54/55 | 98.2% | 52/55 | 94.5% | 7/55 | 12.7% | 2/53 | 3.6% |
|  | Manual | 2/2 | 100% | 1/2 | 50.0% | 0/2 | 0% | 0/2 | 0% |
| SANBS | Automated | - |  | - |  | - |  | - |  |
|  | Manual | 0/3 | 0% | 0/1 | 0% | 0/1 | 0% | 0/3 | 0% |
| UK NEQAS | Automated | (1) |  | 123/123 | 100% | 42/121 | 34.7% | (1) |  |
|  | Manual |  |  | 8/8 | 100% | 6/8 | 75.0% |  |  |
| TTT (manual) |  |  |  |  |  |  |  |  |  |
| Nasjonalkontrollen |  | 5/5 | 100% | 4/5 | 80.0% | 0/5 | 0% | 0/5 | 0% |
| RfB |  | 9/9 | 100% | 8/9 | 88.9% | 1/8 | 12.5% | 0/9 | 0% |
| SANBS |  | 93/309 | 30.1% | 21/139 | 15.1% | 0/139 | 0% | 0/137 | 0% |
| UK NEQAS |  | (1) |  | 2/3 | 66.7% | 0/4 | 0% | (1) |  |

Footnote: (1) this sample was not used by the EQA provider

Legend: Differences between total numbers and automated + manual results are due to results that were not accompanied by the information on the way they were obtained; counts of these results are not shown

# Supplement 6: Anti-D antibody detection rates of different test cell preparations

The test cells included those that can be used both for manual testing and as part of automated test systems, such as those from Bio-Rad Laboratories Inc., Hercules, CA, USA (Bio-Rad), QuidelOrtho Athens, OH, USA (Ortho Clinical Diagnostics, OCD), and Diagnostic Grifols, S.A., Barcelona, Spain (Grifols). Other test cells were produced without the manufacturer offering devices intended for use with them, such as those from Immulab Pty Ltd, Mt. Waverley, Victoria, Australia (Immulab), the South African National Blood Service, Roodepoort, Gauteng, South Africa (SANBS), and the National Blood Service Zimbabwe, Harare, Zimbabwe (NBSZ). Finally, test cells were used that could only be used on the associated test system for reasons of the determination principle, such as those from Immucor Inc. Norcross, GA, USA (Immucor) and Diagast, Loos, France (Diagast). The pattern of decreasing detection rates with decreasing anti-D concentrations was also observed in the results of all test cells. (Table S5, Figure S2)

### Table S6: Anti-D antibody detection rates of different test cell preparations

| Manufacturer of test cells | Sample 1 ∼0.1 IU/mL |  | Sample 1 ∼0.025 IU/mL |  | Sample 3 ∼0.01 IU/mL |  | Sample 4 ∼0.005 IU/mL |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bio-Rad Laboratories Inc., Hercules, CA, USA | 1103/1108 | (99.5%) | 1204/1266 | (95.1%) | 434/1232 | (35.2%) | 48/1104 | (4.3%) |
| QuidelOrtho Athens, OH, USA | 346/352 | (98.3%) | 454/473 | (96.0%) | 126/462 | (27.3%) | 4/353 | (1.1%) |
| Diagnostic Grifols, S.A., Barcelona, Spain | 136/137 | (99.3%) | 248/254 | (97.6%) | 86/247 | (34.8%) | 9/131 | (6.9%) |
| South African National Blood Service, Roodepoort, Gauteng, South Africa | 91/123 | (74.0%) | 21/132 | (15.9%) | 1/114 | (0.9%) | 0/125 | (0%) |
| Immucor Inc. Norcross, GA, USA | 75/75 | (100%) | 113/115 | (98.3%) | 99/110 | (90.0%) | 30/66 | (45.5%) |
| Immulab Pty Ltd, Mt. Waverley, Victoria, Australia | 41/41 | (100%) | 30/41 | (73.2%) | 11/40 | (27.5%) | 1/42 | (2.4%) |
| Diagast, Loos, France | 38/39 | (97.4%) | 38/39 | (97.4%) | 37/39 | (94.9%) | 38/39 | (97.4%) |
| National Blood Service Zimbabwe, Harare, Zimbabwe | 2/15 | (13.3%) | 0/15 | (0%) | 0/15 | (0%) | n.d. |  |
| Other (<15 results per manufacturer) in total (1) | 11/18 | (61.1%) | 9/27 | (32.1%) | 3/21 | 14.3%) | 0/24 | (0%) |

Footnote: 1) test cells were manufactured by Haukeland University Hospital, Reagents lab, Affirmagen, NHSBT, Oslo University Hospital, Tulip Diagnostics, NAMBTS, NHS, Ilex, Immuno, The Scientific Group, CNTS de Lomé, Davies Diagnostics, KZN, NDSO, SAB, or were in-house products.

### Figure S6: Detection rates of test cells produced by individual manufacturers

Footnote: (1) These test cells were not used for Sample 4.

# Supplement 7: PubILC

|  |  | Item |
| --- | --- | --- |
| Description of activity (1) |  |  |
|  | (1a) | Description of the activity - The description of the activity should be consistent and unambiguous to make the nature of the activity and the publication reporting on it clear. Activity is described in Materials and Methods section. |
|  | (1b) | Additional information on the activity reported on - More detailed information on the activity and the local regulations for participation in EQA/PT should be described. n/a |
|  | (1c) | Initiator (and executor, if different) of the activity reported on; business purpose of the initiator - The initiator of the ILC activity should be named and their business purpose should be declared. This study was initiated by the Immunohaematology Working Group of EQALM, the umbrella organisation of EQA providers for laboratory medicine; executors were EQA providers identified via the professional network of EQALM; |
|  | (1d) | Purpose of the activity - Purpose of the activity should be clearly defined and reflected in the publication. If an activity is reported with a purpose different from what is ultimately intended (incidental finding), the original purpose and the unexpected outcome should be stated. The purpose of the study was to evaluate and compare the sensitivity for anti-D antibodies of different RBC antibody screening systems in routine use; |
|  | (1e) | Time span of activity - The time span within which the samples were analysed should be provided; this provides information about the prevailing knowledge at the time of the study; the usual methodical performance of the analysis, or the predominant strains and variants of infectious agents that were used as sample material at the time of the study. Dates when samples were shipped to EQA participants are given in Supplement 1. |
|  | (1f) | Number of items (samples) included in the sample panel - Information on the number of samples used in the activity reported on should be provided. The sample panel consisted of four samples positive for anti-D; results obtained in analyses for characterization of the samples are given in Materials and Methods section. |
|  | (1g) | ISO 17043 accreditation status of EQA provider / scheme / parameter - Accreditation ensures that the organizer of the study has a universally accepted and peer-reviewed management system for providing EQA/PT services. Not applicable for a super-challenge, as each EQA provider obtains samples and assesses results according to their procedures. |
| Information on items (samples) used (2) |  |  |
|  | (2a) | Type, origin and manufacturer – type of sample (serum, plasma (with detailed information, e.g., EDTA, lithium heparin), urine, any other substance); If commercial material is used, its manufacturer and detailed information on the product(s) should be provided. Items (samples) used in ILC activities may also be data (e.g., virtual microscopy, data sets). The manufacturer, matrix and material used for spiking are reported in Materials and Methods section. |
|  | (2b) | Ethics for samples with human origin - Result of the referral to the competent ethics committee if substances of human or animal origin were used. The basis for this can be found in the Declaration of Helsinki. n/a |
|  | (2c) | Detailed specification and justification for selecting each sample - It should be stated, for which purpose each sample was included in the sample panel. Samples 1-4 had decreasing concentrations of anti-D and were intended to identify detection rates of test systems at different concentrations. |
|  | (2d) | Matrix and additives - Information on components in sample materials except the measurand should be provided. Matrix is described in Materials and Methods section. |
|  | (2e) | For quantitative analyses: Range of concentration of measurand(s) - It should be reported whether a measurand is in the low - medium/normal - high range and/or around clinical decision limits; report quantitative results preferably in SI units. It is explicitly stated that the samples have very low concentrations of anti-D antibodies and that the purpose of the study is to analyze such samples. |
|  | (2f) | For qualitative analyses: Identity of examinand - The identity of examinands in samples for qualitative analyses should be reported. The identity of the examinand is specified (anti-D antibodies) |
|  | (2g) | Attestation of sample homogeneity - Adequate homogeneity of the measurand(s) in samples is required so that all participants receive the same starting materials for their analyzes and thereby equal opportunities to obtain correct results. Vial-to-vial homogeneity was determined according to applicable international standards and were found adequately homogenous; see Materials and Methods section. |
|  | (2h) | Attestation of sample stability - Adequate stability of the measurand(s) in samples is required so that all participants receive the same starting materials for their analyzes and thereby equal opportunities to obtain correct results. Stability testing procedures are described in. Stability of the measurand was determined according to applicable international standards and were found adequately stable; see Materials and Methods section. |
|  | (2i) | Physical properties of samples / conditions during shipment to participants - Environmental conditions during transportation can affect the properties of materials; measures to prevent this should be described. Storage and transport conditions are described in Materials and Methods section. |
|  | (2j) | Activities required to prepare samples prior to analysis should be described. None; see line 183. |
|  | (2k) | Classification of samples according to the requirement to submit correct results for them to pass the activity - Depending on the design of the EQA/PT activity, it may be possible that reporting correct results for some samples is required, but incorrect results are accepted for other samples to still pass the round. If applicable, the reasons should be stated. It was left to the participant EQA providers to decide whether results reported for S1-S4 were to be assessed and what target they used. |
|  | (2l) | Commutability of sample materials if applicable - If commutability of samples was evaluated, details should be provided [60]. If commutability was not evaluated, it should be stated that no definitive conclusions on accuracy can be made from the comparisons of results obtained with various methods. n/a |
|  | (2m) | Volume per sample - If relevant for the results presented, the volume of each sample provided to the participants should be presented. The samples were filled in the usual quantities of the EQA providers. |
| Information and instructions to participants (3) |  |  |
|  | (3a) | Statement how participants were instructed to analyse samples - Participants should use their unmodified routine procedures when analysing EQA/PT samples (as far as practicable). It should be stated that participants were made aware of this. However, in some challenges it might be required to deviate from routine procedures and for example analyse materials in triplicate. Information is given in Materials and Methods section. |
| Information on participant entities (4) |  |  |
|  | (4a) | Number of participant entities / respondents - The number of entities registered for the challenge and the rate of respondents should be provided. Given in Results section. |
|  | (4b) | Types of participant entities - Information on the participant laboratories should be provided. n/a |
|  | (4c) | Location (state(s), country, region) - The geographical place where the EQA/PT activity took place should be described. Given in Supplement 1. |
| Information on participating measurement procedures (5) |  |  |
|  | (5a) | Trade name of devices and reagents used (measurement system) - Participant test systems (devices and reagents) should be identified; regarding an international readership, ideally including Unique Device Identifiers (UDI) Manufacturers of antibody screening test systems are named in the paper. |
|  | (5b) | Methods / techniques applied by the measurement systems - If relevant for the results presented, methods applied by test systems (5a) should be described. Methods are described in the paper. |
|  | (5c) | Lot numbers of reagents - If relevant for the results presented, Lot numbers of reagents should be reported. n/a |
| Submission of results (6) |  |  |
|  | (6a) | Possible ways provided by the EQA/PT organizer to report results - It should be described how results are reported to the EQA/PT provider/organizer. Contact remained with the individual EQA providers, who collected and reported the data to the study centre. |
|  | (6b) | Format, unitage (for quantitative results) and - if applicable - factorized results - It should be described whether results are reported by selecting from a dropdown menu, entering free text (incl. numeric results) and whether quantitative results are to be reported in a specified unit or by selection of units (e.g. from a drop-down menu); preferably SI units should be used. n/a |
|  | (6c) | Applicable limitations for acceptance of results - If applicable, limitations for submission of results should be reported. n/a |
| Evaluation and assessment of results (7) |  |  |
|  | (7a) | Determination of the target /assigned value (including uncertainty of the target value) - The traceability of target values should be described, e.g., reference measurement procedure, certified reference materials, overall mean/median, peer group, consensus values, expert laboratories; including - if applicable - uncertainty of the target value. The specified methods for determining the targets / target values correspond to requirements in ISO 13528:2022. If the targets are determined by one or several laboratories, the methods used by them should be described (they may differ from the methods used by participant laboratories). A WHO Standard was used to spike EQA samples. Concentration of measurand was measured in an expert laboratory. |
|  | (7b) | Acceptance criteria - The basis for deciding which results are accepted (pass) and which are not (fail); analytical performance specifications (APS) and - if applicable - APS per intended use of the test should be described. Determination of acceptance criteria and assessment of results remained with the individual EQA providers. |
|  | (7c) | Outlier removal procedures employed - It should be described how the statistical model deals with outliers. n/a |
|  | (7d) | Statistics applied - Basic and advanced statistics applied in routine assessment of results. Statistical methods used are described in the Materials and Methods section. |
| Reporting of assessment results (8) |  |  |
|  | (8a) | to participants - The way participants are informed about their performance should be described. Remained with the individual EQA providers. |
|  | (8b) | to other interested parties - If applicable, the way of notifying other interested parties should be presented. Remained with the individual EQA providers. |
| Findings (9) |  |  |
|  | (9) | Information gained according to the purpose of the activity - Study results, answering the study question, observed variability between laboratories and potential clinical impact. Study objectives and outcome are clearly described; the potential clinical impact needs to be evaluated. |
| Limitations (10) |  |  |
|  | (10) | EQA/PT related limitations - A limitation of EQA/PT studies may be that results can only be analysed as they were reported by participants. It must be trusted that they were generated properly and with the test system specified. Included in the paper. |
| Impact of the outcome of the activity (11) |  |  |
|  | (11a) | Educational aspects for participants, clinical impact of the findings - Educational aspects can be an integral part of ILC activities and publications about them. If the activity being reported has such aspects, they should be mentioned in the publication. The entire paper is educational. |
|  | (11b) | Knowledge gain for other interested parties - If applicable, information revealed by the study and relevant for other interested parties should be made clear. Authorities and manufacturers could become aware of the current disregard of the limits of detection of RBC antibody screening tests. |
|  | (11c) | Identified areas for improvement (for the EQA/PT provider) - If applicable, improvement potentials revealed by the study should be reported. Attention of European Reference Laboratories, which will be responsible for the approval of individual lots of RBC antibody screening tests. |