## CRRT Protcol: Citrate Anticoagulation (PICU Use only)

## Prescriptions Flow Diagram

Decision to perform CRRT

Default Anticoagulant is CITRATE, see page 14

Draw up 50mL into syringe

Always use Prismaflex syringe driver & specific calcium line

NEAT: Calcium Gluconate

0.22mmol/mL

Haemofilter/Circuit Size

see Page 12

Check Patients ionized Calcium and haematocrit

(By ABG or VBG)

Heparin

Check with Intensivist

Use Heparin protocol

5L bag placed on PBP scale

Prismocitrate 18/0

Check circuit volume if>10% of patient’s blood volume consider a blood transfusion as the patient goes on to CRRT.

- Check min 1 unit blood available

Patient’s Blood Volume

= 80mL/kg

3-8kg

Consider

GamCath

6.5FG Double Lumen

>15kg

GamCath

11FG Double Lumen

Or 12FG Triple Lumen

8-15kg

Consider

GamCath

8FG Double Lumen

Venous Catheter, See Table 1 on page 11

Default

CVVHDF

<15kg

HF20

Circuit Vol 60mL

15-30kg

ST60

Circuit Vol 93mL

>30kg

ST100

Circuit Vol 152mL

Contraindications to Citrate

- Severe Hepatic Failure

- Severe Hypocalcaemia (Correct before starting Citrate)

- Citrate Allergy

###### Flow Rates: Use the following for Citrate Based Anticoagulant

|  | Fluid |  |  | Prism0cal B22 | 0.9% saline | Prism0cal B22 |
| --- | --- | --- | --- | --- | --- | --- |
|  | Blood |  |  | Dialysate | Post-filter Replacement | Post-filter Replacement |
| Weight (kg) | Default mL/min |  |  | Default mL/hr | Default mL/hr | Optional* mL/hr |
| 2 ─ 6 | 20 |  |  | 150 ─ 300 | 20 | 200 |
| 6 ─ 7 | 22 |  |  | 150 ─ 350 | 20 | 220 |
| 7 ─ 8 | 26 |  |  | 250 ─ 400 | 20 | 260 |
| 8 ─ 9 | 28 |  |  | 250 ─ 400 | 20 | 280 |
| 9 ─ 10 | 30 |  |  | 300 ─ 450 | 20 | 300 |
| 10 ─ 11 | 34 |  |  | 300 ─ 500 | 40 | 340 |
| 11 ─ 12 | #36 |  |  | 350 ─ 550 | 40 | 360 |
| 12 ─ 13 | #40 |  |  | 400 ─ 600 | 40 | 400 |
| 13 ─ 14 | #44 |  |  | 400 ─ 700 | 100 | 440 |
| 14 ─ 15 | #46 |  |  | 450 ─ 700 | 100 | 460 |
|  | #50min |  |  | if ST 60 | is used |  |
| 15 ─ 16 | 50 |  |  | 500 ─ 800 | 100 | 500 |
| 16 ─ 18 | 55 |  |  | 550 ─ 850 | 100 | 550 |
| 18 ─ 20 | 60 |  |  | 600 ─ 950 | 100 | 600 |
| 20 ─ 21 | 65 |  |  | 650 ─ 1000 | 100 | 650 |
| 21 ─ 23 | 70 |  |  | 700 ─ 1100 | 100 | 700 |
| 23 ─ 25 | 75 |  |  | 750 ─ 1200 | 100 | 750 |
| 25 ─ 26 | 80 |  |  | 800 ─ 1250 | 100 | 800 |
| 26 ─ 28 | 85 |  |  | 850 ─ 1350 | 100 | 850 |
| 28 ─ 30 | 90 |  |  | 900 ─ 1400 | 100 | 900 |
| 30 ─ 36 | 110 |  |  | 1100 ─ 1700 | 100 | 1100 |
| 36 ─ 40 | 120 |  |  | 1200 ─ 1900 | 100 | 1200 |
| 40 ─ 43 | 130 |  |  | 1300 ─ 2000 | 100 | 1300 |
| 43 ─ 46 | 140 |  |  | 1400 ─ 2200 | 100 | 1400 |
| 46 ─ 50 | 150 |  |  | 1500 ─ 2400 | 100 | 1500 |
| 50 ─ 53 | 160 |  |  | 1600 ─ 2500 | 100 | 1600 |
| 53 ─ 56 | 170 |  |  | 1700 ─ 2700 | 100 | 1700 |
| 56 ─ 60 | 180 |  |  | 1800 ─ 2800 | 100 | 1800 |
| 60 ─ 63 | 190 |  |  | 1900 ─ 3000 | 100 | 1900 |
| 63 ─ 66 | 200 |  |  | 2000 ─ 3200 | 100 | 2000 |
| >66 | 210 |  |  | 2100 ─ 3300 | 100 | 2100 |

# Recommended circuit for <15 kg is HF 20. If not available and only ST 60 is available PICU, Intensivist must authorise use of the ST 60 set for Pt 11-15 kg, as min blood flow is 50mL/min.

NOTE: Pre Blood Pump (PBP) = flow rates are dependent on citrate dosing therefore set by the Prismaflex machine, fluid used is Prismocitrate 18/0 (5L bag). CVVH with different flows may be ordered for patients >10kg but this must be ordered by PICU intensivist.

## Ongoing Monitoring and Action of Patient and Circuit Ionized Calcium

Target patient ionized calcium level is 1.0-1.1mmol/L

| Patient iCa2+(mmol/L) | < 0.8 | 0.8 - 1.0 | TARGET 1.0 – 1.1 | >1.1 |
| --- | --- | --- | --- | --- |
| Calcium compensation adjustment | Follow hypocalcaemia flowchart (page 82) | Increase current Ca Compensation by 10% Recheck bloods in 1 hr Monitor pH & Anion Gap | No changes required | Decrease current Ca compensation by 10% Recheck bloods in 1 hr |

Target Circuit pre filter ionized calcium is 0.3 – 0.44mmol/L

| Pre filter iCa2+(mmol/L) | < 0.2 | 0.2 - 0.29 | TARGET 0.3 - 0.44 | 0.45 - 0.5 | > 0.5 |
| --- | --- | --- | --- | --- | --- |
| Citrate dose adjustment | Decrease citrate dose by 0.4 Recheck ABG in 1 hr | Decrease citrate dose by 0.2 Recheck ABG in 1 hr | No changes required | Increase citrate dose by 0.1 Recheck ABG in 1 hr | Increase citrate dose by 0.3 Recheck ABG in 1 hr |

## Emergency Kit

Emergency Kit

- This must be prepared before commencing CRRT, ready for any emergency or disconnection. Ensure all items in box are present and expiry dates checked. It should be kept in at the bedside at all times. Includes a printed copy of MANUAL TERMINATION OF TREATMENT from procedure.

- Y line × 1

- Gambro spike × 1

- Chlorhexidine 2% in 70% alcohol (Chlorhexidine 0.1% if < 2 months of age)

- Dressing pack

- Sterile towels x 2

- Sterile gauze squares × 4

- Sterile gloves

- Syringes 2mL leur lock × 2 (These are for injecting accurate amounts of heparin to flush lumens. NOT to be used for attempts at unblocking)

- Syringes 10mL leur lock × 4

- Drawing up needles x 4

- Interlink bungs x 2

- Interlink blunt cannula x 2

- Plastic clamps x 2

- Medication added label x 2

- Catheter lock label x 3

- Medi-sponge x 1

Please collect and add the following:

- Calcium gluconate vial (you will need to collect this from drug cupboard)

- IV guidelines for Calcium gluconate

- 1 litre bag of 0.9% Normal Saline (required if you need to perform a blood return to patient)

- 100mL bag of 0.9 % Normal Saline (required if you need to perform a blood recirculation)

- Ampoules of 0.9% Normal Saline 10mL x 2

- Ampoule of Heparin Sodium 10units/mL x 2

Calcium Syringe

Calcium Gluconate:

50 ml of neat solution 2.2mmol/10mL

Citrate Anticoagulation

Citrate Anticoagulation

Check patient iCa2+ prior to commencing CRRT

Blood Sampling

Sample Sites:

- always need 2 samples

  - Circuit- via pre filter

  - Patient- via arterial/venous blood

Frequency * wait at least 1 hour between samples*

- 1/24 samples until stable

(3 consecutive readings)

- Then 2/24 for 4 hours

- 6/24 thereafter

**any abnormal results must return to 1/24 sampling**

Step 1. Check patient iCa2+

Patient iCa2+   < 0.8 mmol/L

DO NOT Commence dialysis.  Must correct first.  Give 0.11mmol/kg of 10% Calcium Gluconate.  (Max 4.5mmol/dose)

.

Patient iCa2+   ≥ 0.8 mmol/L

Continue to CRRT set up & settings

Step 2. Set Citrate dose & Calcium Compensation

| Patient iCa2+ | Citrate Dosing mmol/L |
| --- | --- |
| < 0.9 | 2.0 |
| 0.9 – 1.10 | 2.5 |
| 1.11 -1.3 | 3.0 |
| >1.3 (double check) | 3.4 |

| Patient iCa2+ mmol/L | Calcium Compensation % |
| --- | --- |
| < 1.0 | 100 |
| ≥ 1.0 | 90 |

Step 3. Ongoing Monitoring & Action

| < 0.8 | 0.8 - 1.0 | TARGET 1.0 – 1.1 | >1.1 |
| --- | --- | --- | --- |
| Follow hypocalcaemia Flowchart SEE BACK OF CHART!! | Increase current Ca Compensation by 10% Recheck bloods in 1 hr Monitor pH & Anion Gap | No changes required | Decrease current Ca compensation by 10% Recheck bloods in 1 hr |

| < 0.2 | 0.2 - 0.29 | TARGET 0.3 - 0.44 | 0.45 - 0.5 | > 0.5 |
| --- | --- | --- | --- | --- |
| Decrease citrate dose by 0.4 Recheck ABG in 1 hr | Decrease citrate dose by 0.2 Recheck ABG in 1 hr | No changes required | Increase citrate dose by 0.1 Recheck ABG in 1 hr | Increase citrate dose by 0.3 Recheck ABG in 1 hr |

** If iCa2+ consistently drops, check infusion & connections.  Consider Citrate Toxicity **

Patient Arterial iCa2+									Circuit iCa2+

## Appendix: CRRT Heparin Anticoagulation (PICU use only)

## Heparin Prescriptions Flow Diagram

Decision to perform CRRT

Venous Catheter, See table 1

3-8 kg

Consider

GamCath

6.5FG Double Lumen

>15 kg

GamCath

11FG Double Lumen or 12FG Triple Lumen

8-15 kg

Consider

GamCath

8FG Double Lumen

Haemofilter/Circuit Size

>30 kg

ST100

Circuit Vol 152mL

<15 kg

HF20

Circuit Vol 60mL

15-30 kg

ST60

Circuit Vol 93mL

No-anticoagulant

Check with Intensivist

Epoprostenol

(Prostacyclin)

If heparin contraindicated or inadequate,

Check with Intensivist

Heparin

Note: Only if anti-coagulant is needed

Check with Intensivist

Patient’s Blood Volume

= 80mL/kg

Check circuit volume if>10% of patient’s blood volume consider a blood transfusion as the patient goes on to CRRT.

- Check minimum 1 unit blood available

Anticoagulant,

Citrate anticoagulant is default (please use citrate protocol).

However if there is contraindications to citrate use Heparin.

If ACT < 210 seconds

Commence heparin infusion at 10units/kg/hr to a max of 50 kg.

≤ 50 kg: 250 units/kg in 50 mL 0.9% NaCl

> 50 kg: 12500 units in 50mL 0.9% NaCl

## Flow Rates: Use the following for Heparin Based Anticoagulant

SCUF:	This table does not apply to SCUF

Choose 3-4mL/kg/min blood flow (Minimum 20mL/min)

Zero replacement.

Recommended circuit for <15 kg is HF 20. If not available and only ST 60 is available PICU Intensivist must authorise use of the ST 60 set for Pt 11-15 kg as min blood flow is 50mL/min & PBP flow is 600mL/hr

| Weight (kg) | Blood | Pre-Blood-Pump (PBP) | Dialysate | Dialysate | Post-filter Replacement | Post-filter Replacement |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Hemosol B0 | Hemosol B0 | Hemosol B0 | Hemosol B0 | Hemosol B0 |
|  | Default mL/min | Default mL/hr | Default mL/hr | Optional mL/hr | Default mL/hr | Optional mL/hr |
| 2 ─ 6 | 20 | 240 | 0 | 200 ─ 300 | 20 | 200 |
| 6 ─ 7 | 22 | 270 | 0 | 200 ─ 350 | 20 | 220 |
| 7 ─ 8 | 26 | 320 | 0 | 250 ─ 400 | 20 | 260 |
| 8 ─ 9 | 28 | 340 | 0 | 250 ─ 400 | 20 | 280 |
| 9 ─ 10 | 30 | 360 | 0 | 300 ─ 450 | 20 | 300 |
| 10 ─ 11 | 34 | 410 | 0 | 300 ─ 500 | 40 | 340 |
| 11 ─ 12 | #36 | 440 | 0 | 350 ─ 550 | 40 | 360 |
| 12 ─ 13 | #40 | 480 | 0 | 400 ─ 600 | 40 | 400 |
| 13 ─ 14 | #44 | 530 | 0 | 400 ─ 700 | 100 | 440 |
| 14 ─ 15 | #46 | 560 | 0 | 450 ─ 700 | 100 | 460 |
|  | #50 | and 600 | For PBP | if ST 60 | is used |  |
| 15 ─ 16 | 50 | 600 | 0 | 500 ─ 800 | 100 | 500 |
| 16 ─ 18 | 55 | 700 | 0 | 550 ─ 850 | 100 | 550 |
| 18 ─ 20 | 60 | 750 | 0 | 600 ─ 950 | 100 | 600 |
| 20 ─ 21 | 65 | 800 | 0 | 650 ─ 1000 | 100 | 650 |
| 21 ─ 23 | 70 | 850 | 0 | 700 ─ 1100 | 100 | 700 |
| 23 ─ 25 | 75 | 900 | 0 | 750 ─ 1200 | 100 | 750 |
| 25 ─ 26 | 80 | 1000 | 0 | 800 ─ 1250 | 100 | 800 |
| 26 ─ 28 | 85 | 1050 | 0 | 850 ─ 1350 | 100 | 850 |
| 28 ─ 30 | 90 | 1100 | 0 | 900 ─ 1400 | 100 | 900 |
| 30 ─ 36 | 110 | 1400 | 0 | 1100 ─ 1700 | 100 | 1100 |
| 36 ─ 40 | 120 | 1500 | 0 | 1200 ─ 1900 | 100 | 1200 |
| 40 ─ 43 | 130 | 1600 | 0 | 1300 ─ 2000 | 100 | 1300 |
| 43 ─ 46 | 140 | 1700 | 0 | 1400 ─ 2200 | 100 | 1400 |
| 46 ─ 50 | 150 | 1800 | 0 | 1500 ─ 2400 | 100 | 1500 |
| 50 ─ 53 | 160 | 2000 | 0 | 1600 ─ 2500 | 100 | 1600 |
| 53 ─ 56 | 170 | 2100 | 0 | 1700 ─ 2700 | 100 | 1700 |
| 56 ─ 60 | 180 | 2200 | 0 | 1800 ─ 2800 | 100 | 1800 |
| 60 ─ 63 | 190 | 2300 | 0 | 1900 ─ 3000 | 100 | 1900 |
| 63 ─ 66 | 200 | 2400 | 0 | 2000 ─ 3200 | 100 | 2000 |
| >66 | 210 | 2600 | 0 | 2100 ─ 3300 | 100 | 2100 |

## The decision to start CRRT

Prescriptions

All CRRT prescriptions must be made on a CRRT Prescription Chart – Heparin

A new prescription sheet must be written every 24 hours.

Every change in the prescription including Heparin infusion rate changes must be entered and signed on the prescription sheet.

Every new circuit must be recorded on the DPCCM Troubleshooting Log Form and in the patient’s case notes by the CRRT nurse. Recording date and time CRRT ceased and recommenced, including reason for circuit change.

Choosing Heparin as Anticoagulant

Citrate anticoagulation is the default

(Please use citrate protocol)

However if there is contraindications to citrate then use heparin.

Always discuss with the Intensivist which anticoagulant to use.

If there are contraindications to heparin then consider Epoprostenol (Prostacyclin) (see page 55).

STANDARD DILUTION

| 50 kg | 250 units/kg of heparin sodium in 50mL 0.9% NaCl Infusion @1mL/hr =5 units/kg/hr |
| --- | --- |
| If > 50 kg | 12,500 units of heparin sodium in 50mL 0.9% NaCl Infusion @1mL/hr =250 units/hr |

ALTERNATIVE DILUTION

(for patients requiring lower amounts of heparin to maintain target ACT)

| 50 kg | 100 units/kg of heparin sodium in 50 mL 0.9% NaCl Infusion @1mL/hr =2 units/kg/hr |
| --- | --- |
| If > 50 kg | 5000 units of heparin sodium in 50 mL 0.9% NaCl Infusion @1mL/hr =100 units/hr |

The smallest bolus volume allowed by the syringe pump is a 2mL bolus.

The lowest infusion rate (other than zero) is 2mL/hr.

Flow chart : Heparin Anticoagulant CRRT Circuit Guide

COMMENCING CONTINUOUS HEPARIN INFUSION RATE

(SEE ADJACENT TABLE FOR DILUTANT OPTIONS)

ANTICOAGULATION SYRINGE

*For all patients greater than 50 kg, calculate heparin units on a max weight limit of 50 kg.

STANDARD DILUTION

< 50 kg

250 units/kg of heparin in 50mL 0.9% Normal Saline.  @ 1ml = 5 units/kg

> 50 kg

Maximum dosage of 12500 heparin units in 50mL 0.9% Normal Saline @ 1mL = 250 units

ALTERNATIVE DILUTION

(if smaller heparin dose required)

<50kg

100 units/kg of heparin in 50mL 0.9% Normal Saline

@ 1mL =2 units /kg

> 50 kg

Maximum dosage of 5000 heparin units in 50mL 0.9% Normal Saline @ 1ml = 100 units

NB Syringe Driver on PRISMAFLEX

- Lowest infusion rate 2mL/hr

- Lowest Bolus volume 2mL.

ACT SAMPLING

Sampling is hourly, then 4 hourly once stabilised, as documented by MO

If no anticoagulation infusion, ACT checking may not be required.

Please discuss with PICU consultant.

Sample site default is ARTERIAL LINE

Blood may be taken from CVAD or first access port closest to patient  on circuit ONLY after discussion with PICU consultant.

Dead space clearance

< 10 kg Remove 5mL

> 10kg Remove 5mL -10mL

Remember to ¾ turn tap to prevent heparin contamination of sample

ACT < 160 SECONDS

ONCE CRRT COMMENCED AND DESIRED FLOW RATE IS REACHED

GIVE BOLUS 25 UNITS/KG to a max of 50 kg

INTO THE CIRCUIT PROVIDED NO ACTIVE BLEEDING SEEN

ACT > 160 SECONDS

NO BOLUS

STEP 2: BOLUS

STEP 1: INFUSION

STEP 3 ONGOING MONITORING

IF ACT IN STEP 1 < 210 SECONDS

1st Circuit setup

COMMENCE HEPARIN INFUSION AT THE INITIAL RATE OF 10 UNITS/KG/HR to a max of 50kg WHEN COMMENCING CRRT

Subsequent Circuit setups

CONSIDER previous heparin infusion rate and adjust commencement rate accordingly

IF ACT IN STEP 1 > 210 SECONDS

DO NOT COMMENCE INFUSION

REPEAT ACT IN 1 HR THEN

FOLLOW APPROPRIATE BOX BELOW

ACT MONITORING

ACT   180 - 210 SECONDS

AND NO ACTIVE BLEEDING

MAINTAIN HEPARIN INFUSION RATE

OR

COMMENCE HEPARIN INFUSION AT A

RATE OF 5 UNITS/KG/HR

TO A MAX 50 KG

OR

MININUM INFUSION RATE

ACT < 160 SECONDS

AND NO ACTIVE BLEEDING

COMMENCE/INCREASE HEPARIN INFUSION IN

5 UNITS/KG/HR INCREMENTS

(A BOLUS OF

HEPARIN MAY BE ORDERED IN CONSULTATION WITH INTENSIVIST)

(> 50 KG Wt = MAX 250 UNITS)

ACT 160 -180 SECONDS

AND NO ACTIVE BLEEDING

COMMENCE/INCREASE HEPARIN INFUSION

2 -5 UNITS/KG/HR INCREMENTS

TO A MAX WEIGHT 50 KG

ACT 210 -240 SECONDS

AND NO ACTIVE BLEEDING

1. NOTIFY MEDICAL OFFICER

2. CONSIDER SAMPLE ERROR, RECHECK ACT, IF SAME CONSIDER OPTIONS BELOW

3. CONSIDER AND CHOOSE ONE OPTION

- DECREASING RATE by  2 -5 UNITS/KG/HR

- CHANGING CONCENTRATION TO LOWER DOSE

- CEASING HEPARIN INFUSION FOR 1 HR

4. RECHECK ACT IN 1 HOUR

IF > 240 SECS SEE APPROPRIATE BOX

IF STILL 210 -240 SECS REPEAT PROCESS ABOVE

IF ACT 180 -210 SECS SEE APPROPRIATE BOX

IF ACT <180 SECS RECHECK SAMPLE AND SEE APPROPRIATE BOX

ACT > 240 SECONDS

1. NOTIFY MEDICAL OFFICER

2. CEASE HEPARIN INFUSION

3. RECHECK ACT HOURLY UNTIL < 240 SECS

4. ONCE ACT

LESS THAN 240 SECONDS THEN FOLLOW APPROPRIATE BOX.