# Appendix: Evaluation of spectral transmittance of Bangerter 1.0 and 0.8 filters

To evaluate the possibility that the spectral transmittance of the Bangerter filters might be significantly non uniform and might influence the results of a colour naming experiment, the following examination methods were applied

## Method

The spectral transmittances of the two Bangerter filters were measured and analysed according to the method that is common to ISO 12609-1[25] and AS 16321.4 [54] and validated as providing no significant interference with clinical observation of the colour of patients.[23] This is the most stringent requirement used in eye and face protection standards and substantially more stringent than the colour signal detection criterion found in sunglass, eye and face protection and ophthalmic lens standards.

The total and diffuse spectral transmittances were measured using an Agilent (Santa Clara, CA, USA) Varian Cary 5000 UV-Vis-NIR dual beam, double monochromator spectrophotometer (ORLAB ID #278) fitted with a diffuse reflectance attachment (ORLAB ID #137). The wavelength accuracy has been established with spectral lines as ±0.1 nm.  The uncertainty of measurement (UNC) of spectral transmittance has been established with filters for which the spectral transmittances have been provided by the National Research Council of Canada.

UNC(λ)%=0.1+0.002∙τλ+0.1∙dτdλ

Where	t(l) is the % spectral transmittance at wavelength l

Measurements were made at 1 nm intervals with a spectral half-power bandwidth of 1nm.

The mean and minimum spectral transmittance (450-650 nm) and the minimum to mean ratio were calculated.

The haze of each filter was measured using a haze meter (ORLAB #).  This provides haze and luminous transmittance measurements for CIE Illuminant D65 (ISO 14782 [55]) and CIE Illuminant A (ASTM D1003 [56]).  The meter has been calibrated using a set of five haze reference artefacts for which haze values (Illuminants A, C and D65) have been provided by the National Physical Laboratory, UK. The uncertainty of measurement is ±0.2% absolute.

ORLAB is ISO/IEC 17025 [32] accredited by the National Association of Testing Authorities for these measurements.

## Results

Results for mean and minimum spectral transmittance and the ratio of the two with a compliance statement are set out in Table A.

Table A:  Spectral transmittance characteristics of the Bangerter filters 1.0 and 0.8 and the coloration requirements of ISO 12069 [25] and AS 16321.4. [53]

|  | Bangerter filter 1.0 |  |  | Bangerter filter 0.8 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | % spectral transmittance in range 450-650 nm |  |  |  |  |  |
| Measure | Total | Diffuse | Direct | Total | Diffuse | Direct |
| Mean | 87.6 | 2.2 | 85.4 | 87.9 | 5.0 | 82.9 |
| Minimum | 86.4 | 2.0 | 83.9 | 86.6 | 4.8 | 81.2 |
| Minimum/Mean | 0.99 | 0.94 | 0.98 | 0.99 | 0.97 | 0.98 |
| Requirement (ISO 12609-1 [25] and AS 16321.4) [53] | ≥ 0.80 |  |  |  |  |  |
| Compliance | Pass | Pass | Pass | Pass | Pass | Pass |

Results for haze of the two Bangerter filters are set out in Table B

Table B: Wide angle scatter (haze) measurements of the Bangerter filters 1.0 and 0.8.

| Illuminant | Measure | Bangerter filter 1.0 | Bangerter filter 0.8 |
| --- | --- | --- | --- |
| D65 | Haze (wide angle scatter) % | 2.8 | 9.5 |
|  | Luminous transmittance % | 90.8 | 90.9 |
| A | Haze (wide angle scatter %) | 2.8 | 9.4 |
|  | Luminous transmittance % | 90.9 | 91.0 |

## Discussion

The ratios of minimum to mean transmittance (total, diffuse and direct, 450-650 nm) are very close to 1.0 and easily comply with the ISO 12609 and AS 16321.4.

The haze and luminous transmittances to CIE Illuminants A (a relatively yellow illuminant) and D65 (a relatively blue illuminant) are, within the uncertainty of measurement, the same.  This is a further illustration of the neutrality of transmittance of the Bangerter filters.

## Conclusion

The spectral transmittance characteristics Bangerter filters will not cause any significant effect on the visibility of the red, yellow and green signals.