System suitability:

Table S1: System Suitability Parameters

| S. No. | System suitability parameter | Result |
| --- | --- | --- |
| 1 | Theoretical Plates | 3604 |
| 2 | Tailing Factor | 1.18 |
| 3 | % RSD | 0.10 |

Table S2: System Suitability Areas

| Injection No. | Area (mAU) |
| --- | --- |
| 1 | 2142870 |
| 2 | 2142653 |
| 3 | 2142869 |
| 4 | 2144817 |
| 5 | 2147595 |
| Average | 2144160.68 |
| STDEV | 2111.42 |
| % RSD | 0.10 |

System Precision

Table S3: System Precision Areas

| Injection No. | Area (mAU) |
| --- | --- |
| 1 | 2142870 |
| 2 | 2142653 |
| 3 | 2142869 |
| 4 | 2144817 |
| 5 | 2147595 |
| 6 | 2145009 |
| Average | 2144302.14 |
| STDEV | 1920.03 |
| % RSD | 0.09 |

Specificity:

Table S4: Absolute difference between % Assay during specificity

| Name of the sample | Assay (% w/w) of Rotigotine |
| --- | --- |
| As such sample | 100.53 |
| Spiked sample | 99.86 |
| Difference | 0.67 |

The absolute difference between the % Assay in such sample and the spiked sample is less than ±2.0%.

Ruggedness (Intermediate Precision):

Table S5: Details of Variations during Ruggedness

| S. No. | Changes | Method Precision | Ruggedness |
| --- | --- | --- | --- |
| 1 | Triethylamine | Merck (B. No. TF1Q786019) | Merck (B. No. TF1Q786019) |
| 2 | Orthophosphoric Acid | Merck (B. No. CF0C700316) | Merck (B. No. CB0C700056) |
| 3 | Methanol | Rankem (B. No. S176F76) | Merck (B. No. TM1ER24623) |
| 4 | Acetonitrile | Rankem (B. No. TO66G20) | Rankem (B. No. T038F20) |
| 5 | Date | 11 MAR 2024 | 08 APR 2024 |
| 6 | Column Details | LCC017/23 | LCC010/23 |
| 7 | Instrument Details | HPLC007 | HPLC019 |

Table S6: System Suitability Parameters during Ruggedness

| S. No. | System suitability parameter | Result |
| --- | --- | --- |
| 1 | Theoretical Plates | 3812 |
| 2 | Tailing Factor | 1.21 |
| 3 | % RSD | 0.12 |

Table S7: System Suitability Areas for Ruggedness

| Injection No. | Area (mAU) |
| --- | --- |
| 1 | 2132857 |
| 2 | 2134772 |
| 3 | 2130727 |
| 4 | 2134796 |
| 5 | 2137561 |
| Average | 2134142.51 |
| STDEV | 2540.94 |
| % RSD | 0.12 |

Table S8: Results of Method Precision & Intermediate Precision

| Injection No. | Assay (%w/w) in Method Precision | Assay (%w/w) in Intermediate Precision |
| --- | --- | --- |
| 1 | 99.86 | 99.54 |
| 2 | 100.27 | 99.34 |
| 3 | 99.67 | 99.66 |
| 4 | 100.07 | 99.75 |
| 5 | 100.20 | 99.35 |
| 6 | 99.72 | 100.02 |
| Average | 99.97 | 99.61 |
| STDEV | 0.25 | 0.26 |
| %RSD | 0.25 | 0.26 |
| Overall Average | 99.79 |  |
| Overall SD | 0.31 |  |
| Overall RSD | 0.31 |  |

Linearity:

Figure S1: Calibration curve of Rotigotine

Robustness:

Table S9: Parameters varied during Robustness

| Parameter | Actual | Low | High |
| --- | --- | --- | --- |
| Column Temperature | 40°C | 38°C | 42°C |
| Flow rate | 1.0 mL/min | 0.9 mL/min | 1.1 mL/min |

Table S10: System Suitability Parameters observed during variation in column oven temp.

| S. No | System Suitability Parameters | Actual Column Temperature (40°C) | Low Column Temperature (38°C) | High Column Temperature (42°C) |
| --- | --- | --- | --- | --- |
| 1 | Theoretical Plates | 3812 | 4529 | 2984 |
| 2 | Tailing Factor | 1.21 | 1.26 | 1.23 |
| 3 | %RSD | 0.12 | 0.11 | 0.19 |

Table S11: System Suitability Parameters observed during variation in Flow Rate

| S. No | System Suitability Parameters | Actual Flow Rate (0.8 mL/min) | Low Flow Rate (0.7 mL/min) | High Flow Rate (0.9 mL/min) |
| --- | --- | --- | --- | --- |
| 1 | Theoretical Plates | 3812 | 4312 | 2874 |
| 2 | Tailing Factor | 1.21 | 1.30 | 1.24 |
| 3 | %RSD | 0.12 | 0.13 | 0.16 |

Bench Top Solution Stability of Standard and Sample Solutions:

Table S12: Summary Results for solution stability study of standard solution and sample solution for Rotigotine

| Time Interval | Similarity Factor for Standard | Sample Assay (%w/w) | Difference in %Assay of Sample |
| --- | --- | --- | --- |
| Initial | Not applicable | 99.50 | Not applicable |
| 48 hours at RT | 1.00 | 98.99 | 0.51 |

Bench Top Stability of Mobile Phase:

Table S13: Summary Results for Mobile phase Stability Study

| System suitability parameter | Result |
| --- | --- |
| Theoretical plates | 3931 |
| Tailing factor | 1.14 |
| % Relative standard deviation (RSD) | 0.12 |

Filter Validation:

Table S14: Filter Details

| Filter description | Filter |  |
| --- | --- | --- |
|  | Nylon | PVDF |
| Manufacturer Name | mdi Membrane Technologies | mdi Membrane Technologies |
| Lot Number | SN7353G | SV2462L |
| Size in µm | 0.45 µm | 0.45 µm |

Table S15: Assay Results for Filter Validation

| Filter details | % Assay | Absolute difference |
| --- | --- | --- |
| Centrifuged | 99.50 | Not Applicable |
| Nylon filter | 99.72 | 0.22 |
| PVDF filter | 99.62 | 0.12 |