Supplementary Table of Contents

Methods	2

Blood collection	2

Microbiology	2

Reporting of CRP concentrations outside of the limit of quantitation	2

Supplementary tables.	3

Supplementary table 1.	3

References	4

STARD reporting checklist for diagnostic test accuracy study.	5

# Methods

## Blood collection

Prior to any TB microbiology, participants gave fingerprick blood collected using a micropipette (LumiraDx). Participants also had ≥1 9 ml Serum Vacutainer tube (BD, Franklin Lakes, USA) filled with venous blood. Serum was isolated per manufacturer instructions (1) and stored at -80°C.

## Microbiology

Contacts with presumptive TB (≥1 symptom if PLWH; ≥2 if HIV-negative; possible symptoms include cough of any duration, haemoptysis, weight loss, fever, or night sweats) or able to produce ≥ non-salivary sputum provided two sputum samples (≥1 mL each). If necessary, sputum induction was performed using nebulised 3-10% NaCl. One sputum was used for MGIT960 liquid culture (BD, Franklin Lakes, USA) and the other Xpert MTB/RIF Ultra (Ultra) (2). Culture growth was confirmed as Mycobacterium tuberculosis complex using Ultra or MTBDRplus (Hain-Bruker Lifesciences, Nehren, Germany).

## Reporting of CRP concentrations outside of the limit of quantitation

LumiraDx has a reportable range of 5-250 mg/L. For continuous comparisons, values reported by the instrument as “<5 mg/L” [below the limit of detection (LOD) and limit of quantitation (LOQ] were arbitrarily assigned a midpoint value of 2.5 mg/L. Values reported “>250 mg/L” were assigned 251 mg/L. The reportable range for CRPHS is 0.2-1440 mg/L. Values reported “<0.2 mg/L” (LOD) were assigned 0.1 mg/L. No values between 0.2 mg/L (LOD) and 0.3 mg/L (LOQ) nor >1440 mg/L occurred.

# Supplementary tables.

## Supplementary table 1.

Results of Passing-Bablok regression and Bland-Altman plots for the different testing modalities. Fingerprick versus serum LumiraDx show strong agreement with minimal bias and similar slopes. Both exhibit comparable performance versus serum CRPHS, with no evidence of nonlinearity in any comparison

|  | Passing-Bablok (PB) regression |  |  |  | Bland-Altman (BA) plot |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Slope (95% CI) | Intercept (95% CI) | Lin’s CCC | CUSUM test | Mean difference (95% CI) | Lower LoA (95% CI) | Upper LoA (95% CI) |
| Fingerprick LumiraDx and serum LumiraDx | -0.27 (-0.82, 0.20) | +1.1 (1.0, 1.1) | 0.85 | P=0.15 | +1 (-1, 3) | -25 (-29, -22) | 27 (24, 31) |
| Serum CRPHS and serum LumiraDx | 1.1 (-1.1, 2.3) | +1.4 (1.3, 1.6) | 0.79 | P>0.20 | +11 (9, 14) | -20 (-25, -16) | 43 (39, 48) |
| Serum CRPHS and fingerprick LumiraDx | 1.3 (-0.6, 2.5) | +1.4 (1.3, 1.5) | 0.75 | P>0.20 | +10 (8, 13) | -27 (-32, -22) | 48 (43, 53) |

Abbreviations: CRPHS, cobas C‑Reactive Protein (Latex) High Sensitive; CCC, concordance correlation coefficient; CUSUM, cumulative sum control chart; LoA, Limit of agreement.

# References

1.	BD. BD Vacutainer serum tubes  [cited 2025 August 13]. Available from: https://www.bd.com/en-us/products-and-solutions/products/product-page.367815.

2.	Cepheid. Xpert MTB RIF Ultra Brochure CE IVD 3098 English  [Available from: https://infomine.cepheid.com/sites/default/files/2024-01/302-5776%20Rev.%20D%20IFU%20MTB.RIF%20Ultra%20ENGLISH%20MII.pdf.

# STARD reporting checklist for diagnostic test accuracy study.

|  |  | Reporting Item | Page Number |
| --- | --- | --- | --- |
| Title or abstract |  |  |  |
| None | #1 | Identification as a study of diagnostic accuracy using at least one measure of accuracy (such as sensitivity, specificity, predictive values, or AUC) | 1 |
| Abstract |  |  |  |
| None | #2 | Structured summary of study design, methods, results, and conclusions (for specific guidance, see STARD for Abstracts https://www.equator-network.org/reporting-guidelines/stard-abstracts/) | 3 |
| Introduction |  |  |  |
| None | #3 | Scientific and clinical background, including the intended use and clinical role of the index test | 4 |
| None | #4 | Study objectives and hypotheses | 4 |
| Methods |  |  |  |
| Study design | #5 | Whether data collection was planned before the index test and reference standard were performed (prospective study) or after (retrospective study) | 5 |
| Participants | #6 | Eligibility criteria | 5 |
| Participants | #7 | On what basis potentially eligible participants were identified (such as symptoms, results from previous tests, inclusion in registry) | 5 |
| Participants | #8 | Where and when potentially eligible participants were identified (setting, location and dates) | 5 |
| Participants | #9 | Whether participants formed a consecutive, random or convenience series | 5 |
| Test methods | #10 | Index and reference tests in sufficient detail to allow replication | 5-6 |
| Test methods | #11 | Rationale for choosing the reference standard (if alternatives exist) | 6 |
| Test methods | #12 | Definition of and rationale for test positivity cut-offs or result categories of the index and reference tests, distinguishing pre-specified from exploratory | 6 |
| Test methods | #13 | Whether clinical information and reference standard results were available to the performers / readers of the index test; Whether clinical information and index test results were available to the assessors of the reference standard | 5 |
| Analysis | #14 | Methods for estimating or comparing measures of diagnostic accuracy | 6 |
| Analysis | #15 | How indeterminate index test or reference standard results were handled | 5 |
| Analysis | #16 | How missing data on the index test and reference standard were handled | 5-6 |
| Analysis | #17 | Any analyses of variability in diagnostic accuracy, distinguishing pre-specified from exploratory | 6 |
| Analysis | #18 | Intended sample size and how it was determined | N/A |
| Results |  |  |  |
| Participants | #19 | Flow of participants, using a diagram | N/A |
| Participants | #20 | Baseline demographic and clinical characteristics of participants | 7 |
| Participants | #21 | Distribution of severity of disease in those with the target condition, and distribution of alternative diagnoses in those without the target condition | 7 |
| Participants | #22 | Time interval and any clinical interventions between index test and reference standard | 6 |
| Test results | #23 | Cross tabulation of the index test results (or their distribution) by the results of the reference standard | 7 |
| Test results | #24 | Estimates of diagnostic accuracy and their precision (such as 95% confidence intervals) | 7 |
| Test results | #25 | Any adverse events from performing the index test or the reference standard | N/A |
| Discussion |  |  |  |
| None | #26 | Study limitations, including sources of potential bias, statistical uncertainty, and generalisability | 10 |
| None | #27 | Implications for practice, including the intended use and clinical role of the index test | 10 |
| Other information |  |  |  |
| None | #28 | Registration number and name of registry | 5 |
| None | #29 | Where the full study protocol can be accessed | 5 |
| None | #30 | Sources of funding and other support; role of funders | 11 |

The STARD checklist is distributed under the terms of the Creative Commons Attribution License CC-BY. This checklist was completed on 04 November 2025 using https://www.goodreports.org/, a tool made by the EQUATOR Network in collaboration with Penelope.ai