Additional file 1: Supplementary Methods

Supplementary Methods 1. Step-by-step Algorithm for HB Computation (Data sources, SpO₂ pre-processing, event-based and second-by-second HB, implementation, and QC)

1. Step‑by‑step Algorithm for HB Computation

1.1 Data sources and definitions

Signals: (a) Peripheral oxygen saturation (SpO₂) recorded overnight; (b) Respiratory event (RE) annotations labeled as apnoea or hypopnoea per AASM rules; (c) Optional sleep staging to define total sleep time (TST) and exclude wake epochs.

Sampling: Raw device sampling ≥8–10 Hz before vendor filtering; analysis resampled to 1 Hz.

Units: HB is expressed as %·min/h (percentage‑minutes of desaturation per hour of effective sleep). Areas are first computed in %·s and converted to %·min/h.

1.2 Time base and t = 0 (“point zero”)

Convert device timestamps to an integer second counter.

Define t = 0 as the earliest of: (i) first scored sleep epoch, or (ii) explicit device marker (if present). If staging is unavailable, use the recording start after basic Q and approximate TST as the valid analysis window duration after artefact exclusion.

Express all times relative to t = 0.

1.3 Analysis window and TST

Sleep period: [t = 0, last scored sleep epoch].

Validity mask: seconds retained after artefact handling (Valid(t) = 1; see §2.5).

TST (hours): TST = (Σ Valid(t)) / 3600.

2. Pre‑processing of SpO₂

Resampling: resample to 1 Hz (nearest or linear interpolation after QC).

Physiological range: values <50% or >100% are set missing.

Artefact repair: gaps ≤5 s linearly interpolated; longer gaps remain missing and are excluded (Valid(t) = 0).

Optional smoothing: 3–5 s moving median (default: median, 3 s) applied to reduce single‑sample spikes without blunting nadirs.

Validity flags: Valid(t) = 1 iff SpO₂ available post‑QC; otherwise 0.

Edge cases: clock roll‑over across midnight; duplicate timestamps; monotone missing segments ≥30 s near device detachments (left missing).

3. Event‑based HB (HB_event — primary index)

3.1 Event eligibility & windows

Candidate REs are those with AASM‑concordant annotations (apnoea or hypopnoea).

Baseline window: [−30 s, −10 s] before the scored event onset ts_i.

Recovery window: extend forward from ts_i up to ts_i + 120 s or until recovery is detected (§3.3).

3.2 Baseline saturation

S_base,i = median{ SpO₂(t) : t ∈ [ts_i − 30, ts_i − 10], Valid(t) = 1 }.

If <50% of baseline window is valid, mark event ineligible.

3.3 Recovery detection

Define recovery when SpO₂(t) returns to within 1% of S_base,i and remains there for ≥2 consecutive seconds (hysteresis to avoid chatter). If no recovery by +120 s, truncate at +120 s.

3.4 Instantaneous deficit and minimum depth

Per‑second deficit: D_i(t) = max(0, S_base,i − SpO₂(t)) for t ∈ [ts_i, t_rec,i].

Minimum depth rule: discard event if the peak desaturation < 3% below baseline (i.e., max_t D_i(t) < 3).

3.5 Event area and aggregation

Event area in %·s: A_i = Σ_{t = ts_i}^{t_rec,i} D_i(t).

HB_event (%·min/h): HB_event = ( Σ_i A_i / 60 ) / TST_hours.

3.6 Quality control (QC) outputs

Counts: scored REs, retained events, excluded events (reasons: insufficient baseline, <3% depth, missing data >50% in window, truncated recovery).

Per‑night summaries: median baseline SpO₂, median nadir, median event duration.

Note: Primary and secondary analyses use HB_event exclusively. The continuous algorithm (§4) is documented for reproducibility but was not applied to the analytic dataset.

4. Second‑by‑second HB (HB_sec — documented, not used in analyses)

4.1 Dynamic baseline

B(t) = 95th percentile of SpO₂ in a rolling 5‑min window centered at t (use asymmetric window near edges).

4.2 Continuous deficit

C(t) = max(0, B(t) − SpO₂(t)).

4.3 Optional threshold

Enforce C(t) ≥ 3% (optional) to align with desaturation criteria; default off for sensitivity analyses.

4.4 Integration and scaling

HB_sec (%·min/h) = ( Σ_{t} C(t) / 60 ) / TST_hours, restricted to Valid(t) = 1 during sleep.

4.5 Sensitivity analyses

Alternative baselines: 90th–98th percentiles; rolling windows 3–10 min.

Alternative thresholds: 2–4%.

5. Implementation details (Stata to Python)

5.1 Development in Stata: Event loops with recovery detection and area summation using rangestat/by‑group operations; integer‑second time base; robust handling of midnight roll‑over.

5.2 Translation to Python: Vectorized Pandas/Numpy implementation with boolean masks for Valid(t), rolling quantiles (bottleneck/fallback); unit tests to match Stata outputs (absolute/relative tolerance ≤1e‑6 on areas and HB indices).

5.3 Reproducibility artefacts

Versioning: semantic versions for code and JSON config.

Config (JSON): sampling Hz, baseline/recovery windows, smoothing kernel, min depth, max recovery horizon, thresholds, QC toggles.

QC logs (CSV): per‑night counters, event‑level reasons for exclusion, timing truncations, and missingness rates.

6. Outputs

Primary indices: HB_event (%·min/h); HB_sec (%·min/h, if computed).

Auxiliary: A_total (%·s), C_total (%·s, if continuous), TST (h).

Counts: n_scored, n_retained, n_excluded (by reason).

Stratified (optional): position‑specific HB (supine, right, left, prone) using synchronized body‑position channels.

Per‑event table: start/end times, baseline, nadir, max depth, duration, area, position/stage label.

7. Interpretation

Higher HB indicates greater integrated nocturnal hypoxaemic load.

Event‑based HB captures desaturations tied to scored REs; more specific to OSA pathophysiology under AASM rules.

Second‑by‑second HB captures the full hypoxaemic profile (may include unscored fluctuations, e.g., non‑respiratory dips), useful for sensitivity analyses.

8. Parameter defaults and justifications

| Parameter | Default | Rationale |
| --- | --- | --- |
| Resampling | 1 Hz | Align with second‑by‑second event integration; avoids aliasing of nadirs |
| Baseline window | −30 to −10 s | Stable pre‑event plateau while avoiding immediate pre‑onset slope |
| Recovery horizon | 120 s | Captures typical re‑oxygenation; prevents indefinite tails |
| Recovery criterion | within 1% of baseline for ≥2 s | Hysteresis to avoid chatter |
| Min depth | 3% | Aligns with AASM desaturation criterion |
| Interp. of gaps | ≤5 s | Small gaps unlikely to bias areas; longer gaps excluded |
| Smoothing | 3‑s median | Reduces single‑sample spikes while preserving nadirs |

9. Mathematical definitions (explicit)

- Event area: (A_i = \sum_{t = t_{s,i}}^{t_{rec,i}} \max{0, S_{base,i} - S(t)}) (units: %·s)

- HB_event: (HB_{event} = \frac{\sum_i A_i / 60}{TST; (h)}) (units: %·min/h)

- Continuous baseline: (B(t) = Q_{0.95}\big({S(u) : u \in [t-150, t+150]}\big))

- Continuous burden: (HB_{sec} = \frac{\sum_t \max{0, B(t) - S(t)}/60}{TST; (h)})

10. Pseudocode (event‑based, simplified)

# Inputs: SpO2[sec], RE list with (ts, te), Valid[sec]

for each event i:

if valid_fraction(SpO2 in [ts-30, ts-10]) < 0.5: exclude; continue

S_base = median(SpO2[ts-30:ts-10])

t_end = min(ts + 120, detect_recovery(SpO2, S_base))

D = clip(S_base - SpO2[ts:t_end], lower=0)

if max(D) < 3: exclude; continue

A_i = sum(D)

HB_event = (sum(A_i)/60) / TST_hours

11. Quality‑control thresholds and reports

Event exclusion reasons logged at event‑level.

Night‑level flags: (i) missingness >20% during sleep; (ii) <4 retained events; (iii) TST < 240 min (study exclusion, per protocol).

Position channels audited: % of TST with valid position; if <10% in non‑supine, mark positional HB as unreliable.

12. File I/O and artefacts

Inputs: vendor‑exported CSV/XLSX with second‑by‑second data; optional event and position channels; optional stage channel.

Outputs: (i) per‑event table; (ii) per‑night summary; (iii) optional per‑position HB summary; (iv) QC log.

Metadata: code version, config JSON snapshot, device firmware/software versions, oximetry averaging time.

13. Reproducibility & versioning

Analyses run with fixed configuration snapshots; all thresholds and windows declared above.

Cross‑validated Stata→Python parity using unit tests on canonical samples; absolute/relative errors ≤1e‑6 on areas and ≤1e‑4 on HB indices.

14. Notes on generalizability and altitude

The algorithm is device‑agnostic if inputs provide second‑by‑second SpO₂ and synchronized annotations; cut‑off values may shift with different oximetry dynamics or averaging.

At moderate altitude (≈1,500–2,600 m a.s.l.), lower baseline SpO₂ may inflate absolute HB; report local reference percentiles alongside thresholds.

15. Worked example (unit consistency)

Example event: S_base = 95%, nadir 88%, duration 19 s.

Triangle approximation area ≈ ½ × 7% × 19 s = 66.5 %·s (if approximated triangular). With full second‑wise integration, area equals Σ per‑second deficits.

Single‑night aggregation: if ΣA_i = 35,000 %·s and TST = 6.0 h → HB_event = (35,000/60)/6 = 97.2 %·min/h.

16. Declarations for this Supplement

Primary index used: HB_event.

Continuous index: documented, not applied to this cohort’s results.

Manual review: index events edited per COMPASS/AASM rules prior to HB computation; REs with spurious 3% dips removed as a strict prerequisite

Supplementary Note — Corrected ODI3% Events

.For the corrected computation of ODI3% events, each additional desaturation was mapped to a fixed-duration element to align with HB units. Specifically, we assumed a constant duration of 10 s for each event and weighted it by the desaturation depth (in % below baseline). The nightly increment was then expressed in %·min/h as:

where depthj\mathrm{depth}_jdepthj​ is the event’s SpO₂ drop in percentage points relative to its baseline and TST is the total sleep time in hours.