Table 1 Repeatability and within-lab precision for latex reagent

|  | Samples |
| --- | --- |
| Measurement frequency | 80 |
| Mean (ug/g) | 49.60 |
| Maximum value (X max) (ug/g) | 54.00 |
| Minimum value (X min) (ug/g) | 45.00 |
| Standard deviation (SD) | 2.300 |
| S(R) (ug/g) | 2.455 |
| CV(R) | 4.9% |
| S(WL) (ug/g) | 2.317 |
| CV(WL) | 4.7% |

Table 2 Distribution of interference rates of four substances

The interference testing protocol involved measuring low-concentration sample(n=2) and high-concentration samples(n=2) at each level of the interfering substance. The average value from these measurements was used for subsequent analysis.

Table 3 Expected Bias at Medical Decision Levels