| Copies / reaction | Colorimetric RT-LAMP Mean TTP (min) ± SEM (N) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Orf1a |  | N+E |  | Orf1a+N+E |  |
| 10,000 | 8.2 ± 0.16 (8/8) |  | 13.2 ± 0.48 (8/8) |  | 10.0 ± 0.28 (8/8) |  |
| 1,000 | 10.0 ± 0.22 (8/8) |  | 16.0 ± 0.78 (8/8) |  | 12.0 ± 0.35 (8/8) |  |
| 500 | 10.7 ± 0.35 (8/8) |  | 15.8 ± 0.78 (8/8) |  | 12.7 ± 0.28 (8/8) |  |
| 100 | 12.4 ± 0.64 (3/8) |  | 19.4 ± 1.23 (7/8) |  | 16.2 ± 1.72 (5/8) |  |
| 50 | 13.4 ± 0.94 (4/8) |  | 24.5 ± 1.42 (4/8) |  | 20.2 ± 2.15 (5/8) |  |
| 10 | na |  | na |  | 22.5 ± 7.75 (2/8) |  |
| NTC | na |  | na |  | na |  |

Supplementary Table 1. Limit of detection of colorimetric RT-LAMP reactions targeting the Orf1a, dual N+E gene and multiplex Orf1a+N+E of SARS-CoV-2 using a serially diluted Twist Bioscience synthetic positive control RNA (control 2 GenBank ID MN908947.3, GISAID Wuhan-Hu-1). Synthetic RNA was serially diluted from stock to result in 10,000, 1,000, 500, 100, 50 and 10 copies of viral sequence per 25 µl reaction. Water no template control (NTC) were included in each reaction. Reactions were performed at 65°C for 40 minutes on the Applied Biosystems StepOnePlus platform. Representative colorimetric reactions for each primer set are shown whereby yellow indicates positive amplification and pink indicates no amplification (‘na’). Data summary is an average of 3 independent experiments, performed in duplicate and presented as mean TTP ± S.E.M. Numeration in parentheses indicates the number of repeat reactions achieving the amplification threshold required to report a TTP.

| RT-LAMP target: | Dse (95% CI) | Dse (95% CI) |
| --- | --- | --- |
|  | DSp (95% CI) Ct < 30 (N32) | DSp (95% CI) Ct < 40 (N41) |
| SARS-CoV-2 Orf1a | 100% (80.49% to 100.00%) | 85.7% (63.7% to 97.0%) |
|  | 100 % (91.19% to 100.00%) | 100.00% (91.19% to 100.00%) |
| SARS-CoV-2 Dual N+E | 100% (80.49% to 100.00%) | 85.7% (63.7% to 97.0%) |
|  | 100 % (91.19% to 100.00%) | 100.00% (91.19% to 100.00%) |
| SARS-CoV-2 Orf1a+N+E | 100% (89.11% to 100 %) | 90.24% (76.87% to 97.28%) |
|  | 100% (91.19% to 100.00%) | 100.00% (91.19% to 100.00%) |

Supplementary Table 2. Diagnostic sensitivity (Dse) and specificity (DSp) of a single (Orf1a) and dual (N+E) SARS-CoV-2 colorimetric end-point RT-LAMP assay amplifying residual RNA extracted from clinical patient samples with comparator RT-PCR Ct values. Reactions were performed at 65°C for 30 minutes on a StepOne PCR platform against 21 UHL positive samples & 40 UHL negative samples for the Orf1a and N+E RT-LAMP assays and 41 UHL positive samples & 40 UHL negative samples for the Orf1a+N+E RT-LAMP assay. Diagnostic test evaluation calculations performed using the Medcalc Scientific Software. Sensitivity and specificity with 95% CI are shown for samples with a corresponding RT-PCR Ct < 30 and Ct < 40.

| ORF1a HMSe F3 | CGGTGGACAAATTGTCAC |
| --- | --- |
| ORF1a HMSe B3 | CTTCTCTGGATTTAACACACTT |
| ORF1a HMSe FIP | TCAGCACACAAAGCCAAAAATTTATTTTTCTGTGCAAAGGAAATTAAGGAG |
| ORF1a HMSe BIP | TATTGGTGGAGCTAAACTTAAAGCCTTTTCTGTACAATCCCTTTGAGTG |
| ORF1a HMSe LF | TTACAAGCTTAAAGAATGTCTGAACACT |
| ORF1a HMSe LB | TTGAATTTAGGTGAAACATTTGTCACG |
| E-F3 | TGAGTACGAACTTATGTACTCAT |
| E-B3 | TTCAGATTTTTAACACGAGAGT |
| E-FIP | ACCACGAAAGCAAGAAAAAGAAGTTCGTTTCGGAAGAGACAG |
| E-BIP | TTGCTAGTTACACTAGCCATCCTTAGGTTTTACAAGACTCACGT |
| E-LB | GCGCTTCGATTGTGTGCGT |
| E-LF | CGCTATTAACTATTAACG |
| N-F3 | ACCAGGAACTAATCAGACAAG |
| N-B3 | GACTTGATCTTTGAAATTTGGATCT |
| N-FIP | TTCCGAAGAACGCTGAAGCGGAACTGATTACAAACATTGGCC |
| N-BIP | CGCATTGGCATGGAAGTCACAATTTGATGGCACCTGTGTA |
| N-LB | GGGGGCAAATTGTGCAATTTG |
| N-LF | CTTCGGGAACGTGGTTGACC |
| ACTB F3 | AGTACCCCATCGAGCACG |
| ACTB B3 | AGCCTGGATAGCAACGTACA |
| ACTB FIP | GAGCCACACGCAGCTCATTGTATCACCAACTGGGACGACA |
| ACTB BIP | CTGAACCCCAAGGCCAACCGGCTGGGGTGTTGAAGGTC |
| ACTB LF | TGTGGTGCCAGATTTTCTCCA |
| ACTB LB | CGAGAAGATGACCCAGATCATGT |

Supplementary Table 3. RT-LAMP primer sequences targeting the large open reading frame.  ORF1a-HMSe designed by Rabe and Cepko 2020, and the nucleocapsid, envelope and human beta-actin gene designed by NEB (Tanner at al., 2020 & Zhang 2020).