## Supplementary Information

### Theoretical Background

#### Signal Description

Interferometric NIRS records, at detector position r, optical frequency ν=c/λ (c is the speed of light in vacuum and λ is the wavelength), and acquisition time td (sweep index), the spectral interferogram:

S(r,ν,td)=SDC(r,ν,td)+2ReWrsr,ν,td, 	(S1)

where SDC is the DC offset, the sum of the sample- and reference-arm intensities, and Wrs is the cross-spectral density function between the optical fields from the two arms. Its real part carries the spectral fringe information. The symbol r marks the detector position. It is typically omitted in single-channel iNIRS (one detector) and becomes the pixel coordinate in parallel iNIRS.

After subtracting SDC and resampling to uniform optical frequency (k-linearization), inverse Fourier transformation with respect to ν yields the mutual coherence function:

Γrs(r, τs,td)=Ur*r,ts,tdUsr,ts+τs,tdts, 	(S2)

where Ur and Us are the optical fields in the reference and sample arms, τs is the time-of-flight (a conjugate variable to ν), and ∙ts denotes averaging over the fast sampling time within the sweep ts.

Using Γrs(r, τs,td) we form the two-dimensional autocorrelation:

G1(iNIRS)(r,τs,τd)=Γrs*r, τs,tdΓrsr, τs,td+τdtd. 	(S3)

with τd the lag between successive sweeps (sampling step δτd=1/fl, set by the sweep rate fl). The measured correlation is the true TOF-resolved autocorrelation G1(τs,τd)=Us*τs,tdUsτs,td+τdtd convolved along τs with the instrument response function (IRF):

G1(iNIRS)r,τs,τd=G1r,τs,τd*IRFτs, 	(S4)

where the IRF is governed by the by the laser tuning bandwidth Δλ. The wider the bandwidth the narrower the IRF, and thus finer the TOF resolution.

Setting τd=0 gives the temporal point spread function (TPSF):

IsiNIRS(r,τs)=G1(iNIRS)r,τs,τd=0, 	(S5)

which equals the true distribution of times of flight (DTOF) convolved with the same IRF.

For parallel iNIRS (πNIRS), we spatially average TOF-resolved quantities:

G1(πNIRS)τs,τd=G1iNIRSr,τs,τdr, 	(S6)

IsπNIRSτs=IsiNIRSr,τsr. 	(S7)

Integrating out TOF yields the continuous-wave limit (CW-πNIRS):

G1(CW-πNIRS)τd=0∞dτsG1(πNIRS)τs,τd, 	(S8)

IsCW-πNIRS=0∞dτsG1(CW-πNIRS)τd=0. 	(S9)

Thus CW-πNIRS captures the same integrated intensity and field-autocorrelation as parallel iNIRS, but without TOF resolution. For plotting and parameter estimation, we normalize to zero lag:

g1(CW-πNIRS)τd=G1CW-πNIRSτdG1CW-πNIRSτd=0, 	(S10)

g1(πNIRS)τs, τd=G1πNIRSτs,τdG1πNIRSτs,τd=0. 	(S11)

#### Temporal IRF, TOF resolution and the sampling-limited TOF range

In iNIRS (single‑channel or parallel), the instrument response function (IRF) is set by the laser tuning bandwidth Δλ. For a Gaussian spectrum and a Mach–Zehnder detection geometry, the full width at half maximum (FWHM) of the temporal IRF is:

δτs=22ln⁡(2)πλc2cΔλ≈0.62λc2cΔλ, 	(S12)

where λc is the centre wavelength. Thus, increasing Δλ sharpens the TOF resolution δτs in direct proportion to Δλ-1. So, a wider bandwidth yields finer resolution.

In a swept-source acquisition, each sweep provides N samples:

N= fs/fl, 	(S13)

where fs is the sampling rate and fl is the laser sweep (tuning) rate. The corresponding spectral step is δλ≈Δλ/N, which sets the unaliased TOF window (maximum measurable TOF):

τs,max(sampling)=λc2cδλ=fsflλc2cΔλ. 	(S14)

Finally, the autocorrelation-lag sampling is fixed by the sweep rate,

δτd = 1fl. 	(S15)

These relations expose two trade-offs:

- Bandwidth trade-off. Increasing Δλ improves δτs (finer depth resolution) but reduces τs,max(sampling).

- Speed trade-off. A faster sweep (larger fl) improves the autocorrelation sampling interval δtd = 1/fl needed for flow measurements but reduces τs,max(sampling) unless fs is increased proportionally. Raising fs, however, demands wider-bandwidth electronics and faster data transfer, adding cost and complexity.

Therefore, the iNIRS operating point must balance (i) fine TOF resolution for optical-property fits, (ii) sufficient TOF range to reach cortex, and (iii) high temporal sampling for dynamic flow, and all within the constraints of laser tuning speed and electronic bandwidth.

#### Sampling-limited sweep-rate ceiling

To set a practical target TOF window for cortical sensitivity, we adopt the standard TD-NIRS framework under diffusion theory for a semi-infinite head. Representative analyses and measurements [16] show usable DTOF windows of 1–5 ns for adult source–detector separations (SDS). Because iNIRS typically operates at shorter SDS than TD-NIRS, enabled by its markedly narrower IRF (15–50 ps FWHM in iNIRS versus 200–500 ps in TD-NIRS), we take τs,target≈ 1-1.5 ns as a design point for cortex-level sensing.

Using Eq. (S14) for the sampling-limited unaliased TOF window, we impose τs,max(sampling)≥ τs,target and obtain the range-limited sweep ceiling:

fl<fl,max= fsλc2cΔλτs,target. 	(S16)

Expressing Δλ in pm, τs,target in ns, and fl,max in kHz, we get:

fl,max kHz=Kτs,target ns Δλ [pm],     K=fsλc2c×1018 [kHz⋅ns⋅pm]. 	(S17)

Equation (S17) makes explicit that fl,max scales linearly with the sampling rate and quadratically with centre wavelength and decreases as 1/Δλ. Thus, widening Δλ (to sharpen δτs) reduces the allowable sweep rate unless fs is increased proportionally. At fixed fs and τs,target, halving Δλ doubles fl,max but simultaneously doubles δτs (coarser TOF resolution), reflecting the core bandwidth–range trade-off.

In our parallel-iNIRS prototype the dominant hardware limit is the camera sampling rate fs(πNIRS) = 1.1 MHz. By contrast, single-channel iNIRS typically employs a fast digitizer with fs(iNIRS) = 100-125 MS/s but with orders of magnitude fewer detection channels. Using Eq. (S17), such a fast digitizer raises K by ~100 relative to the camera, so range-limited sweep ceilings of the order of hundreds of kHz to ~1 MHz become theoretically admissible for the same Δλ and τs,target. In practice, however, tuneable-laser mechanics and electronics rarely support such high fl while maintaining sweep linearity, and one must retain a minimum spectral sampling density Nmin (typically Nmin≥32-128 samples across Δλ), imposing

fl,max kHz=minKτs,target Δλ ,fsNmin 	(S18)

Thus, even with a fast digitizer, N-density often limits fl,max to the tens–hundreds of kHz regime.

Figure S1 plots the sampling-limited sweep-rate ceiling fl,max as a function of the tuning bandwidth Δλ [Eq. (S16)] for λc=780 and 850 nm with fs=1.1 MHz. The black curves follow fl,max for τs,target=1 ns (solid) and 2 ns (dashed). Horizontal lines show the sampling-density caps fs/Nmin for Nmin∈32,64,128. Feasible operation lies below both the black curve and the relevant Nmin line. Specifically, the active limit is therefore the minimum of the range curve and the horizontal fs/Nmin line. The crossover bandwidth at which the cap begins to dominate is

Δλth(τs,target)=Kτs,targetfs/Nmin. 	(S19)

For τs,target=2 ns and Nmin=64, we obtain Δλ≈65 pm (780 nm) and ≈77 pm (850 nm). For τs,target=1 ns these thresholds double to Δλ≈130 pm and ≈154 pm, respectively. Thus, at narrower Δλ the cap fs/Nmin limits fl (curve above the line), whereas at wider Δλ the range relation limits fl (curve below the line). Increasing λc or fs raises K and shifts Δλth to larger bandwidths. Increasing Nmin (more conservative sampling) lowers the cap and shifts Δλth left, making the cap active over a wider Δλ range.

Fig. S1. Sampling-limited sweep-rate ceiling for camera-based πNIRS. Sweep-rate fl versus tuning bandwidth Δλ for (left) λc=780 and (right) 850 nm, assuming the detector sampling fs=1.1 MHz. Black curves show the range constraint (K/τs,target Δλ, Eq. (S17)) with τs,target=1 ns (solid) and 2 ns (dashed). Colored horizontal lines indicate sampling-density caps fs/Nmin for Nmin=32 (blue), 64 (red), and 128 (green). The feasible region lies below both the corresponding black curve and the cap. For example (right): at Δλ≈70 pm the 2 ns range curve gives fl,max≈18.9 kHz, close to the Nmin=64 cap (17.2 kHz, red line). Increasing wavelength raises the curves as fl,max∝λc2; widening Δλ improves TOF resolution but lowers the allowable sweep rate, illustrating the bandwidth–range trade-off.

#### Sweep-rate limitations imposed by decorrelation dynamics

Beyond the sampling-limited TOF window, the ability to resolve blood-flow dynamics is set by the decorrelation rate of scattered photons. According to Diffusing Wave Spectroscopy (DWS) theory, the field autocorrelation at a given photon time-of-flight, τs decays exponentially in lag τd:

g1τs,τd=exp-ξτsτd, 	(S20)

with a decay rate:

ξτs=2k2αDBμs'cτs/n, 	(S21)

where k=2πn/λc is the in-medium wavenumber, μs' the reduced scattering coefficient, n the refractive index, and αDB the blood-flow index. The proportionality ξτs∝τs indicates that longer-path photons decorrelate faster, so deeper tissue gates require faster temporal sampling to capture their dynamics.

The sweep rate fl sets the autocorrelation lag step δτd=1/fl. To adequately resolve the exponential decay, the sampling step δτd must be small compared to the decay time 1/ξ(τs). A practical criterion is

δτd≤  1m ξτs    ⇔   fl≥ mξ(τs), 	(S22)

where m≈ 3-5 ensures stable curve fitting. This condition defines a decorrelation-limited maximum measurable TOF,

τs,max(ξ)≈ flmn2k2αDBμs'c. 	(S23)

Substituting k=2πn/λc yields a compact form,

τs,max(ξ)≈ flmλc28π2cnαDBμs'×1012    [ns] 	(S24)

Equation (S24) provides several guidelines that are useful for experimental design. The maximum correlation time, τs,max(ξ), exhibits distinct scaling behavior with respect to both instrumental and tissue parameters. It increases linearly with the sampling frequency, such that a doubling of fl results in a twofold increase of τs,max(ξ). The dependence on the center wavelength λc follows a quadratic relationship: for example, an increase from 780 to 1060 nm raises τs,max(ξ) by a factor of 10607802≈ 1.85, consistent with the reduced sensitivity of longer wavelengths to rapid decorrelation dynamics. In contrast, the reduced scattering coefficient μs' contributes inversely; higher scattering shortens the accessible correlation time. For cortical tissue, where μs'≈0.9 mm-1 at 850 nm, an increase to μs'≈1.2 mm-1 reduces τs,max(ξ),  by approximately 25%. A similar inverse relationship holds for the flow index, αDB, where faster flow diminishes the ceiling: a twofold increase in blood flow index effectively halves the decorrelation-limited time-of-flight window.

It is convenient to define the decorrelation slope (rate per unit TOF) as

Sλc ≡ ξτs=8π2cnλc2αDBμs'×10-12    [kHz/ns] 	(S25)

Then, from Eq. (S22),

flτs≈mSλc τs 	(S26)

Inverting Eq. (S26) gives

τs,max(ξ)≈flmSλc. 	(S27)

This expression quantifies the largest TOF gate whose dynamics can be reliably captured at a given sweep rate, explicitly formalizing the trade-off between wavelength, scattering, flow, and sweep frequency.

The usable TOF range is now the smaller of the sampling-limited and decorrelation-limited ceilings:

τs≤minτs,max(sampling),τs,max(ξ) . 	(S28)

Although a sufficiently high camera sampling rate may, in principle, support a long τs,max(sampling), rapid decorrelation at large τs can render such measurements noise-dominated unless the sweep rate is increased to satisfy fl≥ mξ(τs). Increasing fl improves sensitivity to faster dynamics but simultaneously reduces τs,max(sampling) unless the camera sampling frequency is proportionally raised. In practice, fl is bounded by laser tuning capabilities, sweep linearity, and the bandwidth and throughput limits of the detection electronics. Consequently, the optimal sweep rate must balance decorrelation resolvability and TOF coverage. Faster sweeps enhance the ability to measure rapid blood-flow changes (higher ξ(τs)) but reduce the maximum TOF range accessible within the same sampling budget. Selecting fl to navigate this trade-off is thus critical for achieving quantitative, depth-resolved cerebral blood-flow measurements.

Figure S2 summarizes the constraints imposed by decorrelation dynamics on the choice of sweep rate. The left panels show how the maximum measurable TOF window is jointly limited by sampling (black and gray dashed curves) and decorrelation (colored bands), with the usable range defined by the smaller of the two ceilings. At higher sweep rates, the decorrelation ceiling increases linearly with fl, while the sampling ceiling decreases inversely with the sweep range Δλ, leading to a crossover point that sets the effective operating limit. The right panels depict the complementary perspective. The minimum sweep rate required to robustly capture dynamics at a given TOF. Longer wavelengths reduce the required sweep rate owing to their quadratic scaling with λc, whereas larger blood-flow indices (αDB) demand proportionally faster sweeps due to more rapid decorrelation. Likewise, wider tuning ranges (Δλ) tighten the sampling ceiling and thus also necessitate higher sweep rates. Together, these results highlight the intrinsic trade-off between sensitivity to fast dynamics and coverage of longer photon path lengths, emphasizing that optimal sweep-rate selection must balance these competing constraints.

Fig. S2. Sweep rate limitations imposed by decorrelation dynamics. (Left) Maximum measurable TOF τs,max versus fl, showing the decorrelation-limited ceiling τs,max(ξ) [Eq. (S24)] and the sampling-limited ceiling τs,max(sampling)  [Eq. (S14)] for representative Δλ. Shaded regions indicate parameter ranges for different αDB. (Right) Minimum sweep rate required to robustly resolve a given τs, fl,min(τs) [Eq. (S26)]. Longer wavelengths reduce the required fl (∝λc2), whereas larger αDB demands faster sweeps.

### Design Considerations

#### Selecting operating parameters for parallel iNIRS

The dominant hardware limit in our parallel iNIRS prototype is the camera sampling. The camera sensor clocks fs = 1.1 MHz. So, every combination of sweep bandwidth (Δλ) and sweep rate (fl) must meet the required TOF resolution, TOF range, and autocorrelation lag- step without exceeding that 1.1 MHz ceiling.

Table S1 illustrates the trade-offs assuming λc=785 nm and fs=1.1 MHz. A wider bandwidth sharpens δτs but squeezes τs,max. Speed row (fl) shows that faster sweeps give finer δtd for flow, yet shrink τs,max. Any attempt to push fl higher or to regain TOF range would require faster, higher-bandwidth electronics.

|  | Unit | High- time-of-flight resolution | Balanced | Fast-flow sensing | Extended fast-flow |
| --- | --- | --- | --- | --- | --- |
| Tuning bandwidth, Δλ | pm | 100 | 50 | 25 | 25 |
| Sweep rate, fl | kHz | 10 | 20 | 50 | 100 |
| Time-of-flight resolution, δτs | ps | 13 | 25 | 51 | 51 |
| Sampling-limited TOF ceiling, τs,maxsampling | ns | 2.3 | 2.3 | 1.8 | 0.9 |
| Decorrelation-limited TOF ceiling, τs,maxsampling αDB=0.5×10-8cm2/s | ns | 0.12 | 0.25 | 0.62 | 1.24 |
| Decorrelation-limited TOF ceiling, τs,max(ξ) αDB=1.0×10-8cm2/s | ns | 0.06 | 0.12 | 0.31 | 0.62 |
| Autocorrelation step, δtd | μs | 100 | 50 | 20 | 10 |
| Comments |  | Resolution-limited, decorrelation dominates | Balanced trade-off | Fast flow captured but TOF coverage reduced | Electronics-limited, sampling ceiling dominates |

Table S1. Example operating points for parallel iNIRS at λc=785 nm and fs=1.1 MHz, showing the impact of sweep bandwidth and sweep rate on TOF resolution, TOF coverage, and flow sensitivity. The limiting ceiling in each configuration is shown in bold. At low sweep rates, decorrelation dynamics dominate (τs,max(ξ)), while at high sweep rates the sampling ceiling (τs,max(sampling)) becomes more restrictive.

#### Motivating the hybrid system

The parameter analysis above highlights the fundamental trade-offs in interferometric NIRS. Classical single-mode iNIRS offers sub-20 ps TOF resolution but suffers from photon starvation at centimeter-scale source–detector separations, forcing integration times >1 s and blurring fast hemodynamic responses. Conversely, parallel iNIRS (πNIRS) captures thousands of speckle modes simultaneously, dramatically improving photon throughput, yet the limited sampling rate of CMOS cameras (1.1 MHz in our setup) restricts the sweep rate to ~20 kHz. As shown in Figure S2 and Table S1, such rates fall short of resolving nanosecond-scale decorrelation dynamics, leaving default πNIRS apparently unable to resolve fast rapid cerebral blood-flow fluctuations.

To overcome these complementary bottlenecks, we engineered a hybrid interferometer architecture that shares a single swept source but distributes it across three coordinated detection paths:

- Single-mode iNIRS, digitized at 125 MS/s, providing precise wavenumber linearization, phase stability, and sub-20 ps TOF fidelity.

- Wide-field heterodyne camera arm, imaging ~2,000 speckle modes at 1.1 MHz, which boosts photon throughput and SNR through spatial averaging without requiring fiber bundles or costly detector arrays.

- Auxiliary interferometer arm, dedicated to real-time wavenumber calibration and source fluctuation monitoring, ensuring timing stability across sweeps.

By integrating these complementary detection strategies into one system, the hybrid design simultaneously achieves the high depth resolution of iNIRS, the photon throughput of πNIRS, and the timing stability necessary for fast sweeping. Joint operation enables practical sweep rates of 50–100 kHz while maintaining nanosecond TOF coverage, effectively balancing resolution, throughput, and temporal sampling. This purpose-built systems approach addresses the core hardware bottlenecks of interferometric NIRS and establishes a compact, fiber-sparse platform suited for bedside clinical monitoring of cerebral blood flow and oxygenation.

## Operating regimes

For in vivo measurements of the human brain, the expected time-of-flight (TOF) distribution spans approximately 1–3 ns. Therefore, we performed numerical simulations to establish the optimal system parameters. For the calculations, the sampling frequency of 125 MS/s for reference interferometer and iNIRS subunits, the refractive index of 1.46 (RI of SMF core), and the geometrical path difference of 27 cm were assumed. The laser tuning bandwidth (here – 50 pm) constrains the allowable laser sweep frequency fl to ~15–40 kHz (Fig. S3, top left). By selecting an operating frequency within this interval, for example 20 kHz, the TOF imaging window can be further tailored to the desired TOF imaging range by changing the sweep bandwidth, as illustrated in Fig. S3 (top right). However, at 20 kHz the πNIRS subunit cannot provide a fully time-resolved temporal point-spread function (TPSF) because of the limited sampling rate (Fig. S3, bottom left). Importantly, the laser tuning bandwidth Δλ not only determines the accessible TOF range, but also governs the decay rate ξ(τs), which characterizes the temporal loss of coherence in the detected light field. This decay rate encodes the dynamics of scattering particles within the tissue – such as red blood cell motion – that we aim to quantify.

Fig. S3. Operating regimes of the hybrid interferometric NIRS system. The available laser sweep frequencies for the human brain in vivo measurements (top, left). The available TOF imaging range for the sampling frequency of 20 kHz in the function of the sweep bandwidth for iNIRS (top, right) and πNIRS (bottom, left) subunit. Laser sweep frequency and decay rate in the function of the largest measurable TOF at the sweep bandwidth of 25 pm (bottom, right).

## Summary

The theoretical framework presented in this Supplement provides a unified description of interferometric NIRS that connects optical design parameters with measurable performance metrics. Starting from the field formulation of the spectral interferogram, we derived the mutual coherence function and the TOF-resolved field autocorrelation G1(τs,τd), which together define how depth information and temporal dynamics are encoded in interferometric signals. The analysis reveals that the practical performance of any iNIRS or πNIRS implementation is governed by two fundamental trade-offs: the bandwidth trade-off, linking spectral tuning range to TOF resolution and penetration depth, and the speed trade-off, linking sweep rate to the ability to sample decorrelation dynamics.

By expressing these limits through simple scaling laws and design ceilings, we provide an explicit roadmap for selecting operating points that balance TOF resolution, dynamic range, and photon throughput under real hardware constraints. The sampling- and decorrelation-limited windows derived here [Eqs. (S14)–(S28)] apply broadly to any swept-source interferometric technique, extending beyond iNIRS to full-field OCT, interferometric DWS, and emerging multi-speckle systems. Figures S1–S3 and Table S1 translate these relations into experimentally accessible parameters, enabling rapid estimation of feasible sweep frequencies, tuning ranges, and camera or digitiser requirements for a given application.

For the present study, this analysis justified the hybrid design: combining a single-mode iNIRS arm for high TOF fidelity with a parallel camera arm for photon efficiency and an auxiliary reference for stability. More generally, the framework can guide other researchers developing next-generation interferometric instruments, whether aimed at cerebral monitoring, tissue-flow mapping, or non-biological diffuse media. The equations and scaling relations derived here define the quantitative limits within which any swept-source interferometer can operate effectively, and thus serve as a practical reference for optimising performance in photon-limited, dynamic scattering environments.