## Sheet1
Author | Title | Intensity Normalisation | Pre-processing scheme utilised | Limitations of ML models | Size of external validation data set (K. pneumoniae) | Size of external validation data set KP | MALDI-TOF MS instrument used | Culture Medium Used For MALDI TOF MS spectra | Study Design | MALDI-TOF Technical Approach | Machine Learning methods | Target Reistance | Sample Characteristics | Performance Metrics | Clinical Advantages | Key Limitations | Number of Isolates | Antibiotic resistance prediction | Source of the isolates | Number of sources of isolates, i.e. different hospitals | Specimen Source Isolation | MALDI-TOF MS Instrument used | Machine Learning Preprocessing Steps | Performance metrics of models | Best model developed | External Validation Performed | Improved time to detection | Top AMR peaks m/z/ m/z biomarker
Zeng et al., 2023 | Machine learning model of imipenem-resistant Klebsiella pneumoniae based on MALDI-TOF-MS platform: An observational study | Not mentioned | - Standardization and normalization of m/z data - Setting starting m/z value to 1962DA and ending m/z value to 19998DA - Resampling to 15000 points - Gaussian filtering for data smoothing with σ = 1 - White Top Hat algorithm for baseline correction with window value of 5 - Equal-distance binning for mass spectral peak processing - Peak extraction using Python scipy package | - OPLS-DA analysis was not suitable for predicting drug sensitivity due to complex mass spectrometry information. - SVM algorithm had lower accuracy in the test set compared to LASSO. - Need for multicenter big data studies to improve overall accuracy. - Application of machine learning in clinical diagnosis is still in its early stages. | Not mentioned (no information on an external validation dataset is included in the paper) |  | Bruker | MacConkey agar | - Study type: Observational study - Setting and timeframe: Laboratory of microbiology department of Tianjin Haihe Hospital, 2019 January to 2020 December - Data collection approach: MALDI-TOF-MS and imipenem sensitivity data from clinical specimens - Comparison groups: 70 imipenem-sensitive and 70 resistant cases for training; 17 imipenem-sensitive and 17 resistant cases for testing - Validation strategy: Grid search and 3-fold Cross-validation; test set confusion matrix | - Equipment/platform used: Microflex LT spectrometer (Bruker Daltonik) - Sample preparation methods: Transfer to MALDI-TOF MS target plate with α-cyano-4-hydroxycynnamic acid-matrix solution, dried at room temperature - Spectral data preprocessing steps: Standardization, normalization, Gaussian filtering, White Top Hat algorithm for baseline correction, equal-distance binning - Mass range analyzed: 2 to 20 kDa - Number of technical replicates per isolate: Three - Quality control measures: Mass spectrum identification scores ≥2.0 considered high confidence - Spectral data processing software/methods: Python 3.8.8 with scipy package for peak extraction, MetaboAnalyst 5.0 for OPLS-DA | - Specific algorithms used: LASSO, LR, SVM, NN - Training/validation/test set split ratios and sizes: Training set - 70 sensitive, 70 resistant; Test set - 17 sensitive, 17 resistant - Feature selection or extraction methods: Standardization, normalization, Gaussian filtering, White Top Hat baseline correction, equal-distance binning - Hyperparameter optimization approach: Grid search - Cross-validation strategy: 3-fold Cross-validation - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: Python 3.8.8 with Scikit-learn package | - Specific antibiotics tested: imipenem - Resistance mechanisms: Not mentioned - Reference standard used for resistance determination: Minimum inhibitory concentration (MIC) method and VITEK-2 compact system - Resistance breakpoints or criteria applied: Sensitive ≤1 μg/mL, Resistant ≥4 μg/mL (CLSI guidelines) - Multi-drug resistance patterns: Not assessed | - Total sample size: 174 cases - Breakdown by resistance status: 87 resistant, 87 susceptible - Source of isolates: Clinical specimens from Tianjin Haihe Hospital - Isolation timeframe: January 2019 to December 2020 - Patient population: Individuals with K. pneumoniae infection - Inclusion/exclusion criteria: Non-repetitive samples with written informed consent - Distribution of resistant vs. susceptible isolates: 87 resistant, 87 susceptible | - R²Y and Q² of OPLS-DA: 0.546 and 0.0178 - AUC of training set: 0.9726 - AUC of test set: Not reported (0) - Comparative performance: LASSO had better accuracy than SVM in training set | - Time to result compared to conventional methods: Not explicitly mentioned. - Cost considerations or comparisons: The model is described as "inexpensive." - Ease of implementation in clinical labs: The model uses existing MALDI-TOF-MS technology, suggesting ease of implementation. - Integration with existing workflows: The model is designed to assist clinical antimicrobial application decision-making. - Real-time or rapid result capabilities: Not explicitly mentioned, but implied through integration of clinical expertise and big data analysis. - Potential impact on antimicrobial stewardship: The model aims to assist clinical antimicrobial application decision-making, potentially improving stewardship. | - Technical limitations: OPLS-DA analysis not suitable for prediction due to complexity of mass spectrometry information. - Performance issues: SVM algorithm had lower accuracy in test set. - Generalizability concerns: Need for multicenter big data studies to improve accuracy. - Need for larger validation studies: To drive clinical application and improve overall accuracy. - Implementation barriers: Early stage of exploration for AI diagnostic methods. - Areas requiring further research: Integration of clinical expertise and big data analysis. | 174 | The antibiotic resistance prediction in this study is focused on imipenem, a carbapenem antibacterial drug, in Klebsiella pneumoniae using machine learning models based on MALDI-TOF-MS data. | Clinical specimens from Tianjin Haihe Hospital | 1 (Tianjin Haihe Hospital) | - Broncho-alveolar lavage - Blood - Pleural fluid - Pus - Secretion - Sputum - Urine | Microflex LT spectrometer (Bruker Daltonik) | - Standardization and normalization of m/z data - Setting start m/z value to 1962DA and end m/z value to 19998DA - Resampling with 15000 points - Data smoothing using Gaussian filtering with σ = 1 - Baseline correction using White Top Hat algorithm with window value of 5 - Mass spectral peak m/z processed by equal-distance binning - Peak extraction using Python's scipy package | - R²Y and Q² of OPLS-DA: 0.546 and 0.0178 - AUC of best training set model: 0.9726 - AUC of best test set model: 0 - Accuracy verified by test set confusion matrix - LASSO algorithm had better accuracy and AUC compared to SVM | LASSO algorithm | Not mentioned (no information on external validation using an independent dataset) | Not mentioned (the paper does not provide specific information on improved time to detection) | Not mentioned (the paper does not explicitly mention a "Top AMR peaks m/z/ m/z biomarker")
Iskender et al., 2023 | Special Section on Innovative Diagnostic Tools in Infectious Diseases Rapid determination of colistin resistance in Klebsiella pneumoniae by MALDI-TOF peak based machine learning algorithm with MATLAB | Intensity normalization was done using the highest intensity, applied per spectrum rather than globally across the dataset. | 1. Sample preparation: Culture on MacConkey agar, McFarland suspensions, centrifugation. 2. Protein extraction: Ethanol, formic acid, acetonitrile treatment. 3. MALDI-TOF MS preparation: Dispensing on metallic plate, drying, adding HCCA matrix. 4. Spectral preprocessing: Removing high-frequency noise, reducing spectrum range to 2000-12,000 Da. 5. Normalization: Bringing spectra to a common low frequency, normalizing densities. 6. Baseline estimation: Subtracting high-frequency noise and signal peaks. 7. Feature selection: Genetic algorithm to select 30 discriminative features. | - Small training data size may prevent the model from capturing complexity. - Imbalance between colistin-resistant and colistin-sensitive isolates affects performance. - Small sample size for external validation. - Inability to evaluate prevalence of plasmid and chromosome-encoded resistance. - Low sensitivity rate of LDA model in application test set. - Initial high costs of molecular methods. | 60 | 60 | Bruker | MacConkey agar | - Study type: Retrospective - Setting and timeframe: Department of Medical Microbiology at Istanbul University-Cerrahpasa, 2015-2020 - Data collection approach: Collection of K. pneumoniae isolates from clinical samples - Comparison groups: Colistin-resistant and colistin-susceptible isolates - Validation strategy: 10-fold cross-validation during training; external validation with 60 isolates | - Equipment/platform used: microflex LT/SH model MALDI-TOF instrument (Bruker Daltonics, Germany) - Sample preparation methods: Culture on MacConkey agar, McFarland suspensions, centrifugation, suspension in ethanol, formic acid, and acetonitrile, vortexing, and centrifugation to remove cellular debris - Spectral data preprocessing steps: Signal resampling, baseline estimation, feature selection using genetic algorithm - Mass range analyzed: Initially 2000 to 20,000 Da, reduced to 2000 to 12,000 Da - Number of technical replicates per isolate: 240 laser shot measurements - Quality control measures: Bruker Biotyper automation control and software libraries - Spectral data processing software/methods: fleXcontrol 3.4 software, MATLAB 2020B | - Specific algorithms used: Linear Discriminant Analysis (LDA), Support Vector Machine (SVM), Ensemble - Training/validation/test set split ratios and sizes: Training set = 200 isolates, Test set = 60 isolates - Feature selection or extraction methods: Genetic algorithm to find 30 discriminative features - Hyperparameter optimization approach: Not mentioned - Cross-validation strategy: 10-fold cross-validation - Any ensemble methods or model combinations: Not mentioned beyond the use of the three algorithms - Software/programming languages used: MATLAB 2020B | - Specific antibiotics tested: Colistin - Resistance mechanisms: Not specified - Reference standard used for resistance determination: Broth microdilution method - Resistance breakpoints or criteria applied: MIC values for colistin-resistant isolates range from 4 to >128 mg/mL - Multi-drug resistance patterns: Not assessed | - Total sample size: 260 K. pneumoniae isolates - Breakdown by resistance status: 134 colistin-resistant, 126 colistin-susceptible - Source of isolates: Various clinical samples at Istanbul University-Cerrahpasa, Cerrahpasa Medical Faculty, Department of Medical Microbiology - Isolation timeframe: 2015 to 2020 - Clinical samples: Urine (45.8%), blood (21.9%), endotracheal aspirate (15.4%), other sources (16.9%) - Inclusion/exclusion criteria: Repetitive isolates excluded - Distribution of resistant vs. susceptible isolates: 134 resistant, 126 susceptible | - Training Phase: - LDA: Accuracy = 99.3% - SVM: Accuracy = 93.1% - Ensemble: Accuracy = 88.3% - Test Set: - LDA: Accuracy = 81.6%, Sensitivity = 66.7%, Specificity = 91.7%, Precision = 84.2% - AUC Values: - LDA: 0.99 - Ensemble: 0.94 - Performance Differences: - Discrepancy between training and test set performance due to small training data size and class imbalance. - Comparative Performance: - LDA shows the best performance in terms of accuracy and AUC. | - Time to result: Rapid compared to conventional methods like broth microdilution that take days. - Cost considerations: No additional cost when used for routine bacterial identification; potential for long-term cost savings. - Ease of implementation: Described as "easy" and "straightforward," suggesting ease of integration into clinical labs. - Integration with existing workflows: Can be used simultaneously with bacterial identification. - Real-time or rapid result capabilities: Emphasized as "rapid." - Potential impact on antimicrobial stewardship: Believed to help in early management of life-threatening infections. | - Sample size limitations: Small sample size for external validation due to time and technical constraints. - Technical limitations: Small training data size and class imbalance affecting model performance. - Generalizability concerns: Limited evaluation of plasmid and chromosome-encoded resistance. - Performance issues: Low sensitivity rate of the LDA model. - Need for larger validation studies: Planned increase in dataset size to mitigate class imbalance and improve performance. - Implementation barriers: Initial high costs of molecular methods. - Areas requiring further research: Evaluation of plasmid and chromosome-encoded resistance, addressing class imbalance. | 260 | Colistin resistance prediction using machine learning algorithms applied to MALDI-TOF data in Klebsiella pneumoniae isolates. | - Source: Clinical samples from Istanbul University-Cerrahpasa, Cerrahpasa Medical Faculty, Department of Medical Microbiology - Types: Urine (45.8%), Blood (21.9%), Endotracheal aspirate (15.4%), Other sources (16.9%) | 1 (Istanbul University-Cerrahpasa) | - Urine: 119 (45.8%) - Blood: 57 (21.9%) - Endotracheal aspirate: 40 (15.4%) - Other sources: 44 (16.9%) | microflex LT/ SH model MALDI-TOF instrument (Bruker Daltonics, Germany) | 1. Removal of high-frequency noises from the spectrum data. 2. Reduction of spectrum range from 2000 to 20,000 Da to 2000 to 12,000 Da. 3. Normalization of spectrum data to a common low frequency and density normalization. 4. Signal resampling to remove noise and redundant values. 5. Baseline estimation and subtraction to bring MS signals to a common baseline. 6. Feature selection using a genetic algorithm to select 30 discriminative features. | - LDA: Accuracy = 81.6%, Sensitivity = 66.7%, Specificity = 91.7%, Precision = 84.2% - SVM: Accuracy = 76.7%, Sensitivity = 62.5%, Specificity = 88.9%, Precision = 78.9% - Ensemble: Accuracy = 78.3%, Sensitivity = 58.3%, Specificity = 88.9%, Precision = 77.8% | LDA (Linear Discriminant Analysis) | - Total isolates: 60 - Colistin-susceptible: 24 - Colistin-resistant: 36 | The MALDI-TOF MS screening assay is described as "rapid and accurate," providing a faster detection method compared to the traditional broth microdilution method, which takes days to complete. The exact time improvement is not quantified in the paper. | Not mentioned (the paper does not specify any "Top AMR peaks m/z/ m/z biomarker")
Yu et al., 2023 | Prediction of methicillin-resistant Staphylococcus aureus and carbapenem-resistant Klebsiella pneumoniae from flagged blood cultures by combining rapid Sepsityper MALDI-TOF mass spectrometry with machine learning | Intensity normalization was performed using min-max scaling to keep signal intensity within the 0-1 range. The specific method (e.g., TIC, vector norm, base-peak) is not explicitly mentioned. Normalization was likely applied globally across the dataset, but this is not explicitly stated. | - Baseline correction using TopHat - Intensity normalization using square root transformation - Smoothing using Savitzky-Golay-filter - Conversion to m/z ratios - Resampling to a fixed sampling rate of 1 Da - Min-max scaling to keep signal intensity within 0-1 range - Division of spectra into 10-Da bins with maximum peak intensity as representative feature | The study identified a limitation of the ML models as the limited number of samples used in the study, which requires caution in interpreting the effectiveness of the methodology. Further research with larger sample sizes is recommended to validate and refine the approach. | Not mentioned (the paper does not specify the size of an external validation dataset) |  | Bruker Microflex LT/SH, Bruker Daltonics GmbH | BD BBL trypticase soy agar with 5% sheep blood, Levine eosin methylene blue agar, BD BBL CDC anaerobe with 5% sheep blood agar, and chocolate agar | - Study type: Retrospective - Setting and timeframe: China Medical University Hospital, October to December 2022 - Data collection approach: Flagged blood culture bottles, conventional MALDI-TOF MS, Rapid Sepsityper protocol - Comparison groups: Rapid Sepsityper protocol vs. conventional methods - Validation strategy: Five-fold cross-validation | - Equipment/platform used: Bruker Microflex LT/SH and Rapid Sepsityper Kit from Bruker Daltonics GmbH - Sample preparation methods: Transfer of 1.0-mL sample from positive blood culture bottle to 1.5-mL centrifuge tube - Spectral data preprocessing steps: Baseline correction, intensity normalization, conversion to m/z ratios, resampling to 1 Da, min-max scaling - Mass range analyzed: Not explicitly mentioned - Number of technical replicates per isolate: Ten 1.0-mL samples per isolate for data augmentation - Quality control measures: Quality check steps before data incorporation - Spectral data processing software/methods: Biotyper 3.0 software and LightGBM | - Specific algorithms used: LightGBM - Training/validation/test set split ratios and sizes: Not explicitly mentioned, but five-fold binary classifiers were used - Feature selection or extraction methods: Baseline correction, intensity normalization, conversion to m/z ratios, min-max scaling, division into 10-Da bins with maximum peak intensity as representative feature - Hyperparameter optimization approach: Not explicitly mentioned - Cross-validation strategy: Five-fold binary classifiers - Any ensemble methods or model combinations: Not explicitly mentioned - Software/programming languages used: Not explicitly mentioned | - Specific antibiotics tested: Methicillin for MRSA, Carbapenem for CRKP - Resistance mechanisms: Not specified - Reference standard: 2022 guidelines of the Clinical and Laboratory Standards Institute - Resistance breakpoints or criteria: Not specified beyond CLSI guidelines - Multi-drug resistance patterns: Not assessed | - Total sample size: 461 patients - Breakdown by resistance status for K. pneumoniae: 53 carbapenem-resistant, 79 carbapenem-susceptible - Source of isolates: Flagged blood culture bottles from China Medical University Hospital - Isolation timeframe: October 1 to December 31, 2022 - Patient population: Patients with monomicrobial bacteraemia - Inclusion/exclusion criteria: Not explicitly mentioned - Distribution of resistant vs. susceptible isolates: 53 resistant, 79 susceptible | - Accuracy for MRSA/MSSA: 0.875, 0.83 - AUC for MRSA/MSSA: 0.898, 0.85 - F1 score for MRSA/MSSA: 0.904, 0.87 - Accuracy for CRKP/CSKP: 0.766, 0.81 - AUC for CRKP/CSKP: 0.828, 0.82 - F1 score for CRKP/CSKP: 0.795, 0.86 | - Time to result: Within 1 hour, significantly faster than conventional methods. - Cost considerations: No additional costs, integrates into existing workflow without extra equipment or manual processes. - Ease of implementation: Integrates seamlessly into current workflow. - Integration with existing workflows: Yes, integrates well with existing clinical microbiology laboratory workflows. - Real-time or rapid result capabilities: Enables rapid identification within 1 hour. - Potential impact on antimicrobial stewardship: Enables earlier initiation of targeted antimicrobial therapy, improving patient outcomes and reducing mortality. | - Sample size limitations: The study had a limited number of samples, which may affect the generalizability of the findings. - Generalizability concerns: The effectiveness of the methodology should be interpreted with caution due to the small dataset. - Need for larger validation studies: Further research with larger sample sizes is needed to validate and refine the approach. | 461 | - Methicillin-resistant Staphylococcus aureus (MRSA) - Carbapenem-resistant Klebsiella pneumoniae (CRKP) | The source of the isolates is the blood cultures from 461 patients with monomicrobial bacteraemia at the China Medical University Hospital from 1 October to 31 December 2022. | 1 (China Medical University Hospital) | Blood | Bruker Microflex LT/SH, Bruker Daltonics GmbH | - Baseline correction - Intensity normalization - Conversion to m/z ratios - Resampling to a fixed sampling rate of 1 Da - Baseline correction using TopHat - Normalization using square root transformation - Smoothing using Savitzky-Golay-filter - Conversion to m/z ratios - Mass spectra extraction within the range of 20,000-20,000 Da - Min-max scaling to keep signal intensity within 0-1 range - Division into 10-Da bins to mitigate peak shifting - Use of maximum peak intensity as a representative value | - MRSA/MSSA: Accuracy = 0.875, AUC = 0.898, F1 score = 0.904 - CRKP/CSKP: Accuracy = 0.766, AUC = 0.828, F1 score = 0.795 - Detailed metrics in the table on page 4 | - Model: Machine learning (ML) approach combined with MALDI-TOF MS and Rapid Sepsityper protocol - Performance Metrics: - MRSA/MSSA: Accuracy = 0.875, AUC = 0.898, F1 score = 0.904 - CRKP/CSKP: Accuracy = 0.766, AUC = 0.828, F1 score = 0.795 | Not mentioned (the paper does not provide information on external validation performed using an independent dataset) | Within 1 hour | - MRSA/MSSA: 7740-7749 - CRKP/CSKP: 4520-4529
Lin et al., 2025 | Accelerating antimicrobial stewardship: An AI-CDSS approach to combating multidrug-resistant pathogens in the era of increasing resistance | Intensity normalization involved log transformation and was applied within each segment across resistant and susceptible strains. The specific method (e.g., TIC, vector norm, base-peak) is not mentioned. | - Extraction of m/z segments and intensity values using Pyopenms library - Formatting data into Pandas DataFrame - Coding resistance as '1' and susceptibility as '0' - Segmenting m/z segments into 1unit intervals with ±2 units extension and merging overlapping ranges - Log-transformation and normalization of peak intensity values - Prioritization of features based on 'feature_importances' attribute - Exclusion of less informative features with a threshold of 0.01 - Filling missing values with 0 | - Reliance on empirical antimicrobial prescribing protocols leading to potential suboptimal initial treatment. - Limited integration of model outputs into practical clinical decision support tools. - Complexity of bacterial resistance mechanisms challenging diagnosis and treatment. - Limited generalizability due to reliance on samples from specific hospitals. - Severe data imbalances restricting applicability for certain antibiotics. - Exclusion of intermediate resistance results limiting nuanced resistance pattern handling. - Need to address barriers like regulatory approval, clinician trust, and cost considerations for clinical adoption. | Not mentioned (the paper does not provide the size of an external validation dataset distinct from the internal validation set) |  | VITEK MS (bioMérieux, France) | Not reported | - Study type: Retrospective - Setting and timeframe: Conducted at the Tri-Service General Hospital and four regional secondary urban hospitals from January 2021 to December 2023 - Data collection approach: Clinical samples collected from patients with suspected bacterial infections - Validation strategy: Dataset segmented into training (January to September) and validation (October to December) sets; GridSearchCV with 5-fold cross-validation used for model optimization | - Equipment/platform used: bioMérieux, France - Sample preparation methods: Culturing bacteria and using MALDI-TOF MS for identification - Spectral data preprocessing steps: Extracting m/z segments and intensity values using Pyopenms, formatting data into Pandas DataFrame, segmenting m/z segments into 1-unit intervals, log-transforming and normalizing peak intensity values, evaluating features based on importance - Mass range analyzed: m/z range of 2000-2500 - Number of technical replicates per isolate: Not mentioned - Quality control measures: Applying a strict threshold to exclude less informative features, filling missing values with 0 - Spectral data processing software/methods: Pyopenms, Pandas DataFrame | - Specific algorithms used: Random Forest Classifier (RFC), Light Gradient Boosting Machine (LGBM) - Training/validation/test set split ratios and sizes: Samples from January to September for training, October to December for validation - Feature selection or extraction methods: Extracted m/z segments and intensity values using Pyopenms; feature selection based on 'feature_importances' attribute of RFC - Hyperparameter optimization approach: GridSearchCV with 5-fold cross-validation - Cross-validation strategy: 5-fold cross-validation - Any ensemble methods or model combinations: None mentioned - Software/programming languages used: Pyopenms, 'sklearn' | - Specific antibiotics tested: amikacin (AN), gentamicin (GM), piperacillin/tazobactam (TZP), ceftazidime (CAZ), ceftriaxone (CRO), cefoperazone/sulbactam (SFP), flomoxef (FLO), cefepime (FEP), imipenem (IPM), doripenem (DOR), ciprofloxacin (CIP), levofloxacin (LEV) - Resistance mechanisms: Not specified - Reference standard used for resistance determination: VITEK-2 system - Resistance breakpoints or criteria: Binary coding system (resistance as '1', susceptibility as '0') - Multi-drug resistance patterns: Assessed across different antibiotic classes | - Total sample size: 12,967 K. pneumoniae isolates - Source of isolates: Clinical specimens (blood, urine, throat swabs, sputum, body fluids, wound swabs, catheter samples) from Tri-Service General Hospital and four regional secondary urban hospitals - Isolation timeframe: January 2021 to December 2023 - Patient population: Individuals over 20 years of age - Inclusion criteria: Specimens with definitive bacterial growth and clear identification by MALDI-TOF MS - Distribution of resistant vs. susceptible isolates: Detailed in Table 1 | - AUC values: Predominantly ranging from 0.91 to 0.95 - Sensitivity, specificity, PPV, NPV: Generally above 80% - Specific AUC values: - KP: Approximately 0.95 for IPM, DOR, and GM - PA: Exceeding 0.9 for GM, FEP, and LEV - LGBM slightly outperforms RFC in data-intensive scenarios - Lower sensitivity to AN; poor specificity for CRO and FLO in PA | - Time to result: The AI-CDSS provides results up to one day faster than conventional AST methods. - Cost considerations: Not mentioned. - Ease of implementation: The system is implemented as a web-based interface that integrates with existing hospital information systems. - Integration with existing workflows: The AI-CDSS is designed to work with existing systems, facilitating seamless integration. - Real-time or rapid result capabilities: The system provides real-time recommendations. - Potential impact on antimicrobial stewardship: The AI-CDSS supports timely and precise antibiotic administration, potentially reducing antibiotic resistance. | - Generalizability concerns due to reliance on samples from specific hospitals - Technical limitations related to data imbalances for certain antibiotics (AN in KP and AN, CRO, and FLO in PA) - Need for expansion of spectral analysis method to address resistance mechanisms - Exclusion of intermediate resistance results - Implementation barriers: regulatory approval, clinician trust, cost considerations - Need for larger validation studies and expanded datasets - Development of user-friendly interfaces for clinicians | - Klebsiella pneumoniae: 12,967 - Pseudomonas aeruginosa: 9,429 | - Aminoglycosides: Amikacin (AN), Gentamicin (GM) - Penicillins: Piperacillin/Tazobactam (TZP) - Cephalosporins: Ceftazidime (CAZ), Ceftriaxone (CRO), Cefoperazone/Sulbactam (SFP), Flomoxef (FLO), Cefepime (FEP) - Carbapenems: Imipenem (IPM), Doripenem (DOR) - Quinolones: Ciprofloxacin (CIP), Levofloxacin (LEV) | The source of the isolates was clinical samples collected from patients with suspected bacterial infections at the Tri-Service General Hospital and four regional secondary urban hospitals. | 5 | - Blood - Urine - Throat swabs - Sputum - Body fluids - Wound swabs - Catheter samples | bioMérieux, France | - Extraction of m/z segments and intensity values from MALDI-TOF MS data using Pyopenms and formatting into a Pandas DataFrame. - Coding resistance as '1' and susceptibility as '0', excluding intermediate results. - Segmenting m/z segments into intervals to reduce misclassifications. - Log-transforming and normalizing peak intensity values. - Evaluating and prioritizing features based on their importance. - Applying a threshold to exclude less informative features. - Filling missing values with '0'. - Addressing data imbalance through resampling techniques. | - AUC values: predominantly ranging from 0.91 to 0.95 - Sensitivity, specificity, PPV, NPV: primarily exceeded 80% - Detailed metrics: accuracy, sensitivity, specificity, PPV, NPV, F1-score (visualized in Fig. 4) - Validation performance: AUC values >0.8 for most antibiotics - Exceptions: lower sensitivity to AN, poor specificity for CRO and FLO in PA - Specific resistance and sensitivity percentages: provided in table on page 4 | Light Gradient Boosting Machine (LGBM) | Not mentioned (the paper does not mention the use of an independent external validation dataset) | 1 day | Not mentioned (the paper does not provide specific details on the "Top AMR peaks m/z/ m/z biomarker")
Xu et al., 2025 | Rapid detection of carbapenem-resistant Escherichia coli and carbapenem-resistant Klebsiella pneumoniae in positive blood cultures via MALDI-TOF MS and tree-based machine learning models | Intensity normalization was achieved by binning data every 10 Da and using the maximum peak value in each region to represent intensity. Gaussian smoothing was applied to reduce noise. The method of normalization (e.g., TIC, vector norm, base-peak) and whether it was applied per spectrum or globally is not specified. | - Binning of mass spectral peaks every 10 Da within the range of 2000-20,000 Da. - Conversion of data from '.mzML' to '.csv' format with Gaussian smoothing to reduce noise. - Use of SMOTE algorithm to balance sample sizes. - Division of dataset into training and test sets using train_test_split. | - Poor training efficacy of the decision tree model. - Lower efficacy for K. pneumoniae due to high mucus phenotype affecting MALDI-TOF data quality. - Inherent uncertainties and varying effects of different parameters on the models. - Limited generalizability due to single-centre study. | K. pneumoniae: 127 | 127 | VITEK MS IVD system (bioMérieux, Lyon, France) | Columbia blood Agar | - Study type: Retrospective - Setting and timeframe: Laboratory Department of Zhejiang Rong Jun Hospital, Jiaxing, China; October 2021 to September 2024 - Data collection approach: Collection of MALDI-TOF MS data from positive blood cultures of E. coli and K. pneumoniae - Comparison groups or control methods: Not mentioned - Validation strategy: Test score on 30% of original data, ten-fold cross-validation | - Equipment/platform used: VITEK MS IVD system (bioMérieux) - Sample preparation methods: Serum isolation gel method - Spectral data preprocessing steps: Binning data every 10 Da, taking maximum value of peaks to represent intensity, Gaussian smoothing - Mass range analyzed: 2000 Da to 20,000 Da - Number of technical replicates per isolate: 100 shots - Quality control measures: Peak plots with more than 90% identification - Spectral data processing software/methods: Conversion from '.mzML' to '.csv' format using Python libraries (pyteomics, numpy, pandas, scipy, sklearn, matplotlib, imblearn) | - Specific algorithms used: Decision Tree (DT), Random Forest (RF), Gradient Boosting Machine (GBM), eXtreme Gradient Boosting (XGBoost), Extremely Randomized Trees (ERT) - Training/validation/test set split ratios and sizes: 70% training, 30% testing - Feature selection or extraction methods: Binning mass spectral peaks every 10 Da, Gaussian smoothing - Hyperparameter optimization approach: Setting random_state to 0 - Cross-validation strategy: Ten-fold cross-validation - Ensemble methods or model combinations: Not explicitly mentioned - Software/programming languages used: Python with pyteomics, numpy, pandas, scipy, sklearn, matplotlib, imblearn | - Specific antibiotics tested: ertapenem, imipenem, meropenem - Resistance mechanisms: Not specified - Reference standard used for resistance determination: CLSI M100 (33rd edition) - Resistance breakpoints or criteria applied: MIC ≥ 2 µg/ml for ertapenem, MIC ≥ 4 µg/ml for imipenem or meropenem - Multi-drug resistance patterns: Not assessed | - Total sample size: 444 - Breakdown by resistance status: 301 carbapenem-sensitive, 143 carbapenem-resistant - Source of isolates: Laboratory Department of Zhejiang Rong Jun Hospital, Jiaxing, China - Isolation timeframe: October 2021 to September 2024 - Patient population: Positive blood cultures - Inclusion/exclusion criteria: Not mentioned - Distribution of resistant vs. susceptible isolates: 67.79% carbapenem-sensitive, 32.21% carbapenem-resistant (total sample); 79.53% carbapenem-sensitive, 20.47% carbapenem-resistant (127 isolates from positive blood cultures) | - Accuracy for E. coli: DT (0.89), RF (0.92), GBM (0.90), XGBoost (0.92), ERT (0.86) - Accuracy for K. pneumoniae: DT (0.76), RF (0.86), GBM (0.81), XGBoost (0.80), ERT (0.76) - AUC for E. coli: DT (0.95), RF (1.00), GBM (0.99), XGBoost (0.99), ERT (1.00) - AUC for K. pneumoniae: DT (0.78), RF (0.95), GBM (0.93), XGBoost (0.90), ERT (0.95) - Ten-fold cross-validation scores and test set scores are provided in tables. | - Time to result: Reduces time to approximately 0.5 hours. - Cost considerations: Cost-effective compared to traditional methods. - Ease of implementation: MALDI-TOF MS is already used in clinical labs, making it easy to implement. - Integration with existing workflows: Integrates with existing workflows by using MALDI-TOF MS for species identification. - Real-time or rapid result capabilities: Provides rapid detection capabilities. - Potential impact on antimicrobial stewardship: Allows for early clinical intervention and rational use of antibiotics. | - Technical limitations: Quality of MALDI-TOF MS and antimicrobial susceptibility testing data; inherent uncertainties in machine learning models. - Sample size limitations: Single-centre study with limited geographical and clinical diversity. - Generalizability concerns: Limited to one hospital setting; need for larger validation studies. - Performance issues: High mucus phenotype in K. pneumoniae affects data quality and model performance. - Need for larger validation studies: To improve generalizability and address technical limitations. - Implementation barriers: High-quality data required; technical issues with sample preparation and analysis. | - E. coli: 640 - K. pneumoniae: 444 | Carbapenems (specifically, resistance to ertapenem, imipenem, and meropenem in E. coli and K. pneumoniae) | The source of the isolates is the Laboratory Department of Zhejiang Rong Jun Hospital in Jiaxing, China, where they were collected from positive blood cultures using the serum separating gel method. | 1 (Zhejiang Rong Jun Hospital) | Blood | VITEK MS IVD system (bioMérieux, Lyon, France) | - Data Binning: Mass spectral peaks were binned every 10 Da. - Data Conversion: Original data converted from '.mzML' to '.csv' with Gaussian smoothing. - SMOTE Algorithm: Used to balance sample sizes. - Data Splitting: Dataset split into training and test sets using train_test_split. - Random State Control: Randomness controlled by setting random_state to 0. - AUROC Calculation: Used to determine the optimal boundary point. | - E. coli: - AUROC: DT=0.95, RF=1.00, GBM=0.99, XGBoost=0.99, ERT=1.00 - Accuracy: DT=0.89, RF=0.92, GBM=0.90, XGBoost=0.92, ERT=0.86 - K. pneumoniae: - AUROC: DT=0.78, RF=0.95, GBM=0.93, XGBoost=0.90, ERT=0.95 - Accuracy: DT=0.76, RF=0.86, GBM=0.81, XGBoost=0.80, ERT=0.76 | Random Forest (RF) | Yes, external validation was performed using an independent test set to validate the models' performance on predicting susceptibility to carbapenems in E. coli and K. pneumoniae. | 0.5 hours | - For E. coli: 7,870-7,879 m/z - For K. pneumoniae: 4,920-4,929 m/z
Xu and Gao, 2025 | Machine learning powered profiling: Rapid identification of Klebsiella Pneumoniae drug resistance from MALDI-TOF MS | Not mentioned | - Conversion of raw spectral data from ".mzML" to ".csv" format. - Application of Gaussian smoothing to reduce variability. - Extraction of peak features (m/z and intensity) within the 2000-12,000 Da spectral range. - Calculation of peak intensity using a ± 2.5 Da tolerance window. - Use of the maximum intensity within each 5 Da interval as the representative intensity. | - Dispersion of cases in the validation set leading to wide confidence intervals. - Models struggle to generalize when training data does not capture the full range of resistance mechanisms. - Limited predictive power for Levofloxacin due to diverse resistance mechanisms. - Technical constraints of MALDI-TOF MS affecting sensitivity for certain markers. - Biological variability and inter-instrument differences reducing consistency. - Dataset derived from a single hospital may not reflect broader diversity of resistance patterns. | 163 | 163 | Vitek MS IVD system (bioMérieux, Lyon, France) | Columbia blood Agar | - Study type: Retrospective - Setting and timeframe: Zhejiang Rongjun Hospital, December 2023 to January 2025 - Data collection approach: Collection of 484 non-duplicate clinical isolates of K. pneumoniae - Validation strategy: Ten-fold cross-validation and independent test set of 163 isolates | - Equipment/platform used: bioMérieux, Lyon, France - Sample preparation methods: Incubation on Columbia Blood Agar for 18-24 hours - Spectral data preprocessing steps: Conversion from ".mzML" to ".csv", Gaussian smoothing, peak feature extraction - Mass range analyzed: Initially 2000-20,000 Da, analyzed range 2000-12,000 Da - Number of technical replicates per isolate: 100 laser shots - Quality control measures: Spectra with identification rate >90% considered acceptable - Spectral data processing software/methods: Python scripts in Jupyter Lab environment | - Specific algorithms used: Random Forest, XGBoost, AdaBoost, Logistic Regression, Multilayer Perceptron, Support Vector Machine - Training/validation/test set split ratios and sizes: 70:30 split; 30% of the dataset used for testing - Feature selection or extraction methods: Conversion to ".csv" format, Gaussian smoothing, peak features (m/z and intensity) extraction from 2000-12,000 Da range - Hyperparameter optimization approach: Tuning of parameters like n_estimators, max_depth, learning_rate, hidden_layer_sizes, regularization strength, solver - Cross-validation strategy: Ten-fold cross-validation - Ensemble methods or model combinations: Tree-based ensemble methods (RF, XGBoost, AdaBoost) - Software/programming languages used: Python scripts in Jupyter Lab environment, scikit-learn library | - Specific antibiotics tested: Piperacillin/Tazobactam, Ceftazidime, Ceftriaxone, Cefotetan, Aztreonam, Imipenem, Amikacin, Levofloxacin, Co-trimoxazole - Resistance mechanisms: Efflux pumps (e.g., AcrAB-TolC, OqxAB), plasmid-mediated Qnr proteins, drug-modifying enzymes, porin deletions, intrinsic resistance clusters - Reference standard: CLSI M100 (33rd edition) - Resistance breakpoints/criteria: Binary outcome: non-susceptible (resistant or intermediate) = 1, susceptible = 0 - Multi-drug resistance patterns: Simultaneous strain-level resistance profiling for nine drugs | - Total sample size: 484 K. pneumoniae isolates - Source of isolates: Blood, urine, thoracic and abdominal drainage fluids, and alveolar lavage samples from Zhejiang Rongjun Hospital - Isolation timeframe: December 2023 to January 2025 - Patient population: Not explicitly mentioned - Inclusion/exclusion criteria: Non-duplicate clinical isolates - Distribution of resistant vs. susceptible isolates: Not explicitly provided | - Accuracy, sensitivity, specificity: Not mentioned - Area Under the Curve (AUC) values: - XGBoost: 0.80-0.99 - RF: 0.79-0.99 - AdaBoost: 0.81-0.98 - LR: 0.64-0.87 - MLP: 0.70-0.95 - SVM: 0.63-0.87 - Positive and negative predictive values: Not mentioned - F1-scores: Not mentioned - Confidence intervals: - Co-trimoxazole: 0.68-0.89 - Amikacin: Not mentioned - Levofloxacin: 0.56-0.73 - Performance differences between training and test sets: Not mentioned - Comparative performance between different ML algorithms: - Tree-based models (RF, XGBoost, AdaBoost) outperformed linear models (LR, SVM) and MLP | - Time to result: Reduced to 1-2 hours compared to conventional methods that take 1-2 days. - Cost considerations: Described as "cost-effective." - Ease of implementation: Leverages existing MALDI-TOF MS technology, suggesting ease of integration into clinical labs. - Integration with existing workflows: Supports personalized regimen design by predicting susceptibility to multiple antibiotics. - Real-time or rapid result capabilities: Enables early targeted therapy and improves patient outcomes. - Potential impact on antimicrobial stewardship: Enhances stewardship by providing comprehensive susceptibility predictions for multiple antibiotics. | - Technical limitations of MALDI-TOF MS: Limited sensitivity for low-abundance peptides and high-molecular-weight resistance determinants. - Sample size limitations: Dataset from a single hospital may not reflect broader resistance patterns. - Generalizability concerns: Need for multicenter data to enhance real-world applicability. - Performance issues: Limited predictive power for Levofloxacin due to diverse resistance mechanisms. - Need for larger validation studies: To improve generalizability and address dispersed validation set distribution. - Implementation barriers: Challenges in clinical implementation due to technical and biological variability. - Areas requiring further research: Optimization to ensure clinically reliable accuracy and understanding of resistance mechanisms. | 484 | - Amikacin - Piperacillin/Tazobactam - Imipenem - Co-trimoxazole - Levofloxacin - Ceftazidime - Ceftriaxone - Cefotetan - Aztreonam | Zhejiang Rongjun Hospital; Sources included blood, urine, thoracic and abdominal drainage fluids, and alveolar lavage samples. | 1 (Zhejiang Rongjun Hospital) | - Blood - Urine - Thoracic drainage fluids - Abdominal drainage fluids - Alveolar lavage samples | bioMérieux, Lyon, France | - Spectral range analysis: 2000-12,000 Da - Peak intensity calculation: ± 2.5 Da tolerance window - Peak extraction: 2000 characteristic peaks - Gaussian smoothing: Applied using Python scripts - SMOTE: Used to address class imbalance - Data split: 70:30 training-test split - Cross-validation: Ten-fold cross-validation | - Test score - Ten-fold cross-validation - Accuracy - Specificity - F1 score - AUROC - 95% CI - AUROC values for each antibiotic: - Piperacillin/Tazobactam: 0.69 - Ceftazidime: 0.71 - Ceftriaxone: 0.70 - Cefotetan: 0.71 - Aztreonam: 0.67 - Imipenem: 0.73 - Amikacin: 0.75 - Levofloxacin: 0.65 - Co-trimoxazole: 0.80 | XGBoost model | Yes, an independent validation dataset of 163 isolates was used for external validation. | 1-2 hours | - Amikacin: 4517.5 ± 2.5 Da - Co-trimoxazole: 6022.5 ± 2.5 Da
Ye et al., 2025 | Rapid Identification of Carbapenemase Subtypes in Klebsiella pneumoniae Using MALDI-TOF MS Combined With Convolutional Neural Networks | Normalization was applied as part of preprocessing, but the specific method (e.g., TIC, vector norm, base-peak) and whether it was applied per spectrum or globally is not mentioned. | The pre-processing scheme utilised for the raw MALDI-TOF MS spectra included smoothing, baseline correction, and normalization. | - Limited global applicability due to exclusion of certain carbapenemase subtypes like VIM and IMP. - Restricted diversity of the cohort may compromise generalizability. - Lack of external validation with independent strains raises concerns about real-world accuracy. - Spectral variability due to technical factors like extraction protocols and instrument parameters. - Need for expanding datasets and optimizing model architectures for better generalizability. | Not mentioned (the paper does not provide information on the size of an external validation dataset) |  | Chongqing ZhongyuanHuiji Co. MALDI-TOF MS spectrometer | Columbia blood Agar | - Study type: Retrospective - Setting and timeframe: Affiliated Hospital of Xuzhou Medical University, 2018-2023 - Data collection approach: Analysis of 205 clinical CRKP isolates using MALDI-TOF MS and CNNs - Comparison groups: Different carbapenemase subtypes (KPC, NDM, OXA-48) and machine learning algorithms (CNN, RF, SVM, AdaBoost) - Validation strategy: Fivefold cross-validation | - Equipment/platform used: Standard MALDI-TOF MS spectrometer - Sample preparation methods: Suspended in sterile water, ethanol treatment, centrifugation, resuspension in formic acid and acetonitrile - Spectral data preprocessing steps: Smoothing, baseline correction, normalization - Mass range analyzed: 2000-20,000 m/z - Number of technical replicates per isolate: 12 - Quality control measures: Reference strains for calibration, confidence score ≥ 2.0 for species identification - Spectral data processing software/methods: CNN implemented in PyTorch | - Specific algorithms used: Convolutional Neural Network (CNN), Random Forest (RF), Support Vector Machine (SVM), AdaBoost - Training/validation/test set split ratios and sizes: Fivefold cross-validation with 80% training and 20% validation - Feature selection or extraction methods: Automated through hierarchical convolutional architecture of CNN - Hyperparameter optimization approach: Cross-entropy loss function with stochastic gradient descent optimizer - Cross-validation strategy: Fivefold cross-validation - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: PyTorch | - Specific antibiotics tested: cephalosporins (e.g., ceftazidime), carbapenems (e.g., meropenem, imipenem), sulfamethoxazole/trimethoprim, aminoglycosides (e.g., amikacin) - Resistance mechanisms: KPC, NDM, OXA-48 - Reference standard: Clinical and Laboratory Standards Institute document M100 (CLSI M100, 31st edition, 2021) - Resistance breakpoints: Interpreted per CLSI guidelines - Multi-drug resistance patterns: Near-universal resistance to certain antibiotics among KPC and NDM-producing isolates | - Total sample size: 205 K. pneumoniae isolates - Breakdown by resistance status: 111 KPC-producing, 47 NDM-producing, 47 OXA-48-producing - Source of isolates: Clinical specimens from the Affiliated Hospital of Xuzhou Medical University, China - Isolation timeframe: 2018 to 2023 - Patient population: KPC-producing strains more prevalent in adults; NDM-producing strains more prevalent in pediatric patients - Inclusion criteria: Clinical CRKP isolates from the specified hospital and timeframe - Distribution of resistant vs. susceptible isolates: All isolates are carbapenem-resistant | - Accuracy: 96.06% - Precision: 96.10% - Recall: 96.06% - F1 score: 96.03% - AUC: 0.99 - Subtype-specific accuracies: - KPC: 98.88% - NDM: 85.84% - OXA-48: 96.77% | - Time to result: Results can be obtained within 20 minutes after bacterial culture. - Cost considerations: Significantly lower cost compared to molecular detection methods. - Ease of implementation: Modular design allows for easy integration with local strain databases. - Integration with existing workflows: Potential for deep integration with hospital information systems. - Real-time or rapid result capabilities: Results are available within 20 minutes. - Potential impact on antimicrobial stewardship: Reduces reliance on empirical broad-spectrum regimens, mitigating resistance selection pressures. | - Technical limitations: Regional variations in carbapenemase subtypes limit global applicability. - Sample size limitations: Restricted diversity of strains from a single center may compromise generalizability. - Generalizability concerns: Absence of external validation with independent strains raises concerns about real-world diagnostic accuracy. - Performance issues: Spectral variability due to technical factors requires standardized workflows. - Need for larger validation studies: To improve generalizability and address underrepresented subtypes. - Implementation barriers: Need for further research to optimize model architectures and integrate regional epidemiological patterns. - Areas requiring further research: Expanding datasets to include underrepresented subtypes and optimizing model architectures. | 205 | - β-lactams (e.g., cephalosporins, carbapenems) - Sulfamethoxazole/trimethoprim - Aminoglycosides (e.g., amikacin) - Resistance profiles: - Near-universal resistance to cephalosporins and carbapenems among KPC and NDM-producing isolates. - Lower resistance rates to sulfamethoxazole/trimethoprim and amikacin. - Notable variation in imipenem resistance among subtypes. | The Affiliated Hospital of Xuzhou Medical University | 1 (the Affiliated Hospital of Xuzhou Medical University) | Not mentioned (the paper does not specify the source of the specimens, such as blood or urine) | Chongqing ZhongyuanHuiji Co. MALDI-TOF MS spectrometer | - Smoothing - Baseline correction - Normalization - Protein extraction using formic acid - Spectral acquisition using MALDI-TOF MS - Exclusion of low-confidence spectra | - CNN: Accuracy = 96.06%, Precision = 96.10%, Recall = 96.06%, F1 score = 96.03%, AUC = 0.99 - RF: Accuracy = 0.9289, AUC = 0.9917 - SVM: Accuracy = 0.8394, AUC = 0.9558 - AdaBoost: Accuracy = 0.9126, AUC = 0.9656 | Convolutional Neural Network (CNN) model | Not mentioned (the paper explicitly states that external validation was not performed) | 20 minutes | Not mentioned (the paper does not specify any top AMR peaks or m/z biomarkers)
Xu , 2024 | Modelling the rapid detection of Carbapenemase-resistant Klebsiella pneumoniae based on machine learning and matrix-assisted laser desorption/ionization time-of-flight mass spectrometry | Not mentioned | K-means clustering to obtain the best "feature peaks" and determine the "center of mass" for data organization; no dimensionality reduction was used. | The machine learning models may exhibit limitations in their applicability due to their reliance on local datasets generated by the same instrument. There are potential biases in the model's predictive accuracy when applied to data from instruments of varying brands or when assessing bacterial strains from regions with distinct endemic profiles. | Not mentioned (the paper does not provide information on an external validation dataset) |  | Vitek MS RUO system from bioMérieux, Lyon, France | Columbia blood Agar | - Study type: Retrospective - Setting and timeframe: Zhejiang Rong Jun Hospital, China; January 2022 to December 2023 - Data collection approach: Collection of MALDI-TOF MS peaks from selected strains - Comparison groups: CR-KP vs. CS-KP - Validation strategy: Train-test split and ten-fold cross-validation | - Equipment/platform used: bioMérieux MALDI-TOF MS system - Sample preparation methods: Cultured on Columbia blood plates for 18-24 hours, then applied to target plate with α-cyano-4-hydroxycinnamic acid matrix solution - Spectral data preprocessing steps: VITEK MS RUO system, matched against SARAMIS version 4.12 - Mass range analyzed: 2000 Da to 20,000 Da - Number of technical replicates per isolate: 100 shots - Quality control measures: Peak plots with more than 90% identification - Spectral data processing software/methods: Imported into Excel, K-means algorithm for clustering | - Specific algorithms used: Random Forest, Support Vector Machine, Logistic Regression, Xgboost - Training/validation/test set split ratios and sizes: 70% training, 30% testing - Feature selection or extraction methods: K-means clustering - Hyperparameter optimization approach: GridSearchCV - Cross-validation strategy: Ten-fold cross-validation - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: Python with sklearn and xgboost libraries | - Specific antibiotics tested: ertapenem, imipenem, meropenem - Resistance mechanisms: carbapenemase activity - Reference standard: CLSI M100 (32nd edition) - Resistance breakpoints: MIC ≥ 2 μg/ml for ertapenem, MIC ≥ 4 μg/ml for imipenem or meropenem - Multi-drug resistance patterns: Not mentioned | - Total sample size: 240 - Breakdown by resistance status: 80 CR-KP, 160 CS-KP - Source of isolates: Clinical specimens from Zhejiang Rong Jun Hospital, Jiaxing, China - Isolation timeframe: January 2022 to December 2023 - Patient population: Not specified - Inclusion criteria: Based on drug sensitivity results (MIC ≥ 2 μg/ml for ertapenem and MIC ≥ 4 μg/ml for imipenem or meropenem) - Distribution of resistant vs. susceptible isolates: 80 CR-KP to 160 CS-KP | - Accuracy: Not mentioned - Sensitivity and Specificity: - Random Forest: Sensitivity: 90.48%, Specificity: 98.04% - Support Vector Machine: Sensitivity: 90.48%, Specificity: 92.16% - Logistic Regression: Sensitivity: 95.24%, Specificity: 88.24% - Xgboost: Sensitivity: 95.24%, Specificity: 92.16% - Area Under the Curve (AUC) values: - Random Forest: 0.99 - Support Vector Machine: 0.97 - Logistic Regression: 0.96 - Xgboost: 0.97 - Positive and negative predictive values: Not mentioned - F1-scores: Not mentioned - Confidence intervals: Not mentioned - Performance differences between training and test sets: - Random Forest: Test set score: 0.94, Ten-fold cross-validation score: 0.84 - Support Vector Machine: Test set score: 0.91, Ten-fold cross-validation score: 0.81 - Logistic Regression: Test set score: 0.90, Ten-fold cross-validation score: 0.81 - Xgboost: Test set score: 0.93, Ten-fold cross-validation score: 0.85 - Comparative performance between different ML algorithms: - AUC values and test set scores vary between models | - Time to result compared to conventional methods: Shortens in-laboratory reporting time. - Cost considerations or comparisons: Not mentioned. - Ease of implementation in clinical labs: Can be generalized to other bacteria. - Integration with existing workflows: Not mentioned. - Real-time or rapid result capabilities: Provides fast and reliable identification results; can determine CR-KP within 1 hour. - Potential impact on antimicrobial stewardship: Not mentioned. | - Technical limitations: The model may exhibit limitations in its applicability due to being suited to local datasets generated by the same instrument. - Generalizability concerns: Potential biases in predictive accuracy when applied to data from instruments of varying brands or regions with distinct endemic profiles. | 240 | Carbapenem resistance prediction using machine learning and MALDI-TOF MS for Klebsiella pneumoniae | Zhejiang Rong Jun Hospital (Jiaxing, China) | 1 (Zhejiang Rong Jun Hospital) | Not mentioned (the paper does not specify the specimen source isolation) | bioMérieux, Lyon, France (VITEK MS RUO system) with SARAMIS version 4.12 | - K-means clustering to determine the "center of mass" of feature peaks - Use of "center of mass" as a feature peak for classification - Implementation using Python and sklearn library - No dimensionality reduction used; instead, parameterization was used to improve model performance | - Random Forest: AUC = 0.99, Test Set Score = 0.94, Ten-fold Cross-validation Score = 0.84, Specificity = 98.04%, Sensitivity = 90.48% - Support Vector Machine: AUC = 0.97, Test Set Score = 0.91, Ten-fold Cross-validation Score = 0.81, Specificity = 92.16%, Sensitivity = 90.48% - Logistic Regression: AUC = 0.96, Test Set Score = 0.90, Ten-fold Cross-validation Score = 0.81, Specificity = 88.24%, Sensitivity = 95.24% - Xgboost: AUC = 0.97, Test Set Score = 0.93, Ten-fold Cross-validation Score = 0.85, Specificity = 92.16%, Sensitivity = 95.24% | Random Forest model with max_depth=6, max_features=6, min_samples_leaf=1, min_samples_split=5, n_estimators=139 | Not mentioned (no information on external validation dataset is included in the paper) | The paper suggests that the method can shorten the in-laboratory reporting time, but does not provide a specific numerical value for the improved time to detection. A referenced study by Jiaxin Yu et al. achieved detection within 1 hour. | 4519 m/z
Wang et al., 2022 | Large-Scale Samples Based Rapid Detection of Ciprofloxacin Resistance in Klebsiella pneumoniae Using Machine Learning Methods | Vector normalization (scaling by mean and standard deviation) | 1. Peak smoothing using Savitzky-Golay filter 2. Baseline correction using Tof-hat filter 3. Calibration with an external standard calibrator 4. Selection of peaks with signal-to-noise ratio ≥ 2 5. Transformation into pseudo-ion vectors with 900 pseudo-ions 6. Scaling of pseudo-ion vectors by mean and standard deviation | - The models may not generalize to other bacterial species and antibiotics. - Absence of biological validations for selected features. - Difficulty in verifying accurate sub-strain information due to data limitations. - Potential for improvement using advanced techniques like deep learning. | Not mentioned (no information on an external validation dataset is included in the paper) |  | Bruker Microflex LT MALDI-TOF System | Blood specimens: Trypticase soy broth and blood plate (BP) agar Fluid and respiratory specimens: BP agar, eosin methylene blue (EMB) agar, CNA agar, and chocolate agar Wound specimens: BP, EMB, CNA, and chocolate agars Urine specimens: BP and EMB agars | - Study type: Longitudinal study - Setting and timeframe: Change Gung Memorial Hospitals (CGMH) at the Linkou branch, June 2013 to February 2018 - Data collection approach: Specimens collected from patients over the specified timeframe - Comparison groups or control methods: Comparison between different machine learning models (SVM, XGB, LR, RF) - Validation strategy: Temporal cross-validation with training set covering 85% of samples and remaining used for performance testing | - Equipment/platform used: Microflex LT MALDI-TOF System, Bruker Daltonik GmbH - Sample preparation methods: Smearing cultivated colonies onto a MALDI steel target plate with formic acid, drying, adding matrix solution, and drying again - Spectral data preprocessing steps: Calibration with external standard calibrator, peak smoothing (Savitzky-Golay filter), baseline correction (Tof-hat filter), selecting peaks with signal-to-noise ratio 2 - Mass range analyzed: 2,000 to 20,000 m/z - Number of technical replicates per isolate: Not mentioned - Quality control measures: Calibration with external standard, selecting peaks with signal-to-noise ratio 2 - Spectral data processing software/methods: Not explicitly mentioned | - Specific algorithms used: Logistic Regression (LR), Support Vector Machine (SVM), Random Forest (RF), Extreme Gradient Boosting (XGB) - Training/validation/test set split ratios and sizes: Training set includes samples from June 1, 2013 to June 31, 2017 (85% of total samples), remaining samples used for testing - Feature selection or extraction methods: Chi-square test for non-spectrometric features; Wilcoxon rank-sum test, F-test, and KS test for pseudo-ions - Hyperparameter optimization approach: L1-regularization for LR with log (λ) = -4.77; radial kernel with penalty parameter C = 1.5 for SVM; 'softmax' mode with max depth = 3 and 527 iterations for XGB - Cross-validation strategy: Fivefold cross-validation for all models except RF - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: R with packages 'glmnet' for LR, 'kernlab' for SVM, 'randomForest' for RF, 'xgboost' for XGB | - Specific antibiotics tested: Ciprofloxacin - Resistance mechanisms: Not specified - Reference standard used for resistance determination: Disk diffusion method as per CLSI M100 guidelines - Resistance breakpoints or criteria applied: ATS breakpoints listed in CLSI M100 - Multi-drug resistance patterns: Not assessed | - Total sample size: 16,997 - Breakdown by resistance status: 11,354 CISKP, 4,428 CIRKP - Source of isolates: Clinical specimens from Change Gung Memorial Hospitals (CGMH) at the Linkou branch - Isolation timeframe: June 2013 to February 2018 - Patient population: Not specifically detailed - Inclusion/exclusion criteria: Elimination of intermediate and incomplete samples - Distribution of resistant vs. susceptible isolates: 11,354 CISKP, 4,428 CIRKP | - SVM and XGB: AUC = 0.89, Specificity = 0.80, Sensitivity = 0.80, Accuracy = 0.83 (SVM) [95% CI: (0.81, 0.84)], 0.82 (XGB) [95% CI: (0.80, 0.83)] - LR: AUC = 0.86, Sensitivity = 0.76, Specificity = 0.76 - RF: AUC = 0.85 - Comparative performance: SVM and XGB slightly better than LR and RF | The clinical advantages include a significant reduction in time to result (less than 2 hours with MALDI-TOF MS), lower costs compared to traditional methods, potential ease of implementation in clinical labs due to integration with existing workflows, and the ability to provide rapid and accurate results that can aid in antimicrobial stewardship by improving initial antibiotic selection. | - Technical limitations: The models may not generalize to other bacterial species and antibiotics. - Sample size limitations: Data limitations affect the verification of accurate sub-strain information. - Generalizability concerns: Models may not be applicable to other bacterial species and antibiotics. - Performance issues: Lack of biological validation for selected features. - Need for larger validation studies: Biological validation is needed to confirm feature importance. - Implementation barriers: Challenges in verifying accurate sub-strain information. - Areas requiring further research: Use of advanced machine learning techniques like deep learning. | 15,782 | Ciprofloxacin resistance in Klebsiella pneumoniae (quinolone class) | The source of the isolates is the Change Gung Memorial Hospitals (CGMH) at the Linkou branch, collected from a longitudinal study involving 16,697 participants from June 2013 to February 2018. | 1 (Change Gung Memorial Hospitals at the Linkou branch) | - Blood samples (B) - Body fluid samples (F) - Wound samples (W) - Respiratory samples (R) - Urinary samples (U) - Other types (O) | Microflex LT MALDI-TOF System, Bruker Daltonik GmbH | 1. Peak Smoothing and Baseline Correction 2. Signal-to-Noise Ratio Selection 3. Pseudo-ion Vector Creation 4. Grouping of Specimen Carriers 5. Age Transformation 6. Chi-square Test 7. Wilcoxon Rank-sum Test and Fold Change Selection 8. Feature Ranking 9. Final Feature Matrix Formation | - SVM: AUC = 0.89, Specificity = 0.80, Sensitivity = 0.80 - XGB: AUC = 0.89, Specificity = 0.80, Sensitivity = 0.80 - LR: AUC = 0.85, Specificity = 0.76, Sensitivity = 0.76 - RF: AUC = 0.85, Specificity = 0.47, Sensitivity = 0.94 | SVM and XGB models, both achieving an AUC of 0.89 and showing high stability across different clusters. | Not mentioned (no information on external validation using an independent dataset) | Less than 2 hours after cell culture | - 2,069 m/z - 4,367 m/z - 5,382 m/z - 6,291 m/z
Jian et al., 2024a | Pioneering Klebsiella Pneumoniae Antibiotic Resistance Prediction With Artificial Intelligence-Clinical Decision Support System-Enhanced Matrix-Assisted Laser Desorption/Ionization Time-of-Flight Mass Spectrometry: Retrospective Study | Intensity normalization was achieved through logarithmic transformation of peak intensity values within segmented m/z ranges, applied per spectrum. | 1. Extraction of mass-to-charge (m/z) ratios and intensity values from MALDI-TOF MS data. 2. Resampling strategy using downsampling and upsampling to balance the dataset. 3. Segmentation of m/z ranges into 1-unit intervals with a ± 2-unit tolerance. 4. Logarithmic transformation of peak intensity values. 5. Filtering out segments with feature importance below 0.01 as background noise. | - Exclusion of specimens with intermediate AST results may affect model generalizability. - Limitations when resistance mechanisms are not mediated at the protein-peptide level or fall outside the m/z ratio range of 2000-20,000. - Need for future validation studies to further improve and validate the models. | Not mentioned (the paper does not specify the size of an external validation dataset) |  | The MALDI-TOF MS instrument used was provided by bioMérieux, France. | Not reported | - Study type: Retrospective - Setting and timeframe: Conducted from January 2021 to December 2023 at Tri-Service General Hospital and 4 district hospitals - Data collection approach: Examined 165,299 bacterial specimens using MALDI-TOF MS technology - Validation strategy: Used a separate validation set (samples collected from October to December) for model validation | - Equipment/platform used: bioMérieux, France - Sample preparation methods: Not explicitly detailed - Spectral data preprocessing steps: Extracting m/z ratios and intensity values, resampling, segmenting m/z ranges, logarithmic transformation, filtering out low feature importance segments - Mass range analyzed: Not explicitly mentioned - Number of technical replicates per isolate: Not specified - Quality control measures: Data balancing, filtering out background noise - Spectral data processing software/methods: sklearn.utils for resampling and feature engineering | - Specific algorithms used: Logistic Regression, Linear Discriminant Analysis, Random Forest, Gradient Boosting Classifier, AdaBoost Classifier, XGBoost, light gradient boosting machine (LGBM) - Training/validation/test set split ratios and sizes: Training set (January to September), Validation set (October to December) - Feature selection or extraction methods: Segmentation of m/z ranges into 1-unit intervals, logarithmic transformation of peak intensity values, filtering out low-importance segments - Hyperparameter optimization approach: Grid search - Cross-validation strategy: Not explicitly mentioned - Any ensemble methods or model combinations: Not explicitly mentioned - Software/programming languages used: sklearn.utils for resampling | - Specific antibiotics tested: Levofloxacin (LEV) and Ciprofloxacin (CIP) - Resistance mechanisms: Not specified - Reference standard: Clinical and Laboratory Standards Institute (CLSI) breakpoints - Resistance breakpoints or criteria: CLSI breakpoints - Multi-drug resistance patterns: Not assessed | - Total sample size: 165,299 bacterial specimens; 11,996 KP isolates - Source of isolates: Clinical specimens from Tri-Service General Hospital and 4 district hospitals - Isolation timeframe: January 2021 to December 2023 - Patient population: Not specified - Inclusion/exclusion criteria: Excluded specimens with intermediate AST results (869 for LEV, 373 for CIP) - Distribution of resistant vs. susceptible isolates: Not explicitly detailed | - Accuracy: - RF: 0.87 (LEV), 0.86 (CIP) - GBC: 0.87 (LEV), 0.87 (CIP) - XGBoost: 0.87 (LEV), 0.85 (CIP) - LGBM: 0.90 (LEV), 0.85 (CIP) - AdaBoost: 0.77 (LEV), 0.85 (CIP) - LR: 0.78 (LEV), 0.66 (CIP) - LDA: 0.77 (LEV), 0.64 (CIP) - Sensitivity: - RF: 0.88 (LEV), 0.89 (CIP) - GBC: 0.92 (LEV), 0.88 (CIP) - XGBoost: 0.89 (LEV), 0.86 (CIP) - LGBM: 0.88 (LEV), 0.84 (CIP) - AdaBoost: 0.81 (LEV), 0.85 (CIP) - LR: 0.77 (LEV), 0.74 (CIP) - LDA: 0.80 (LEV), 0.81 (CIP) - Specificity: - RF: 0.86 (LEV), 0.82 (CIP) - GBC: 0.84 (LEV), 0.86 (CIP) - XGBoost: 0.86 (LEV), 0.85 (CIP) - LGBM: 0.91 (LEV), 0.85 (CIP) - AdaBoost: 0.74 (LEV), 0.84 (CIP) - LR: 0.78 (LEV), 0.57 (CIP) - LDA: 0.75 (LEV), 0.43 (CIP) - AUC: - RF: 0.95 (LEV), 0.95 (CIP) - GBC: 0.95 (LEV), 0.95 (CIP) - XGBoost: 0.95 (LEV), 0.94 (CIP) - LGBM: 0.95 (LEV), 0.92 (CIP) - AdaBoost: 0.87 (LEV), 0.90 (CIP) - LR: 0.84 (LEV), 0.73 (CIP) - LDA: 0.80 (LEV), 0.68 (CIP) - PPV: - RF: 0.81 (LEV), 0.86 (CIP) - GBC: 0.80 (LEV), 0.89 (CIP) - XGBoost: 0.82 (LEV), 0.87 (CIP) - LGBM: 0.87 (LEV), 0.87 (CIP) - AdaBoost: 0.69 (LEV), 0.87 (CIP) - LR: 0.71 (LEV), 0.68 (CIP) - LDA: 0.69 (LEV), 0.63 (CIP) - NPV: - RF: 0.91 (LEV), 0.86 (CIP) - GBC: 0.93 (LEV), 0.85 (CIP) - XGBoost: 0.92 (LEV), 0.83 (CIP) - LGBM: 0.92 (LEV), 0.82 (CIP) - AdaBoost: 0.85 (LEV), 0.83 (CIP) - LR: 0.83 (LEV), 0.64 (CIP) - LDA: 0.84 (LEV), 0.65 (CIP) - F1-score: - RF: 0.85 (LEV), 0.87 (CIP) - GBC: 0.86 (LEV), 0.88 (CIP) - XGBoost: 0.85 (LEV), 0.86 (CIP) - LGBM: 0.88 (LEV), 0.86 (CIP) - AdaBoost: 0.74 (LEV), 0.86 (CIP) - LR: 0.74 (LEV), 0.70 (CIP) - LDA: 0.74 (LEV), 0.71 (CIP) | - Time to result: Reduced to approximately 3 days compared to conventional 4-day timeline. - Cost considerations: Not explicitly mentioned, but potential for cost benefits through streamlined processes. - Ease of implementation: Facilitated by its ability to interpret complex data and recommend personalized treatment options. - Integration with existing workflows: Streamlines evaluation process from initial culture to final decision-making. - Real-time or rapid result capabilities: Provides real-time data-driven insights and rapid predictions of antibiotic resistance. - Potential impact on antimicrobial stewardship: Enhances stewardship by ensuring treatments are appropriate and effective. | - Technical limitations: Resistance mechanisms not mediated at the protein-peptide level or outside the m/z ratio range of 2000-20,000. - Sample size limitations: Exclusion of specimens with intermediate AST results. - Generalizability concerns: Need to incorporate intermediate cases and diverse data modalities. - Performance issues: Limited applicability for certain resistance types. - Need for larger validation studies: Future studies required to validate results further. | 11,996 | The antibiotic resistance prediction in this study focused on quinolones, specifically levofloxacin and ciprofloxacin. | Tri-Service General Hospital and four district hospitals in Taiwan | 5 | Not mentioned (the paper does not specify the source of the specimens) | bioMérieux, France | 1. Extraction of mass-to-charge (m/z) ratios and intensity values from MALDI-TOF MS data. 2. Addressing data imbalance through resampling (downsampling over-represented classes and upsampling underrepresented classes). 3. Feature engineering: segmenting m/z ranges into 1-unit intervals with a ± 2-unit tolerance. 4. Logarithmic transformation of peak intensity values. 5. Filtering out segments with feature importance value below 0.01. | - Random Forest: Training AUC=0.99, Testing AUC=0.95, Accuracy=0.87, Sensitivity=0.88, Specificity=0.86, PPV=0.81, NPV=0.91, F1-score=0.85 - Light Gradient Boosting Machine: Training AUC=0.99, Testing AUC=0.95, Accuracy=0.90, Sensitivity=0.88, Specificity=0.91, PPV=0.87, NPV=0.92, F1-score=0.88 - Gradient Boosting Classifier: Training AUC=0.99, Testing AUC=0.95, Accuracy=0.87, Sensitivity=0.92, Specificity=0.84, PPV=0.80, NPV=0.93, F1-score=0.86 - XGBoost: Training AUC=0.99, Testing AUC=0.95, Accuracy=0.87, Sensitivity=0.89, Specificity=0.86, PPV=0.82, NPV=0.92, F1-score=0.85 - AdaBoost: Training AUC=0.99, Testing AUC=0.87, Accuracy=0.77, Sensitivity=0.81, Specificity=0.74, PPV=0.69, NPV=0.85, F1-score=0.74 - Logistic Regression: Training AUC=0.91, Testing AUC=0.84, Accuracy=0.78, Sensitivity=0.77, Specificity=0.78, PPV=0.71, NPV=0.83, F1-score=0.74 - Linear Discriminant Analysis: Training AUC=0.96, Testing AUC=0.80, Accuracy=0.77, Sensitivity=0.80, Specificity=0.75, PPV=0.69, NPV=0.84, F1-score=0.74 | Random Forest classifier | Yes, external validation was performed using a temporal split of the dataset. Samples collected from October to December were used as the validation set to test the models trained on data from January to September. The validation included 3339 cases of LEV resistance and 3487 cases of CIP resistance, validated against traditional culture-based AST at Tri-Service General Hospital. | The improved time to detection is approximately 3 days, which is a reduction from the traditional 5-day timeline for culturing and AST interpretation. | Not mentioned (the paper does not explicitly list the "Top AMR peaks m/z/ m/z biomarker")
Huang et al., 2020 | Detection of carbapenem-resistant Klebsiella pneumoniae on the basis of matrix-assisted laser desorption ionization time-of-flight mass spectrometry by using supervised machine learning approach | Not mentioned | - Peak matching: Utilized Mass-Up software with a tolerance of 300 ppm to unify peak values among samples. - Feature selection: Used the Student's t-test to compare the distribution of each m/z ratio and generate p-values, selecting significant peaks based on these p-values. | - Limited availability of CRKP isolates leading to a relatively small sample size. - High dimensionality of data compared to sample size, which can cause standard pattern classification algorithms to fail. - Consistent misclassification of certain isolates by some algorithms. - Repeated misclassification of a specific CRKP isolate by the RF algorithm. - Potential limitations in the predictive capabilities of algorithms other than RF due to the lack of techniques like bagging and CART-split criterion. | Not mentioned (the paper does not specify the size of an external validation dataset) |  | Vitek MS system (bioMérieux SA) | Blood agar | - Study type: Retrospective - Setting and timeframe: Clinical microbiology laboratory, January 2016 to October 2017 - Data collection approach: Selection of isolates from frozen stocks, use of MALDI-TOF MS - Comparison groups: CRKP vs. CSKP isolates - Validation strategy: Leave-one-out cross-validation | - Equipment/platform used: Vitek MS system (bioMe ´rieux SA) - Sample preparation methods: Direct colony method with matrix solution (α-cyano-4-hydroxycinnamic acid) - Spectral data preprocessing steps: Baseline correction, denouncing, peak detection - Mass range analyzed: 2,000 to 20,000 Da - Number of technical replicates per isolate: One spectrum per isolate - Quality control measures: E. coli reference strain ATCC 8739 for calibration, data count (DC) score for protein extraction quality - Spectral data processing software/methods: Mass-Up for peak matching, Mathematica software v. 12 for data analysis | - Specific algorithms used: Random Forest (RF), Support Vector Machine (SVM), Nearest Neighbors, Naïve Bayes, Logistic Regression - Training/validation/test set split ratios and sizes: Leave-one-out (L1O) cross-validation - Feature selection or extraction methods: Student's t-test for peak selection - Hyperparameter optimization approach: Default settings used - Cross-validation strategy: Leave-one-out (L1O) cross-validation - Any ensemble methods or model combinations: Random Forest (RF) uses bagging and CART-split criterion - Software/programming languages used: Mathematica software v. 12, Mass-Up software | - Specific antibiotics tested: imipenem, meropenem - Resistance mechanisms: carbapenemase production, impermeability - Reference standard: Vitek 2 system with AST N320 card - Resistance breakpoints: updated 2018 CLSI breakpoints - Multi-drug resistance patterns: carbapenemase (+ or-ESBL), impermeability CARBA (+ESBL or +HL AmpC) | - Total sample size and breakdown by resistance status: 95 isolates (46 CRKP, 49 CSKP) - Source of isolates: Clinical specimens (pus/wound, sputum, urine, blood, body fluids) - Isolation timeframe: January 2016 to October 2017 - Patient population: One isolate per patient - Inclusion/exclusion criteria: CRKP isolates were resistant to imipenem or meropenem with a 'carbapenemase' phenotype - Distribution of resistant vs. susceptible isolates: CRKP (63% carbapenemase + or -ESBL, 37% carbapenemase + impermeability); CSKP phenotypes varied | - Accuracy: 97% (RF with 80 peaks) - Sensitivity: 93% (RF with 80 peaks) - Specificity: 100% (RF with 80 peaks) - Ranges for RF: Accuracy (81%-97%), Sensitivity (76%-96%), Specificity (82%-100%) - Comparative performance: RF outperformed other algorithms | - Time to result: Faster than conventional susceptibility testing methods. - Cost considerations: Cost-effective compared to traditional methods. - Ease of implementation: Simplified process using direct colony method; easy integration into existing workflows. - Integration with existing workflows: Can be incorporated without altering current protocols. - Real-time or rapid result capabilities: Provides rapid results for timely clinical decision-making. - Potential impact on antimicrobial stewardship: Enhances clinical decision-making and patient outcomes by guiding appropriate treatment and infection control. | - Technical limitations: Small sample size of CRKP isolates, variability in peak selection across iterations. - Sample size limitations: Limited availability of CRKP isolates. - Generalizability concerns: Need for additional studies before clinical integration. - Performance issues: Potential misidentification by some algorithms. - Need for larger validation studies: Required before clinical integration. - Implementation barriers: Need for further research to confirm potential biomarkers. - Areas requiring further research: Broad application to other classification problems. | 95 | Carbapenem resistance prediction using MALDI-TOF MS and machine learning algorithms, specifically focusing on differentiating carbapenem-resistant Klebsiella pneumoniae (CRKP) from carbapenem-susceptible Klebsiella pneumoniae (CSKP) isolates. | The source of the isolates is from clinical specimens including pus/wound, sputum, urine, blood, and body fluids like ascitic fluid, bile, pleural fluid, and catheter tips. | Not mentioned (the paper does not provide information on the number of different hospitals from which the isolates were sourced) | - Pus/wound: 7 (7.4%) - Sputum: 5 (5.3%) - Urine: 20 (21.1%) - Blood: 54 (56.8%) - Body fluids (ascitic fluid, bile, pleural fluid): 5 (5.3%) - Other specimens (catheter tips): 4 (4.2%) | Vitek MS system (bioMe ´rieux SA) | 1. Peak matching: Using Mass-Up software to unify peak values among samples with a tolerance of 300 ppm. 2. Feature selection: Using the Student's t-test to compare the distribution of each m/z ratio, generate p-values, and rank them to select the most significant peaks for machine learning analysis. | - Accuracy: 81%-97% - Sensitivity: 76%-96% - Specificity: 82%-100% - Best performance with RF: Accuracy = 97%, Sensitivity = 93%, Specificity = 100% when 80 peaks were selected | Random Forest (RF) algorithm with 80 peaks selected as input features, achieving 97% overall classification accuracy | Not mentioned (the paper uses internal validation methods such as L1O cross-validation and does not mention an independent validation dataset) | Improved time to detection is implied through the use of MALDI-TOF MS and the direct colony method, which are faster and more streamlined than traditional methods. However, specific timeframes or comparisons are not provided. | 7705.009 Da
De Carolis et al., 2017 | A rapid diagnostic workflow for cefotaximeresistant Escherichia coli and Klebsiella pneumoniae detection from blood cultures by MALDI-TOF mass spectrometry | Not mentioned | 1. Centrifugation of positive blood culture fluid to obtain a bacterial pellet. 2. Resuspension of the pellet in HPLC-grade water to achieve a turbidity of 2 McFarland. 3. Centrifugation to separate bacterial cells from the supernatant. 4. Incubation with cefotaxime for β-lactam ring hydrolysis. 5. Centrifugation after incubation. 6. Spotting of the supernatant onto a target plate. 7. Application of HCCA matrix solution. 8. Acquisition of spectra in positive linear mode with a defined mass range. | - The D-MSBL assay only determines susceptibility/resistance to CTX and does not indicate ESBL alone. - It does not guide therapy for carbapenem resistance, which is a limitation in settings with high prevalence of carbapenemase-producing bacteria. - The assay is more valuable for E. coli bacteremia than for K. pneumoniae due to carbapenem resistance prevalence. - Further studies are needed to integrate the assay into laboratory workflows for real-time diagnosis and therapy initiation. | Not mentioned (the paper does not specify an external validation dataset distinct from the test set used in the study) |  | Bruker Microflex LT mass spectrometer | Not reported | - Study type: Prospective - Setting and timeframe: Conducted at Università Cattolica del Sacro Cuore, Fondazione Policlinico Universitario Agostino Gemelli in Rome, Italy, between April and December 2015 - Data collection approach: MALDI-TOF MS identification followed by D-MSBL assay for resistance detection - Comparison groups or control methods: Compared with phenotypic and genotypic testing methods - Validation strategy: Tested on reference strains and compared with established methods | - Equipment/platform used: Microflex LT mass spectrometer (Bruker Daltonics) - Sample preparation methods: Inoculation into VACUETTE Z Serum Sep Clot Activator tube, centrifugation, transfer to Eppendorf tube with HPLC-grade water to reach turbidity of 2 McFarland - Spectral data preprocessing steps: Use of Bruker's Biotyper software and library - Mass range analyzed: 100 to 1000 Da, with focus on 300 to 600 Da for CTX hydrolysis - Number of technical replicates per isolate: Triplicate resuspension in CTX solution - Quality control measures: Reference strains as controls, control without bacteria for spontaneous hydrolysis - Spectral data processing software/methods: ClinProTools for CTX hydrolysis peak analysis | - Specific algorithms used: Novel algorithm for bacterial organism classification based on mass spectrometry data. - Training/validation/test set split ratios and sizes: Not mentioned. - Feature selection or extraction methods: Mass peaks (370.5, 414.5, 396.5, and 456.5 Da) used for classification. - Hyperparameter optimization approach: Not mentioned. - Cross-validation strategy: Not mentioned. - Any ensemble methods or model combinations: Not mentioned. - Software/programming languages used: ClinProTools (Bruker Daltonics) version 3.0. | - Specific antibiotics tested: Cefotaxime (CTX) - Resistance mechanisms: Extended-spectrum β-lactamases (ESBLs), specifically TEM, SHV, and CTX-M family genes - Reference standard: EUCAST breakpoints document version 6.0 2016 - Resistance breakpoints or criteria: EUCAST breakpoints - Multi-drug resistance patterns: Assessed through genotypic characterization of ESBL and other resistance genes | - Total sample size: 35 - Breakdown by resistance status: 19 CTX-resistant, 16 CTX-susceptible - Source of isolates: Blood cultures from patients at Università Cattolica del Sacro Cuore, Fondazione Policlinico Universitario Agostino Gemelli, Rome, Italy - Isolation timeframe: April to December 2015 - Patient population: Patients with bacteremia - Inclusion criteria: Positive blood cultures with E. coli or K. pneumoniae - Distribution of resistant vs. susceptible isolates: 19 resistant, 16 susceptible | - Sensitivity: 86.8% - Specificity: 98.2% - Positive Predictive Value (PPV): 97.0% - Negative Predictive Value (NPV): 91.5% - Agreement with genotypic method: 94.8% for E. coli, 92.5% for K. pneumoniae | - Time to result: Less than 2 hours, 24 to 48 hours faster than conventional methods. - Cost considerations: Cost-effective compared to PCR-based methods. - Ease of implementation: Can be integrated into existing diagnostic laboratory workflows. - Integration with existing workflows: Designed to fit into current lab workflows. - Real-time or rapid result capabilities: Provides rapid results for timely antibiotic therapy decisions. - Potential impact on antimicrobial stewardship: Aids in choosing appropriate antibiotics by rapidly identifying resistant organisms. | - Technical limitations: The D-MSBL assay only determines susceptibility/resistance to cefotaxime and does not indicate ESBL presence. - Generalizability concerns: The assay is more applicable in settings with low prevalence of carbapenemase-producing bacteria. - Performance issues: Less beneficial for K. pneumoniae due to high carbapenem resistance. - Need for larger validation studies: Further studies are needed to confirm integration into clinical practice. - Implementation barriers: Limited by inability to detect resistance to other antibiotics. | 93 | Cefotaxime (CTX), a third-generation cephalosporin antibiotic | The isolates were sourced from blood cultures collected at the Università Cattolica del Sacro Cuore, Fondazione Policlinico Universitario Agostino Gemelli in Rome, Italy. | 1 (the study was conducted at a single hospital: Università Cattolica del Sacro Cuore, Fondazione Policlinico Universitario Agostino Gemelli) | Blood | Microflex LT mass spectrometer (Bruker Daltonics, Bremen, Germany) | - Co-incubation of blood culture samples with cefotaxime. - Analysis of mass spectra to identify peaks associated with resistance or susceptibility. - Use of ClinProTools software for peak analysis. - Application of specific mass ranges (300 to 600 Da). - Algorithm-based classification using statistical equations. | - Sensitivity: 86.8% - Specificity: 98.2% - Positive Predictive Value (PPV): 97.0% - Negative Predictive Value (NPV): 91.5% - Agreement with genotypic method: - E. coli: 94.8% - K. pneumoniae: 92.5% | The best model developed in the study is the Direct Mass Spectrometry assay for Beta-Lactamase (D-MSBL) for identifying cefotaxime-resistant Escherichia coli and Klebsiella pneumoniae from positive blood cultures. | Not mentioned (the paper does not specify an independent validation dataset separate from the 93 clinical isolates used in the study) | 24 to 48 hours earlier than conventional methods | - Resistance peaks: 370.5 Da, 414.5 Da - Susceptibility peaks: 396.5 Da, 456.5 Da
Giordano and Barnini, 2018 | Rapid detection of colistin-resistant Klebsiella pneumoniae using MALDI-TOF MS peak-based assay | Intensity normalisation was performed by normalizing each spectrum to its own Total Ion Count (TIC), applied per spectrum individually. | 1. Protein extraction protocol: - Create a three McFarland suspension from pure cultures. - Centrifuge at 16.200 x g for 2 minutes; discharge supernatant. - Suspend pellet in 75% ethanol; centrifuge at 16.200 x g for 2 minutes; discharge supernatant. - Add 70% formic acid; pipette until dissolved. - Add acetonitrile; vortex at maximum speed for 1 minute. - Centrifuge at 16.200 x g for 2 minutes to remove debris. - Dispense supernatant on metallic plate; dry at room temperature. - Cover with α-cyano-4-hydroxycinnamic acid (HCCA) matrix solution; dry again. 2. Spectra acquisition: - Acquire spectra in positive linear mode. - Laser frequency: 60 Hz. - Acquisition range: 1.960 to 20.000 Da. 3. Post-acquisition processing: - Baseline subtraction and smoothing using FlexAnalysis v3.0 software. - Normalize spectra to Total Ion Count (TIC). - No automatic recalibration of spectra. | - The study identifies limitations in the current methodologies for colistin susceptibility testing. - Identifying reliable peak biomarkers for MALDI-TOF MS is a complication. - The negative predictive value of the assay is 73%, which may produce false positive results. - Increasing the number of isolates could improve the model's accuracy and specificity. - The study is preliminary, requiring further validation. | The size of the external validation dataset is 89 strains, from which 712 spectra were collected, with 158 eliminated due to low identification scores. This results in a usable dataset of 554 mass spectra for external validation. | 89 | Bruker Daltonik GmbH, Bremen, Germany | Not reported | - Study type: Retrospective - Setting and timeframe: Pisa University Hospital, 2015-2017 - Data collection approach: Collection of 139 K. pneumoniae isolates and use of MALDI-TOF MS - Comparison groups or control methods: Training set (50 strains) and test set (89 strains) - Validation strategy: Cross-validation and independent test set | - Equipment/platform used: Bruker Daltonik GmbH, Bremen, Germany - Sample preparation methods: Three McFarland suspension, centrifugation, ethanol treatment, formic acid and acetonitrile addition, drying on metallic plate with HCCA matrix solution - Spectral data preprocessing steps: Baseline subtraction and smoothing using FlexAnalysis v3.0, normalization to Total Ion Count (TIC) using ClinProtTools v3.0 - Mass range analyzed: 1.960 Da to 20.000 Da - Number of technical replicates per isolate: Eight - Quality control measures: Main Spectrum Profile (MSP), Genetic Algorithm (GA) - Spectral data processing software/methods: FlexAnalysis v3.0, MALDI Biotyper v3.0, ClinProtTools v3.0, GA, SNN, QC | - Specific algorithms used: Genetic Algorithm (GA), Supervised Neural Network (SNN), Quick Classifier (QC) - Training/validation/test set split ratios and sizes: Training set - 50 strains with 400 spectra; Test set - 89 strains with 712 spectra - Feature selection or extraction methods: Manually selected mass peaks for GA algorithm - Hyperparameter optimization approach: Number of neighbors set to 5; Mutation rate set to 0.2; Crossover rate set to 0.5 - Cross-validation strategy: Recognition capability and cross-validation percentages mentioned - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: ClinProtTools v3.0 | - Specific antibiotics tested: colistin, amikacin, amoxicillin/clavulanate, ampicillin/sulbactam, cefepime, cefotaxime, ceftazidime, ciprofloxacin, ertapenem, fosfomycin, gentamycin, imipenem, levofloxacin, meropenem, nitrofurantoin, piperacillin-tazobactam, tigecycline, and trimethoprim/sulfamethoxazole - Resistance mechanisms: Not specified - Reference standard used for resistance determination: Broth microdilution (BMD) assay - Resistance breakpoints or criteria applied: EUCAST guidelines (v5.0) - Multi-drug resistance patterns: Assessed for multiple antibiotics, but primary focus on colistin resistance | - Total sample size: 139 K. pneumoniae isolates - Source of isolates: Patients hospitalized at Pisa University Hospital (Azienda Ospedaliero-Universitaria Pisana, Italy) - Isolation timeframe: 2015-2017 - Patient population: Patients from Pisa University Hospital - Inclusion/exclusion criteria: Not mentioned - Distribution of resistant vs. susceptible isolates: Training set - 25 colR-KPn, 25 colS-KPn; Test set - 40 colR-KPn, 49 colS-KPn | - Recognition capability and cross-validation for SNN: 97.3% and 94.6% - Recognition capability and cross-validation for QC: 91.3% and 90.4% - Recognition capability and cross-validation for GA: 99.8% and 97.8% - External validation results: - SNN: colR-KPn 19%, colS-KPn 71% - QC: colR-KPn 79%, colS-KPn 63% - GA: colR-KPn 4%, colS-KPn 86% - AUC values: 0.865 and 0.862 for peaks at 4507,28 Da and 5142,82 Da - Sensitivity and specificity for peak couple 4507,28/5142,84 Da: 78% and 89% - Positive predictive value: 91% - Negative predictive value: 73% | - Time to result: Approximately 90 minutes for sample analysis. - Cost considerations: Annual saving of 87.8% in reagent costs compared to traditional methods. - Ease of implementation: Described as easy to use and accurate. - Integration with existing workflows: Can be used for both bacterial identification and colistin resistance detection. - Real-time or rapid result capabilities: Provides rapid preliminary results on colistin resistance profiles. - Potential impact on antimicrobial stewardship: Enables rapid detection and classification of resistant isolates, informing treatment decisions. | - Technical limitations: Lack of defined methodology for colistin susceptibility testing; need for optimal sample preparation; requirement for high-quality, updated databases. - Sample size limitations: Need for larger validation studies to improve accuracy and specificity. - Generalizability concerns: More clonal lineages need to be tested to enhance robustness across different epidemiological settings. - Performance issues: Negative predictive value of 73% may produce false positives. - Implementation barriers: Initial costs of equipment may be prohibitive for some laboratories. - Areas requiring further research: Understanding lipid A modifications and their implications for colistin resistance. | 139 | Colistin resistance prediction using MALDI-TOF MS technology for Klebsiella pneumoniae isolates. | Azienda Ospedaliero-Universitaria Pisana, Italy, from patients hospitalized during 2015-2017 | 1 (Azienda Ospedaliero-Universitaria Pisana, Italy) | Not mentioned (the paper does not specify the exact source of the specimens) | Bruker Daltonik GmbH, Bremen, Germany | 1. Creation of a training set from 50 strains with 400 spectra. 2. Creation of a new database entry called Main Spectrum Profile (MSP). 3. Matching MSP against existing entries in the Taxonomy Tree Project. 4. Analysis of spectra using ClinProtTools v3.0: - Baseline subtraction and smoothing. - Normalization to Total Ion Count (TIC). 5. Use of the Genetic Algorithm (GA) with: - Maximum of 15 peaks. - 1000 iterations. - Specific settings for neighbors and mutation rates. 6. Peak selection: - Manual checking. - Evaluation using ROC curves and AUC values. | - SNN: Recognition capability = 97.3%, Cross-validation = 94.6% - QC: Recognition capability = 91.3%, Cross-validation = 90.4% - GA: Recognition capability = 99.8%, Cross-validation = 97.8% - Peak 4507.28: Recognition capability = 80.5%, Cross-validation = 74.1% - Peak 5142.82: Recognition capability = 83.3%, Cross-validation = 74.5% - Peak couple 4507.28/5142.82: Recognition capability = 91.8%, Cross-validation = 87.6% - Proportion of correctly classified colR-KPn = 91%, colS-KPn = 73% | The best model developed is based on two manually selected mass peaks: 4507,28 Da and 5142,82 Da, with a recognition capability of 91.8% and cross-validation of 87.6%. | - Test set: 89 strains - Spectra: 282 colR-KPn, 272 colS-KPn - Correct classification: 91% colR-KPn, 73% colS-KPn | 90 minutes | 4507.28 Da and 5142.82 Da
Jian et al., 2024b | Artificial Intelligence-Clinical Decision Support System in Infectious Disease Control: Combatting Multidrug-Resistant Klebsiella pneumoniae with Machine Learning | Log transformation was used for intensity normalization, applied per spectrum. | - Extraction of mass-to-charge (m/z) ratio segments and intensity values from MALDI-TOF MS results using MSExperiment and MzMLFile classes. - Conversion of data into floating-point numbers with Pandas DataFrame. - Segmentation of m/z ranges into 1-unit intervals with a ± 2 error range. - Extraction of peak intensity values for antibiotic-resistant and non-resistant strains. - Normalization of data using log transformation. - Comparative analysis to distinguish significant differences in peak values. - Merging of overlapping segments to ensure a coherent feature set. - Feature importance analysis to rank features and exclude those with low importance values. | - The use of data from a single hospital and four secondary hospitals may limit the generalizability of the findings across different geographical areas and healthcare settings. - The study's accuracy is contingent upon the quality of the underlying data, including MALDI-TOF MS profiles and antibiotic susceptibility tests. - There is an inherent uncertainty in machine learning models, which could affect their predictive accuracy, particularly for rare or novel resistance mechanisms. - The model's performance may be less reliable in detecting new or uncommon resistance patterns that were not well represented in the training data. - The focus on carbapenem and colistin resistance addresses only a segment of the broader antibiotic resistance issue. - The successful integration of AI-CDSS into clinical workflows presents logistical and acceptance challenges. | Not mentioned (the paper does not provide a specific size for the external validation dataset) |  | bioMérieux | Not reported | - Study type: Retrospective - Setting and timeframe: Tri-Service General Hospital and four secondary hospitals, January to December 2022 - Data collection approach: Collection of culture specimens from patients suspected of bacterial infections, analyzed using MALDI-TOF MS and VITEK-2 systems - Comparison groups or control methods: Not mentioned - Validation strategy: Samples from January to September 2022 for training, October to December 2022 for validation; fine-tuning using GridSearchCV with 5-fold cross-validation | - Equipment/platform used: bioMérieux, France - Sample preparation methods: Culturing on non-selective media, transfer to selective media to isolate colonies - Spectral data preprocessing steps: Extracting m/z ratio segments and intensity values using pyopenms library, segmenting m/z ranges into 1-unit intervals with ± 2 error range, extracting peak intensity values, log transformation - Mass range analyzed: Not explicitly mentioned - Number of technical replicates per isolate: Not specified - Quality control measures: Using recommended strains (E. coli ATCC 25922, E. coli ATCC 35218, P. aeruginosa ATCC 27853, K. pneumoniae ssp. pneumoniae ATCC 700603), interpreting susceptibility results according to CLSI breakpoints - Spectral data processing software/methods: pyopenms library, Pandas DataFrame | - Specific algorithms used: Random Forest Classifier, Logistic Regression, Linear Discriminant Analysis, Gradient Boosting Classifier, AdaBoost Classifier, Extreme Gradient Boosting, Light Gradient Boosting Machine, Support Vector Machine - Training/validation/test set split ratios and sizes: Samples from January to September 2022 for training, October to December 2022 for validation - Feature selection or extraction methods: Segmentation of m/z ranges into 1-unit intervals, extraction of peak intensity values - Hyperparameter optimization approach: GridSearchCV with 5-fold cross-validation - Cross-validation strategy: 5-fold cross-validation - Any ensemble methods or model combinations: Not explicitly mentioned - Software/programming languages used: Python version 3.9.18 | - Specific antibiotics tested: colistin, doripenem (DOR), imipenem (IPM) - Resistance mechanisms: carbapenem-resistant Klebsiella pneumoniae (CRKP), colistin-resistant Klebsiella pneumoniae (CoRKP) - Reference standard used for resistance determination: Clinical and Laboratory Standards Institute (CLSI) breakpoints - Resistance breakpoints or criteria applied: CLSI breakpoints with specific susceptibility ranges for each antibiotic - Multi-drug resistance patterns: Focus on both carbapenem and colistin resistance | - Total sample size: 4307 K. pneumoniae isolates - Source of isolates: Clinical specimens (blood, urine, throat swabs, sputum, body fluids, wound swabs, catheters) from patients suspected of bacterial infections - Geographic origin: Tri-Service General Hospital and four secondary hospitals - Isolation timeframe: January to December 2022 - Patient population: Patients suspected of bacterial infections - Inclusion/exclusion criteria: Inclusion - use of MALDI-TOF MS and Vitek-2 systems for AST; Exclusion - specimens lacking MALDI-TOF MS data - Distribution of resistant vs. susceptible isolates: Carbapenem-resistant (1207), carbapenem-susceptible (2803); Colistin-resistant (416), colistin-susceptible (1379) | - Random Forest Classifier (RFC): - CRKP: AUC = 0.96, Accuracy = 0.89, Sensitivity = 0.88, Specificity = 0.90, PPV = 0.89, NPV = 0.89, F1-score = 0.88 - CoRKP: AUC = 0.98, Accuracy = 0.92, Sensitivity = 0.94, Specificity = 0.91, PPV = 0.91, NPV = 0.93, F1-score = 0.93 - XGBoost (XGB): - CRKP: AUC = 0.95, Accuracy = 0.87, Sensitivity = 0.89, Specificity = 0.86, PPV = 0.85, NPV = 0.90, F1-score = 0.87 - CoRKP: AUC = 0.96, Accuracy = 0.93, Sensitivity = 0.95, Specificity = 0.90, PPV = 0.91, NPV = 0.95, F1-score = 0.93 - Gradient Boosting Classifier (GBC): - CRKP: AUC = 0.95, Accuracy = 0.89, Sensitivity = 0.90, Specificity = 0.88, PPV = 0.87, NPV = 0.90, F1-score = 0.88 - CoRKP: AUC = 0.96, Accuracy = 0.91, Sensitivity = 0.93, Specificity = 0.90, PPV = 0.90, NPV = 0.93, F1-score = 0.92 | - Time to result: The AI-CDSS can expedite the detection of KP resistance by approximately 1 day compared to traditional methods, which require at least 4 days. - Cost considerations: Not mentioned in the paper. - Ease of implementation: Not explicitly mentioned, but the system is designed to streamline the diagnostic workflow. - Integration with existing workflows: The AI-CDSS is intended to enhance clinical decision-making and improve treatment outcomes by integrating with existing workflows. - Real-time or rapid result capabilities: The system can significantly reduce the identification time of resistance strains from a day to minutes. - Potential impact on antimicrobial stewardship: The AI-CDSS is seen as a vital step forward in optimizing clinical outcomes and advancing antibiotic stewardship by facilitating targeted antibiotic therapy. | - Technical limitations: Dependence on data quality, uncertainty in machine learning models, particularly for rare or novel resistance mechanisms. - Sample size limitations: Data from a single hospital and four secondary hospitals may limit generalizability. - Generalizability concerns: Limited geographical and institutional scope. - Performance issues: Less reliable for new or uncommon resistance patterns. - Need for larger validation studies: Expansion needed to cover a wider range of pathogens and resistance mechanisms. - Implementation barriers: Logistical and acceptance challenges in integrating AI-CDSS into clinical workflows. - Areas requiring further research: Enhancing data collection methods, extending the system's scope, addressing practical challenges of implementation. | 4307 | The antibiotic resistance prediction in this study focused on carbapenem-resistant Klebsiella pneumoniae (CRKP) and colistin-resistant Klebsiella pneumoniae (CoRKP) using machine learning models like Random Forest Classifier, XGBoost, and Gradient Boosting Classifier. | Tri-Service General Hospital and four secondary hospitals (three in the northern region and one on an island) | 5 | - Blood - Urine - Throat swabs - Sputum - Body fluids - Wound swabs - Catheters | bioMérieux, France | 1. Segmentation of m/z ranges into 1-unit intervals with a ± 2 error range. 2. Extraction of peak intensity values for antibiotic-resistant and non-resistant strains. 3. Log transformation for data normalization. 4. Merging of overlapping segments to ensure a coherent feature set. 5. Feature importance analysis to exclude background noise. 6. Resampling to address data imbalance. | - Random Forest Classifier (RFC): - CRKP: AUC = 0.96, Accuracy = 0.89, Sensitivity = 0.88, Specificity = 0.90, PPV = 0.89, NPV = 0.89, F1-score = 0.88 - CoRKP: AUC = 0.98, Accuracy = 0.92, Sensitivity = 0.94, Specificity = 0.91, PPV = 0.91, NPV = 0.93, F1-score = 0.93 - XGBoost (XGB): - CRKP: AUC = 0.95, Accuracy = 0.87, Sensitivity = 0.89, Specificity = 0.86, PPV = 0.85, NPV = 0.90, F1-score = 0.87 - CoRKP: AUC = 0.96, Accuracy = 0.93, Sensitivity = 0.95, Specificity = 0.90, PPV = 0.91, NPV = 0.95, F1-score = 0.93 - Gradient Boosting Classifier (GBC): - CRKP: AUC = 0.95, Accuracy = 0.89, Sensitivity = 0.90, Specificity = 0.88, PPV = 0.87, NPV = 0.90, F1-score = 0.88 - CoRKP: AUC = 0.96, Accuracy = 0.91, Sensitivity = 0.93, Specificity = 0.90, PPV = 0.90, NPV = 0.93, F1-score = 0.92 | Random Forest Classifier (RFC) | Yes, external validation was performed using samples collected from October to December 2022. | 1 day | - CRKP: m/z segments around 2764, 3150, 4605, 5822, and 6152 - CoRKP: m/z segments of 4154, 4341, 4571, 5252, 6596, 7158, 8993, and 9252
Lopez-Cortez et al., 2024 | MSDeepAMR: antimicrobial resistance prediction based on deep neural networks and transfer learning | Not mentioned | The pre-processing scheme utilized involves binning the raw MALDI-TOF MS spectra with a bin size of 3 Da in the range of 2,000 to 20,000 Da, resulting in a fixed-length vector of 6,000 features. Typical preprocessing steps like smoothing, baseline correction, and peak picking were significantly reduced or not performed. | - Poor reproducibility - Small sample size - Lack of external validation - Challenges with transfer learning due to differences in mass spectrometry systems - Limited cross-device compatibility (specifically with MALDI-TOF from Bruker) | - DRIAMS-B: K. pneumoniae-Ceftriaxone (134 susceptible, 18 resistant) | 152 | Bruker Daltonics MALDI Biotyper | Not reported | The study is a retrospective analysis using the DRIAMS database. It involves training and testing a deep learning model (MSDeepAMR) on raw mass spectra data to predict antibiotic resistance. The setting includes data from four clinical laboratories (DRIAMS-A, B, C, D). The study uses transfer learning and employs 10-fold cross-validation and independent test sets for validation. It compares models trained with local data only, without transfer learning, and with transfer learning. | - Equipment/platform used: MALDI Biotyper by Bruker Daltonics - Sample preparation methods: Incubation with antibiotics, centrifugation, analysis of supernatant - Spectral data preprocessing steps: Minimal preprocessing, binning with a size of 3 Da - Mass range analyzed: 2-20 kDa - Number of technical replicates per isolate: Not mentioned - Quality control measures: Not mentioned - Spectral data processing software/methods: Deep learning models (MSDeepAMR) trained on raw data | - Specific algorithms used: Convolutional Neural Network (CNN) - Training/validation/test set split ratios and sizes: Not explicitly mentioned - Feature selection or extraction methods: Not explicitly mentioned - Hyperparameter optimization approach: Grid search for E. coli-ceftriaxone pair - Cross-validation strategy: 10-fold cross-validation - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: Not explicitly mentioned, but available on GitHub | - Specific antibiotics tested: ciprofloxacin, ceftriaxone, cefepime, piperacillin-tazobactam, tobramycin, meropenem, fusidic acid, oxacillin - Resistance mechanisms: Not specified - Reference standard: DRIAMS database - Resistance breakpoints or criteria: Not specified - Multi-drug resistance patterns: Not explicitly discussed | - Total sample size for K. pneumoniae: 2,800 - Source of isolates: Clinical laboratories (specific geographic origin not mentioned) - Isolation timeframe: Not mentioned - Patient population: Not specified - Inclusion/exclusion criteria: Not mentioned - Distribution of resistant vs. susceptible isolates: Detailed in tables at the end of the paper | - AUROC values: - E. coli-Ciprofloxacin: 0.85 - E. coli-Ceftriaxone: 0.87 - E. coli-Cefepime: 0.88 - K. pneumoniae-Ceftriaxone: 0.82 - K. pneumoniae-Cefepime: 0.83 - K. pneumoniae-Meropenem: 0.83 - S. aureus-Oxacillin: 0.93 - AUPRC values: - E. coli-Ciprofloxacin: 0.75 - E. coli-Ceftriaxone: 0.79 - E. coli-Cefepime: 0.70 - K. pneumoniae-Ceftriaxone: 0.68 - Comparative performance: MSDeepAMR shows a 10% improvement over state-of-the-art models in AUROC. | - Time to result: MSDeepAMR significantly reduces the time for AST compared to conventional methods, which take up to 72 hours. - Cost considerations: Potential decrease in public health costs due to reduced time and resources needed for AST. - Ease of implementation: High sensitivity and ability to use on various clinical samples indicate ease of implementation in clinical labs. - Integration with existing workflows: Can be integrated with MALDI-TOF devices for real-time AMR detection without manual preprocessing. - Real-time or rapid result capabilities: Enables on-the-fly AMR detection. - Potential impact on antimicrobial stewardship: Improved accuracy and rapid detection can guide appropriate antibiotic use. | - Technical limitations: Use of raw MS data reduces preprocessing but may introduce complexity in interpretation. - Sample size limitations: Small sample size and lack of external validation. - Generalizability concerns: Limited to specific bacteria and antibiotics, and reliance on MALDI-TOF from Bruker. - Performance issues: Challenges with interpreting spectra of highly related strains. - Need for larger validation studies: Use of 10-fold cross-validation to avoid overfitting. - Implementation barriers: Limited to specific equipment (MALDI-TOF from Bruker). - Areas requiring further research: Optimizing methodologies, adapting to other MALDI-TOF devices, and creating public databases. | - Escherichia coli: 5,000 - Klebsiella pneumoniae: 2,800 - Staphylococcus aureus: 3,800 | - E. coli: ciprofloxacin, ceftriaxone, cefepime, piperacillin-T, tobramycin - K. pneumoniae: ciprofloxacin, ceftriaxone, cefepime, meropenem, tobramycin - S. aureus: ciprofloxacin, fusidic acid, oxacillin | The source of the isolates is from four different clinical laboratories in Switzerland, as indicated by the DRIAMS database sub-collections: DRIAMS-A, DRIAMS-B, DRIAMS-C, and DRIAMS-D. | 4 | Not mentioned (the paper does not specify the specimen source isolation beyond mentioning blood culture bottles) | MALDI Biotyper from Bruker Daltonics | Binning with a bin size of 3 Da in the range of 2,000 to 20,000 Da to create a fixed-length vector of 6,000 features. | - AUROC (Area Under the Receiver Operating Characteristic Curve) - AUPRC (Area Under the Precision-Recall Curve) - Balanced accuracy - Tables 4, 5, 6 provide AUROC and AUPRC values for transfer learning experiments. - Tables referenced in the paper provide detailed performance metrics for different models and scenarios. | E. coli-Ceftriaxone, K. pneumoniae-Ceftriaxone, and S. aureus-Oxacillin | - DRIAMS-B - DRIAMS-C - DRIAMS-D | The MSDeepAMR model can significantly reduce the time required for antibiotic sensitivity testing (AST), which typically takes up to 72 hours. The exact improved time to detection is not specified in the paper. | - E. coli-Ceftriaxone: Peaks in the range 6,800-6,900Da - K. pneumoniae-Ceftriaxone: Peaks in the range 2,000-3,000 Da - S. aureus-Oxacillin: Peaks 2,414 Da and 4,517 Da
Gato et al., 2023 | Direct Detection of Carbapenemase-Producing Klebsiella pneumoniae by MALDI-TOF Analysis of Full Spectra Applying Machine Learning | Intensity normalisation was performed using the total ion current (TIC) method, where each intensity value was divided by the area under the spectrum value. Normalisation was applied per spectrum. | 1. Noise reduction using Savitzky-Golay filter (window length: 11, polynomial order: 3) 2. Baseline subtraction using Top-Hat filter (factor: 0.02) 3. Alignment and merging of spectra to create an average spectrum per isolate 4. Alignment of average spectra with each other (shift medium; linear tolerance: 900 ppm; constant tolerance: 3 Da) 5. Normalization of peak matrices by total ion current (TIC) method | - Impact of culture medium on predictive performance: The model was trained on isolates cultured on blood agar, and using different media can affect accuracy. - Limited generalizability: The model may not fit all biological and epidemiological variations of bacteria across different regions and over time. - Reliance on pattern matching rather than specific biomarkers: This can be less reliable due to variability in protein expression. - Limitation in detecting proteins beyond the effective mass range of MALDI-TOF MS: Many resistance-related proteins are beyond this range. - Need for individual validation and adaptation: The model requires preliminary validation with known isolates and regular reevaluation to adapt to regional characteristics. | 236 | 236 | Bruker Microflex LT/SH SMART mass spectrometer | Blood agar | - Study type: Retrospective - Setting and timeframe: Clinical setting, within 24 hours of sample collection - Data collection approach: Database of mass spectra profiles from clinical isolates - Comparison groups or control methods: Training set and validation set - Validation strategy: 10-fold cross-validation for CPK differentiation, 7-fold cross-validation for carbapenemase type differentiation, AUROC and AUPRC as performance metrics | - Equipment/platform used: Microflex LT/SH SMART mass spectrometer with FlexControl software 3.4 - Sample preparation methods: "in-target" protein extraction - Spectral data preprocessing steps: Noise reduction with Savitzky-Golay filter, baseline subtraction with Top-Hat filter - Mass range analyzed: 2 to 20 kDa - Number of technical replicates per isolate: Nine replicates (three spots with three spectra each) for training set, one replicate (one spot with one spectrum) for validation set - Quality control measures: External calibration using bacterial test standard (BTS), score above 2.0 for species confirmation - Spectral data processing software/methods: Clover MS Data Analysis Software for preprocessing, MALDI Biotyper Compass for species confirmation | - Specific algorithms used: Random Forest (RF), Partial Least Squares Discriminant Analysis (PLSDA), Support Vector Machine (SVM), k-Nearest Neighbor (KNN) - Training/validation/test set split ratios and sizes: Training set - 479 isolates (246 CPK, 233 NCPK); Validation set - 236 isolates (78 CPK, 158 NCPK) - Feature selection or extraction methods: Full spectrum of mass peaks from 2 to 20 kDa; Preprocessing with noise reduction and baseline subtraction - Hyperparameter optimization approach: Mentioned but not detailed - Cross-validation strategy: 10-fold for CPK differentiation; 7-fold for carbapenemase type differentiation - Any ensemble methods or model combinations: Not explicitly mentioned - Software/programming languages used: Clover MS Data Analysis Software for preprocessing; Other software likely used for machine learning tasks but not explicitly mentioned | - Specific antibiotics tested: Meropenem, Ertapenem - Resistance mechanisms: Carbapenemases (OXA-48, KPC, potential inclusion of VIM and NDM) - Reference standard: European Committee on Antimicrobial Susceptibility Testing (EUCAST) - Resistance breakpoints: MIC > 0.125 mg/L for meropenem or ertapenem - Multi-drug resistance patterns: Association of carbapenemase types with specific sequence types (STs) of K. pneumoniae | - Total sample size: 715 unduplicated clinical isolates - Breakdown by resistance status: 324 CPK, 391 NCPK - Source of isolates: Clinical specimens (biopsy specimen, catheter, bile, peritoneal fluid, abscesses, respiratory samples, wounds, blood cultures, urine samples, samples of unknown origin) from various hospitals in Spain - Isolation timeframe: Two-month period in 2018 - Patient population: Clinical specimens from patients with infections - Inclusion criteria: Part of a nationwide survey or authors' own collection - Exclusion criteria: Isolates not meeting EUCAST criteria for carbapenemase screening - Distribution of resistant vs. susceptible isolates: 324 CPK, 391 NCPK | - Accuracy for CPK prediction: 97.83% - Accuracy for OXA-48 or KPC carriage prediction: 95.24% - RF algorithm with M LINEAR method: Accuracy = 99.79%, Sensitivity = 100%, Specificity = 99.57% - AUC values for RF: 1.00 for both AUROC and AUPRC - Comparative performance: RF and PLS show outstanding discrimination | - Time to result: Identification of CPK isolates within a few minutes, compared to 48 to 72 hours for conventional methods. - Cost considerations: Cost-effective compared to molecular techniques like whole-genome sequencing. - Ease of implementation: Easy to implement in clinical labs using existing MALDI-TOF MS equipment and automated bioinformatics processing. - Integration with existing workflows: Integrates well with existing workflows as it uses the same hands-on procedure and spectra as bacterial identification. - Real-time or rapid result capabilities: Provides rapid results, which can support clinical decision-making and antimicrobial stewardship. - Potential impact on antimicrobial stewardship: Reduces time to therapeutic de-escalation and supports clinical decision-making. | - Technical limitations: Impact of culture medium on CPK prediction; need for consistent culture conditions. - Sample size limitations: Limited validation for VIM and NDM-type carbapenemases. - Generalizability concerns: Model may not fit all biological and epidemiological variations; need for prospective validation in different centers. - Performance issues: Lower accuracy for certain carbapenemase types. - Need for larger validation studies: Individual validation with known isolates to adapt to regional characteristics. - Implementation barriers: Need for adaptation to regional characteristics. - Areas requiring further research: Direct identification from blood cultures; validation for other carbapenemase types. | 715 | Carbapenem resistance (specifically OXA-48 and KPC types) in Klebsiella pneumoniae | - 307 CPK isolates: Collected during a nationwide survey involving 15 hospitals in Spain. - 391 NCPK isolates: Collected from various hospitals in Spain. - 17 isolates: Part of the authors' laboratory collection. - Types of samples: Biopsy specimen, catheter, bile, peritoneal fluid, abscesses, respiratory samples, wounds, blood cultures, urine samples, and samples of unknown origin. | 15 | - Biopsy specimen: 1 - Catheter: 1 - Bile: 2 - Peritoneal fluid: 2 - Abscesses: 4 - Respiratory samples: 13 - Wounds: 17 - Blood cultures: 19 - Urine samples: 171 - Samples of unknown origin: 6 | Microflex LT/SH SMART mass spectrometer | 1. Noise reduction with Savitzky-Golay filter and baseline subtraction with Top-Hat filter. 2. Alignment and merging of spectra to create an average spectrum for each isolate. 3. Shifting peaks to match a reference peak list. 4. Application of two mass selection methods: M THRESHOLD and M LINEAR. 5. Normalization of peak matrices by total ion current (TIC) method. | - Accuracy for CPK prediction: 97.83% - Accuracy for OXA-48 or KPC carriage prediction: 95.24% - RF algorithm accuracy: 99.27% - AUROC and AUPRC for CPK prediction: 1.00 - AUROC for OXA-48-like producing isolates: 0.978 - AUPRC for OXA-48-like producing isolates: 0.99 - AUROC for KPC-producing isolates: 0.975 - AUPRC for KPC-producing isolates: 0.97 | Random Forest (RF) algorithm with the M LINEAR method | Yes, an external validation was performed using a separate dataset of 236 isolates (78 CPK and 158 NCPK) to evaluate the predictive performance of the machine learning algorithms. | Within a few minutes; within 24 hours of sample collection | Not mentioned (the paper does not identify a single "Top AMR peak" or biomarker; instead, it emphasizes the use of the complete proteome and multiple mass peaks for prediction)
Gato et al., 2021 | An Improved Matrix-Assisted Laser Desorption Ionization-Time of Flight Mass Spectrometry Data Analysis Pipeline for the Identification of Carbapenemase-Producing Klebsiella pneumoniae | Normalization method: Total Ion Current (TIC), Application: Applied per spectrum | 1. Noise reduction using Savitzky-Golay filter (window length of 11, polynomial order 3) 2. Baseline subtraction using Top-Hat filter (factor of 0.02) 3. Alignment of spectra in a two-step process 4. Normalization of peak matrices by total ion current (TIC) method | The study identified limitations in the ML models, particularly in the accuracy of identifying specific types of carbapenemases (82% overall, with varying success rates for different types) and the need for further introduction of KPC and NDM isolates into the training set to enhance accuracy. Additionally, further validation is necessary to ensure generalizability to broader populations. | 122 isolates (84 CPK and 38 non-CPK) | 122 | Bruker Daltonics, specifically the MALDI Biotyper and microflex LT/SH smart instrument | Blood agar | - Study type: Prospective study with two stages (reproducibility and CPK identification) - Setting: Multi-center study conducted in three hospitals (CHUAC, HGUGM, HUPM) - Timeframe: Not explicitly mentioned - Data collection approach: Collection of 162 unduplicated clinical isolates (112 CPK, 50 non-CPK) - Comparison groups: CPK vs. non-CPK isolates - Validation strategy: k-fold cross-validation (k = 10) | - Equipment/platform used: MALDI BioTyper system (Bruker Daltonik, Germany) - Sample preparation methods: In-target extraction and full protein extraction - Spectral data preprocessing steps: Noise reduction with Savitzky-Golay filter, baseline subtraction with Top-Hat filter - Mass range analyzed: 2 to 20 kDa - Number of technical replicates per isolate: Nine replicates (3 spots with 3 spectra each) - Quality control measures: External calibration using bacterial test standard, isolates of the same age - Spectral data processing software/methods: Clover MS Data Analysis Software, PLSDA and RF algorithms | - Specific algorithms used: Partial Least Squares Discriminant Analysis (PLSDA), Random Forest (RF) - Training/validation/test set split ratios and sizes: Training set of 110 isolates, final validation with 122 isolates (84 CPK, 38 non-CPK) - Feature selection or extraction methods: Not explicitly mentioned - Hyperparameter optimization approach: Not mentioned - Cross-validation strategy: k-fold cross-validation with k = 10 - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: Clover MS Data Analysis Software | - Specific antibiotics tested: meropenem, ertapenem - Resistance mechanisms: KPC, NDM, OXA-48 - Reference standard: EUCAST screening cutoff values - Resistance breakpoints: MIC ≤0.125 mg/liter for meropenem or ertapenem - Multi-drug resistance patterns: Not explicitly mentioned beyond carbapenemase production | - Total sample size: 162 K. pneumoniae isolates - Breakdown by resistance status: 112 CPK, 50 non-CPK - Source of isolates: Nationwide survey of carbapenemase-producing Enterobacterales from 15 hospitals in Spain and authors' own collection - Geographic origin: 15 different hospitals throughout Spain, with some isolates of unknown origin - Isolation timeframe: Not specified - Patient population: Not specified - Inclusion/exclusion criteria: Not explicitly mentioned - Distribution of resistant vs. susceptible isolates: 47 bla KPC, 58 bla OXA-48, 7 bla NDM | - Accuracy: 100% for CPK identification using RF-M LINEAR - Specificity: 100% for CPK identification using RF-M LINEAR - Kappa value: Mean kappa of 95% for RF-M LINEAR - Accuracy for specific carbapenemases: - OXA-48: 93% - KPC: 69% - NDM: 66% | - Time to result: Real-time direct tracking of CPK isolates is possible. - Cost considerations: Less expensive compared to genomic assays. - Ease of implementation: Fully automated processing requires no specialized skills beyond basic MALDI-TOF MS use. - Integration with existing workflows: Can be integrated into routine laboratory workflow. - Real-time or rapid result capabilities: Enables real-time identification. - Potential impact on antimicrobial stewardship: Contributes to targeted treatments and early outbreak identification. | - Technical limitations: Lower accuracy rates for KPC and NDM isolates compared to OXA-48. - Sample size limitations: Small number of NDM isolates limited significance. - Generalizability concerns: Need for further validation studies with larger collections of K. pneumoniae isolates. - Performance issues: Lower accuracy for KPC and NDM isolates. - Need for larger validation studies: Necessary before full integration into clinical workflows. - Implementation barriers: Replacement of molecular techniques with MALDI-TOF MS not yet feasible. - Areas requiring further research: Expansion to other Enterobacterales and enhancing accuracy for KPC and NDM isolates. | 162 K. pneumoniae isolates (112 CPK and 50 non-CPK) | Carbapenemase-producing Klebsiella pneumoniae (CPK) isolates; Carbapenems and beta-lactam/beta-lactamase inhibitors (e.g., imipenem/relebactam and ceftazidime/avibactam) | - 93 CPK isolates from a nationwide survey of carbapenemase-producing Enterobacterales (CPE) from 15 participating hospitals throughout Spain. - 19 CPK isolates from the authors' own collection. - 50 non-CPK isolates (source not specified in detail). | 15 | Not mentioned (the paper does not specify the specimen source isolation for the clinical isolates) | - MALDI Biotyper (Bruker Daltonics, Germany) - microflex LT/SH smart instrument - FlexControl software version 3.4 - Bacterial test standard (BTS; Bruker Daltonics, Germany) | 1. Noise reduction using the Savitzky-Golay filter. 2. Baseline subtraction using the Top-Hat filter. 3. Alignment of spectra in a two-step process. 4. Normalization of peak matrices by the total ion current (TIC) method. | - Overall accuracy for CPK identification: 100% - Specificity for CPK identification: 100% - Accuracy for OXA-48 identification: 93% - Accuracy for KPC identification: 69% - Accuracy for NDM identification: 66% | The best model developed is the RF-M LINEAR method, which uses the random forest algorithm with total ion current (TIC) normalization and achieved 100% correct identification of CPK isolates. The pipeline includes noise reduction, baseline subtraction, alignment, and TIC normalization. | - Number of isolates: 122 (84 CPK, 38 non-CPK) - Method: RF-M LINEAR algorithm - Accuracy: 100% for CPK identification, 100% specificity - Further validation needed for full integration into clinical practice | Real-time direct tracking of CPK isolates is possible using MALDI-TOF MS, with automated processing allowing for rapid identification. This suggests a significantly improved time to detection compared to traditional methods. | Not mentioned (the paper does not specify any particular "Top AMR peaks m/z/ m/z biomarker")
Lin et al., 2024 | Artificial intelligence-clinical decision support system for enhanced infectious disease management: Accelerating ceftazidime-avibactam resistance detection in Klebsiella pneumoniae | Log-transformation, applied per spectrum | 1. Extraction of m/z segments and intensity values from MALDI-TOF MS results using MSExperiment and MzMLFile classes from pyopenms. 2. Conversion of data into floating-point numbers using Pandas DataFrame. 3. Merging with clinical susceptibility outcome data (resistance labeled as '1', susceptibility as '0'). 4. Segmentation of m/z segment values into 1-unit intervals with +/-2 error range. 5. Log-transformation of peak intensity values to normalize data. 6. Evaluation and ranking of features based on random forest classifier's 'feature_importances' attribute. 7. Application of a critical threshold of 0.01 to exclude less informative features. | - Reliance on data from a limited number of hospitals may limit generalizability across diverse clinical settings. - Performance of machine learning models is contingent on the quality and diversity of the training data. - Need for future research to validate findings in multicenter studies and integrate additional data types to refine predictive capabilities. | Not mentioned (the paper does not provide information on an external validation dataset) |  | bioMérieux, France | Not reported | - Study type: Retrospective - Setting and timeframe: Tertiary hospital and four secondary hospitals, 2022-2023 - Data collection approach: Selected 675 KP specimens from 107,721 bacterial samples - Comparison groups or control methods: Training set (502 specimens) and validation set (173 specimens) - Validation strategy: GridSearchCV with 5-fold cross-validation | - Equipment/platform used: bioMérieux, France - Sample preparation methods: Culturing on non-selective media and transferring to selective media - Spectral data preprocessing steps: Extracting m/z segment and intensity values, converting to floating-point numbers using Pandas DataFrame, merging with clinical susceptibility data - Mass range analyzed: Specific m/z segments (e.g., 2760-2765, 4518-4523) - Number of technical replicates per isolate: Not specified - Quality control measures: Using recommended strains (e.g., E. coli ATCC 25922, K. pneumoniae ssp. pneumoniae ATCC 700603) - Spectral data processing software/methods: pyopenms library, Pandas DataFrame | - Specific algorithms used: Random Forest Classifier, Gradient Boosting Classifier, AdaBoost Classifier, Extreme Gradient Boosting (XGBoost), Light Gradient Boosting Machine (LGBM), logistic regression, linear discriminant analysis - Training/validation/test set split ratios and sizes: Training set = 502 specimens, Validation set = 173 specimens - Feature selection or extraction methods: Random forest classifier's "feature_importances" attribute - Hyperparameter optimization approach: GridSearchCV with 5-fold cross-validation - Cross-validation strategy: 5-fold cross-validation - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: Python with scikit-learn library | - Specific antibiotics tested: Ceftazidime-avibactam (CZA) - Resistance mechanisms: Not specified - Reference standard used for resistance determination: VITEK-2 system - Resistance breakpoints or criteria applied: Clinical breakpoints provided by the Clinical and Laboratory Standards Institute (CLSI) - Multi-drug resistance patterns: Not explicitly mentioned | - Total sample size: 675 - Source of isolates: Clinical specimens (blood, urine, throat swabs, sputum, body fluid, wound swabs, catheters) from a tertiary hospital and four secondary hospitals - Isolation timeframe: 2022 to 2023 - Patient population: Diverse population with various health issues, including critical care and services for immunocompromised patients - Inclusion/exclusion criteria: Specimens with suspected multi-drug resistance; excluded specimens without MALDI-TOF MS data - Distribution of resistant vs. susceptible isolates: 265 CZA-resistant, 410 CZA-susceptible | - Accuracy: LGBM (0.90), Gradient Boosting Classifier (0.87), Random Forest (0.87), XGBoost (0.87) - Sensitivity: Gradient Boosting Classifier (0.92), LGBM (0.88), Random Forest (0.88), XGBoost (0.89) - Specificity: LGBM (0.91), Gradient Boosting Classifier (0.84), Random Forest (0.86), XGBoost (0.86) - AUC: LGBM (0.99 training, 0.95 testing), Gradient Boosting Classifier (0.99 training, 0.95 testing), Random Forest (0.99 training, 0.95 testing), XGBoost (0.99 training, 0.95 testing) - PPV: LGBM (0.87), Gradient Boosting Classifier (0.80), Random Forest (0.81), XGBoost (0.82) - NPV: LGBM (0.92), Gradient Boosting Classifier (0.93), Random Forest (0.91), XGBoost (0.92) - F1-score: LGBM (0.88), Gradient Boosting Classifier (0.86), Random Forest (0.85), XGBoost (0.85) | - Time to result compared to conventional methods: The AI-CDSS enables clinical decisions to be made quicker than traditional methods by a day. - Cost considerations or comparisons: High initial costs for hardware and software, continuous maintenance, and necessary upgrades require rigorous cost-effectiveness studies. - Ease of implementation in clinical labs: Integration requires effective training and overcoming operational integration challenges. - Integration with existing workflows: Strategies to address barriers include transparent communication, pilot programs, and change management techniques. - Real-time or rapid result capabilities: The AI-CDSS provides rapid CZA resistance prediction. - Potential impact on antimicrobial stewardship: Rapid detection helps in choosing effective treatments and reducing reliance on broad-spectrum antibiotics. | - Generalizability concerns: Reliance on data from a tertiary hospital and four secondary hospitals may limit generalizability across diverse clinical settings. - Technical limitations: Performance of machine learning models is contingent on the quality and diversity of training data. - Implementation barriers: Financial, infrastructural, and educational challenges; need for robust computational resources, data storage, and specialized physical space; educational support for healthcare professionals. - Need for larger validation studies: Future research required to validate findings in multicenter studies and integrate additional data types. - Areas requiring further research: Integration of additional data types to refine predictive capabilities. | 675 | Ceftazidime-avibactam (CZA) resistance prediction using MALDI-TOF MS and machine learning | - Tertiary hospital in a large urban area - Four secondary hospitals (three in the northern part of the country, one on an island) - Specimens collected from patients with suspected bacterial infections | 5 (1 tertiary hospital and 4 secondary hospitals) | - Blood - Urine - Throat swabs - Sputum - Body fluid - Wound swabs - Catheters | bioMérieux, France | 1. Extraction of mass-to-charge ratio (m/z) segment and intensity values from MALDI-TOF MS results. 2. Conversion of data into floating-point numbers using Pandas DataFrame. 3. Merging of data with clinical susceptibility outcome data. 4. Segmentation of m/z segment values into intervals to group similar values. 5. Log-transformation of peak intensity values to normalize data. 6. Identification and ranking of key features based on their importance. 7. Application of a threshold to exclude less informative features. 8. Use of resampling techniques to address data imbalance. | - Light Gradient Boosting Machine (LGBM): Training AUC = 0.99, Testing AUC = 0.95, Accuracy = 0.90, Sensitivity = 0.88, Specificity = 0.91, PPV = 0.87, NPV = 0.92, F1-score = 0.88 - Gradient Boosting Classifier: Training AUC = 0.99, Testing AUC = 0.95, Accuracy = 0.87, Sensitivity = 0.92, Specificity = 0.84, PPV = 0.80, NPV = 0.93, F1-score = 0.86 - Random Forest: Training AUC = 0.99, Testing AUC = 0.95, Accuracy = 0.87, Sensitivity = 0.88, Specificity = 0.86, PPV = 0.81, NPV = 0.91, F1-score = 0.85 - XGBoost: Training AUC = 0.99, Testing AUC = 0.95, Accuracy = 0.87, Sensitivity = 0.89, Specificity = 0.86, PPV = 0.82, NPV = 0.92, F1-score = 0.85 - AdaBoost Classifier: Training AUC = 0.99, Testing AUC = 0.87, Accuracy = 0.77, Sensitivity = 0.81, Specificity = 0.74, PPV = 0.69, NPV = 0.85, F1-score = 0.74 - Logistic Regression: Training AUC = 0.91, Testing AUC = 0.84, Accuracy = 0.78, Sensitivity = 0.77, Specificity = 0.78, PPV = 0.71, NPV = 0.83, F1-score = 0.74 - Linear Discriminant Analysis: Training AUC = 0.96, Testing AUC = 0.80, Accuracy = 0.77, Sensitivity = 0.80, Specificity = 0.75, PPV = 0.69, NPV = 0.84, F1-score = 0.74 | Light Gradient Boosting Machine (LGBM) | Not mentioned (the paper does not mention the use of an independent external validation dataset) | The AI-CDSS reduces the time to detection by approximately 1 day compared to traditional methods. It can predict CZA resistance within minutes after receiving MALDI-TOF MS data, whereas traditional methods require 4 days. | - m/z segments: approximately 2024, 3095, 4518, 5416, 5497, 6300, 7124, 9593 - Significant m/z segments: 2000-5500, 7500-10000 - Highlighted m/z segments: near 2760-2765, 4518-4523
Lopez-Cortez et al., 2025 | Integrating Machine Learning with MALDI-TOF Mass Spectrometry for Rapid and Accurate Antimicrobial Resistance Detection in Clinical Pathogens | Not mentioned | - Baseline removal - Smoothing - Peak detection - Binning procedure (grouping measured mass values into discrete ranges or "bins") - Binning applied in the range of 2000 to 10,000 Da with a bin size of 5 Da | - Suboptimal performance of transfer learning due to differences in instrument resolution and characteristics. - Failure of the SMOTE technique to improve prediction accuracy, with a reduction in performance by up to 10%. - Need for standardizing spectrum acquisition and preprocessing procedures to enhance model generalizability and clinical applicability. | Not mentioned (the paper does not specify the size of an external validation dataset distinct from the test set generated by a train-test split) |  | VITEK | Columbia Agar and MacConkey Agar | - Study type: Prospective - Setting and timeframe: Regional Hospital of Talca, September 2022 to July 2023 - Data collection approach: Culturing samples on commercial media, antimicrobial susceptibility testing using disk diffusion method - Comparison groups or control methods: Comparison of different machine learning algorithms (SVM, RF, LR, CatBoost), evaluation of transfer learning - Validation strategy: 10-fold cross-validation, decision threshold calibration | - Equipment/platform used: VITEK ® MS instruments - Sample preparation methods: Cultured on commercial media (Columbia Agar and MacConkey Agar) and incubated at 37 °C for 24 h - Spectral data preprocessing steps: Baseline removal, smoothing, and peak detection by the VITEK ® MS instrument; binning from 2000 to 10,000 Da with a bin size of 5 Da - Mass range analyzed: 2000 Da to 12,000 Da - Number of technical replicates per isolate: Not mentioned - Any quality control measures: Not mentioned - Spectral data processing software/methods: Not mentioned | - Specific algorithms used: Support Vector Machines (SVM), Random Forest (RF), Logistic Regression (LR), CatBoost - Training/validation/test set split ratios and sizes: Not explicitly mentioned - Feature selection or extraction methods: Recursive Feature Elimination with Cross-Validation (RFECV) - Hyperparameter optimization approach: Bayesian search using 'Skopt' library - Cross-validation strategy: 10-fold stratified cross-validation - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: Python with scikit-learn and 'Skopt' libraries | - Specific antibiotics tested: oxacillin for S. aureus, ciprofloxacin for E. coli and K. pneumoniae - Resistance mechanisms: Not specified - Reference standard: Disk diffusion method (Kirby and Bauer) on Muller Hinton II plates, evaluated according to CLSI M100 ED33:2023 guidelines - Resistance breakpoints or criteria: Not specified beyond CLSI guidelines - Multi-drug resistance patterns: Not assessed | - Total sample size: 187 - Breakdown by resistance status: 73 susceptible, 114 resistant to ciprofloxacin - Source of isolates: Regional Hospital of Talca - Isolation timeframe: September 2022 to July 2023 - Patient population: Not specified - Inclusion/exclusion criteria: Not specified - Distribution of resistant vs. susceptible isolates: 73 susceptible, 114 resistant | - S. aureus (Oxacillin): - CatBoost: AUROC = 0.86, AUPRC = 0.73, Balanced Accuracy = 0.77, F1 Score = 0.61 - RF: AUROC = 0.86, AUPRC = 0.68, Balanced Accuracy = 0.74, F1 Score = 0.56 - SVM: AUROC < 0.8, F1 Score < 0.4 - TL: AUROC < 0.8, F1 Score < 0.4 - E. coli (Ciprofloxacin): - CatBoost: AUROC = 0.91, AUPRC = 0.91, Balanced Accuracy = 0.81, F1 Score = 0.78 - LR: AUROC = 0.68, AUPRC = 0.66, Balanced Accuracy = 0.64 - SVM: AUROC = 0.77, F1 Score = 0.58 - RF: AUROC = 0.76, F1 Score = 0.63 - TL: AUROC = 0.70, AUPRC = 0.73 - K. pneumoniae (Ciprofloxacin): - CatBoost: AUROC = 0.73, AUPRC = 0.83, F1 Score = 0.78 - TL: Poor performance - SMOTE: Reduced performance by up to 10% for AUROC and AUPRC metrics | - Time to result compared to conventional methods: The proposed method can provide quicker results compared to conventional AST methods that take 24 to 72 hours. - Cost considerations or comparisons: Not explicitly mentioned. - Ease of implementation in clinical labs: Not explicitly mentioned. - Integration with existing workflows: The method is proposed to enhance clinical decision-making, suggesting integration with existing workflows. - Real-time or rapid result capabilities: The method is designed to provide quicker results, enabling more timely and appropriate therapies. - Potential impact on antimicrobial stewardship: The ability to predict antibiotic resistance quickly can help in optimizing treatments and reducing unnecessary antibiotic use. | - Technical limitations: Need for standardizing data and refining algorithms; challenges with transfer learning due to instrument variability; SMOTE technique did not improve performance. - Generalizability concerns: Need for standardizing spectrum acquisition and preprocessing procedures to enhance generalization capabilities. - Implementation barriers: Challenges in applying pre-trained models across different instruments. - Areas requiring further research: Standardization of data and algorithms; improvement of transfer learning across different instruments. | - Total: 2229 - Staphylococcus aureus: 377 - Staphylococcus epidermidis: 225 - Enterococcus faecalis: 192 - Escherichia coli: 256 - Pseudomonas aeruginosa: 222 - Klebsiella pneumoniae: 187 - S. aureus (oxacillin): Susceptible - 298, Resistant - 72 - E. coli (ciprofloxacin): Susceptible - 145, Resistant - 107 - K. pneumoniae (ciprofloxacin): Susceptible - 73, Resistant - 114 | - Staphylococcus aureus: Oxacillin - Escherichia coli: Ciprofloxacin - Klebsiella pneumoniae: Ciprofloxacin | The Regional Hospital of Talca | 1 (Regional Hospital of Talca) | Not mentioned (the paper does not specify the specimen source isolation) | VITEK ® MS instruments | 1. Baseline removal, smoothing, and peak detection of mass spectra from VITEK MS instrument. 2. Discretization of mass spectra using binning procedure with a bin size of 5 Da. 3. Linking spectra to antibiotic resistance labels using a unique numerical code. | - S. aureus (Oxacillin): - CatBoost: AUROC = 0.86, AUPRC = 0.73, Balanced Accuracy = 0.77, F1 Score = 0.61 - RF: AUROC = 0.86, AUPRC = 0.68, Balanced Accuracy = 0.74, F1 Score = 0.56 - SVM: AUROC < 0.8, F1 Score < 0.4 - TL: AUROC < 0.8, F1 Score < 0.4 - E. coli (Ciprofloxacin): - CatBoost: AUROC = 0.91, AUPRC = 0.91, Balanced Accuracy = 0.81, F1 Score = 0.78 - RF: AUROC = 0.76, F1 Score = 0.63 - SVM: AUROC = 0.77, F1 Score = 0.58 - LR: AUROC = 0.68, AUPRC = 0.66, Balanced Accuracy = 0.64 - TL: AUROC = 0.70, AUPRC = 0.73 - K. pneumoniae (Ciprofloxacin): - CatBoost: AUROC = 0.73, AUPRC = 0.83, F1 Score = 0.78 - RF: AUROC = 0.69, AUPRC = 0.79, Balanced Accuracy = 0.58, F1 Score = 0.78 - SVM: AUROC = 0.65, AUPRC = 0.76, Balanced Accuracy = 0.62, F1 Score = 0.73 - LR: AUROC = 0.60, AUPRC = 0.66, Balanced Accuracy = 0.57, F1 Score = 0.70 - TL: AUROC = 0.58, AUPRC = 0.71, Balanced Accuracy = 0.50, F1 Score = 0.76 | CatBoost | Not mentioned (the paper does not explicitly mention the use of an independent validation dataset) | Not mentioned (the paper does not provide a specific improved time to detection) | - S. aureus: 3890, 5670, 2450 Da - E. coli: 5890 Da - K. pneumoniae: Absence of 6590 Da
Zhang et al., 2023 | Rapid identification of carbapenem-resistant Klebsiella pneumoniae based on matrix-assisted laser desorption ionization time-of-flight mass spectrometry and an artificial neural network model | Intensity was transformed to Z scores using 'StandardScaler' from Python Scikit-learn package. Normalization was applied per spectrum. | 1. Initial preprocessing using R packages MaldiQuant v.1.21 and MALDIquant-Foreign v0.13. 2. Square-root transformation of intensity data. 3. Smoothing using the Savitzky-Golay filter. 4. Baseline reduction using a nonlinear iterative peak clipping algorithm. 5. Trimming spectra to 2000 to 20,000 Da range and assigning zero values to missing data. 6. Binning method to reduce mass shift effects. 7. Standardization of intensity data to Z scores using StandardScaler from Python's Scikit-learn package. | - The potential lack of generalization due to local genetic diversity. - Not examining the expression of genes associated with carbapenem resistance. - Limited data from a single clinical laboratory, which may not be representative of all possible strains or settings. - Need for more diverse data from different hospitals and regions to improve generalization ability. | 148 | 148 | Bruker Microflex LT | BAP (Blood Agar Plate), EMB (Eosin Methylene Blue agar), PEA (Presumptive Enteric Agar), and chocolate agar | - Study type: Retrospective - Setting and timeframe: Clinical microbiology laboratories of Taipei Medical University Hospital (TMUH), April 2018 to March 2022 - Data collection approach: MALDI-TOF MS and AST data from clinical isolates - Comparison groups: CRKP vs. CSKP - Validation strategy: Split into training and independent validation sets; tenfold cross-validation during model training | - Equipment/platform used: Microflex LT from Bruker Daltonik, Germany - Sample preparation methods: Overlaying 1 µL of 70% formic acid, adding α-cyano-4-hydroxycinnamic acid and trifluoroacetic acid, drying - Spectral data preprocessing steps: Square-root transformation, smoothing with Savitzky-Golay filter, baseline reduction with nonlinear iterative peak clipping algorithm - Mass range analyzed: 2000 to 20,000 Da - Number of technical replicates per isolate: 240 laser shots - Quality control measures: Calibration with bacterial test standard, using isolates with identification scores > 2 - Spectral data processing software/methods: MALDI Biotyper Compass software version 4.1, R packages MaldiQuant v.1.21 and MALDIquant-Foreign v0.13, Python's Scikit-learn package | - Specific algorithms used: Artificial Neural Network (ANN) - Training/validation/test set split ratios and sizes: 80% training, 20% independent validation - Feature selection or extraction methods: Interquartile range (IQR) to select vectors rated above Q3 - Hyperparameter optimization approach: Tenfold cross-validation and adjustment of weights during training - Cross-validation strategy: Tenfold cross-validation - Any ensemble methods or model combinations: Not mentioned - Software/programming languages used: Python v3.7, Scikit-learn package, Shap package | - Specific antibiotics tested: ertapenem, imipenem, doripenem, meropenem - Resistance mechanisms: Not specified - Reference standard: Clinical Laboratory Standards Institute M100 guidelines - Resistance breakpoints or criteria: Resistance to any carbapenem antibiotic - Multi-drug resistance patterns: Not assessed | - Total sample size: 2683 - Breakdown by resistance status: 369 CRKP, 2314 CSKP - Source of isolates: Clinical specimens from Taipei Medical University Hospital (TMUH) - Isolation timeframe: April 2018 to March 2022 - Patient population: Not specified - Inclusion/exclusion criteria: Complete AST and MALDI-TOF MS data; excluded if incomplete data - Distribution of resistant vs. susceptible isolates: Detailed in table | - Accuracy: 0.84 - Sensitivity: 0.84 - Specificity: 0.84 - F1-score: 0.84 - AUROC: 0.91 - AUPRC: 0.90 | - Time to result: Results are obtained at least 1 day earlier than traditional AST methods. - Cost considerations: The model is described as low-cost compared to traditional AST methods. - Ease of implementation: The model can be integrated into existing workflows, providing guidance before AST reports are available. - Integration with existing workflows: The model is designed to fit into current clinical practices, aiding in the selection of appropriate antibiotics. - Real-time or rapid result capabilities: The model enables rapid detection of CRKP isolates, helping clinicians make timely treatment decisions. - Potential impact on antimicrobial stewardship: The model can improve antimicrobial stewardship by enabling the selection of appropriate antibiotics, potentially reducing misuse and improving patient outcomes. | - Technical limitations: Did not examine gene expression associated with carbapenem resistance. - Sample size limitations: Data set limited to isolates from a single clinical laboratory. - Generalizability concerns: Model's efficacy needs to be tested with isolates from other laboratories or countries. - Performance issues: Potential biomarkers differ from previous studies due to local genetic diversity and equipment differences. - Need for larger validation studies: More data from diverse sources needed to improve generalization abilities. | 2683 (369 CRKP, 2314 CSKP) | Carbapenem resistance prediction using matrix-assisted laser desorption ionization time-of-flight mass spectrometry (MALDI-TOF MS) and an artificial neural network (ANN) model. | The isolates were collected from the clinical laboratories of Taipei Medical University Hospital (TMUH) from various clinical specimens including urine, sputum, blood, swabs from pus, wound, vagina, and other clinical specimens. | 1 (Taipei Medical University Hospital) | - Urine: 200 (54.2%) for CRKP, 1134 (49.0%) for CSKP - Sputum: 63 (17.1%) for CRKP, 458 (19.8%) for CSKP - Blood: 31 (8.4%) for CRKP, 153 (6.6%) for CSKP - Pus: 26 (7.1%) for CRKP, 283 (12.2%) for CSKP - Wound: 17 (4.6%) for CRKP, 80 (3.5%) for CSKP - Vagina: 4 (1.1%) for CRKP, 53 (2.3%) for CSKP - Others: 28 (7.6%) for CRKP, 153 (6.6%) for CSKP | Microflex LT, Bruker Daltonik, Germany | 1. Preprocessing using R packages MaldiQuant and MALDIquant-Foreign. 2. Square-root transformation, Savitzky-Golay filter, and baseline reduction. 3. Trimming spectra to 2000 to 20,000 Da range and handling missing data. 4. Binning to reduce mass shift effects. 5. Standardizing intensity using StandardScaler. 6. Dividing data into groups based on AST labels and identifying meaningful peaks. 7. Calculating ratios of meaningful peaks and classifying vectors using IQR. 8. Selecting vectors above Q3 threshold as features for the ML model. | - AUROC: 0.91 - AUPRC: 0.90 - Sensitivity: 0.84 - Specificity: 0.84 - F1-score: 0.84 - Accuracy: 0.84 | - Model Type: Artificial Neural Network (ANN) - Performance Metrics: - AUROC: 0.91 - AUPRC: 0.90 - Sensitivity: 0.84 - Specificity: 0.84 - F1-score: 0.84 - Accuracy: 0.84 | Yes, external validation was performed using an independent validation dataset consisting of 20% of the total data. A balanced number of CRKP and CSKP isolates (N = 74, 74) were used for this purpose. | At least 1 day earlier than traditional AST methods, which typically require 3 to 4 days. | 2480, 4967, CITETOKEN0362, CITETOKEN1506, 12
Cordovana et al., 2018 | A Full MALDI-Based Approach to Detect Plasmid-Encoded KPC-Producing Klebsiella pneumoniae | Spectra were normalized and smoothed with standard settings, but the specific method of normalization (e.g., TIC, vector norm, base-peak) is not explicitly mentioned. Normalization appears to be applied per spectrum with internal recalibration. | - Normalization and smoothing of spectra with standard settings - Screening for potential peaks for internal recalibration - Setting intensities three times higher than surrounding noise to count as peaks - Internal calibration and optimization of the algorithm for precise peak detection | - The KPC detection by subtyping depends on the association of the peak to bla KPC gene, which could be impaired if the genetic element moves independently. - The sensitivity of the subtyping approach depends on the regional epidemiology of KPC-producing strains, specifically the prevalence of clones harboring the pKpQIL plasmid. | - Subset for carbapenemase activity confirmation: 243 - Subset for application on positive blood cultures: 204 | 204 | Bruker Daltonik | Bologna: Tryptose Soy Agar with sheep blood Dortmund: Columbia Agar with sheep blood Additional media tested: Sheep Blood Agar, CHR-KPC Agar, Mueller-Hinton-Agar | - Study type: Retrospective and prospective analysis - Setting: Clinical and surveillance isolates from Italy and Germany - Timeframe: 2009 to 2017 - Data collection approach: Analysis of MALDI-TOF mass spectra - Comparison groups: KPC-producing strains, MβL-producers, OXA-48 family producers, carbapenem-resistant non-carbapenemase producers, carbapenem-susceptible isolates - Validation strategy: STAR-Carba imipenem hydrolysis assay for carbapenemase activity confirmation | - Equipment/platform used: MALDI Biotyper system (Bruker Daltonik) - Sample preparation methods: Sepsityper kit for blood cultures; culturing on agar plates - Spectral data preprocessing steps: Normalization and smoothing with standard settings; internal recalibration - Mass range analyzed: 11,109 m/z with a detection window of ± 5 m/z - Number of technical replicates per isolate: Two (spectra recorded from two spots) - Quality control measures: Specific algorithm integrated into MALDI Biotyper software; STAR-Carba assay for carbapenemase confirmation - Spectral data processing software/methods: flexAnalysis software version 3.4; MALDI Biotyper RUO software | - Specific algorithms used: Not explicitly mentioned; uses manual analysis with flexAnalysis software. - Training/validation/test set split ratios and sizes: Not mentioned. - Feature selection or extraction methods: Normalization and smoothing with standard settings; internal recalibration. - Hyperparameter optimization approach: Not mentioned. - Cross-validation strategy: Not mentioned. - Any ensemble methods or model combinations: Not mentioned. - Software/programming languages used: MALDI Biotyper system, flexAnalysis software. | - Specific antibiotics tested: ertapenem, meropenem, imipenem - Resistance mechanisms: KPC, MβL, OXA-48 - Reference standard: EUCAST epidemiological cut-offs, STAR-Carba imipenem hydrolysis assay - Resistance breakpoints or criteria: EUCAST epidemiological cut-offs - Multi-drug resistance patterns: KPC, MβL, OXA-48 producers; non-carbapenemase producers | - Total sample size: 6209 K. pneumoniae isolates - Breakdown by resistance status: - KPC-producers: 2390 - MβL-producers: 179 - OXA-48 family producers: 32 - Carbapenem-resistant non-carbapenemase producers: 221 - Carbapenem-susceptible isolates: 3387 - Source of isolates: Clinical and surveillance isolates from Italy (Bologna) and Germany (Dortmund) - Isolation timeframe: 2009 to 2017 - Patient population: Not specified - Inclusion/exclusion criteria: Not specified - Distribution of resistant vs. susceptible isolates: Detailed in the paper and table | - Sensitivity: 85.1% - Specificity: 100% - Positive Predictive Value (PPV): 100% - Negative Predictive Value (NPV): 91.5% | - Time to result: Turnaround time from 10-15 minutes to 1.5 hours, significantly faster than conventional methods. - Cost considerations: Not explicitly mentioned, but automation and ease of use may reduce costs. - Ease of implementation: Simplified by integration into existing MALDI Biotyper software; no special knowledge or tools needed. - Integration with existing workflows: Automated and integrated into routine identification process. - Real-time or rapid result capabilities: Enables real-time detection of antibiotic resistance markers. - Potential impact on antimicrobial stewardship: Allows for rapid initiation or adjustment of therapeutic and infection control measures. | - Technical limitations: Sensitivity influenced by prevalence of specific plasmids; reliance on association of peak with bla KPC gene; potential for independent genetic movement affecting specificity. - Sample size limitations: Not explicitly mentioned, but regional epidemiological variations imply potential for limited generalizability. - Generalizability concerns: Method's effectiveness may vary by region due to differences in plasmid prevalence. - Performance issues: Detection affected by sample preparation issues like operator handling and mucousity of colonies. - Need for larger validation studies: Implied by regional epidemiological variations and potential for future genetic changes. - Implementation barriers: Seamless integration into commercial software, but potential for regional variations in effectiveness. - Areas requiring further research: Expansion to other species of enterobacteria. | 6209 | Carbapenem resistance prediction using a MALDI-based approach to detect KPC-producing Klebsiella pneumoniae strains. | Italy (Microbiology Unit of the University Hospital of Bologna Policlinico Sant'Orsola-Malpighi) and Germany (Department of Microbiology of the MVZ Dr. Eberhard & Partner Dortmund) | 2 | Blood | MALDI Biotyper system (Bruker Daltonik) | - Normalization and smoothing of spectra using standard settings. - Internal recalibration to screen for potential peaks. - Detection of peaks with intensities 3 times higher than surrounding noise. - Identification of the KPC-related peak within a specific mass range (± 5 m/z around 11,109 m/z). | - Sensitivity: 85.1% - Specificity: 100% - Positive Predictive Value (PPV): 100% - Negative Predictive Value (NPV): 91.5% - STAR-Carba hydrolysis assay sensitivity and specificity: 100% | - Automated detection of a KPC-specific peak (at 11,109 m/z) using an algorithm integrated into the MALDI Biotyper system. - Confirmation of carbapenemase activity by STAR-Carba imipenem hydrolysis assay. - Application to both colonies and positive blood cultures. - Excellent specificity (100%) and good sensitivity (85.1%) for automated peak detection. - 100% sensitivity and specificity for STAR-Carba assay. - Integration into MALDI Biotyper RUO software for simplified routine usage. - Reliable and robust detection with a turnaround time of 10-15 min to 1.5 h. | Yes, external validation was performed using a large dataset of 6209 K. pneumoniae isolates from Italy and Germany, with a subset undergoing confirmation by STAR-Carba assay and application to positive blood cultures. | 10-15 minutes to 1.5 hours | 11,109 m/z