## Sheet1
Key | DOI | Url | Wearable | Type of Wearable Device | SubType of Wearable | Invasive | Applications or Use Cases | Measurement Methods | Subtype of Measurement Methods | Accuracy | Health Conditions | Targeted Health Conditions | Sweat yes/no | Saliva yes/no | Interstitial Fluid Yes/no | Fluid 01 | Fluid 02 | Fluid 03 | Fluid 04 | Fluid 05 | Glucose Yes/NO | Lactate Yes/NO | Cortisol yes/no | Uric Acid yes/no | Urea yes/no | Cholesterol yes/no | pH yes/no | creatine yes/no | Biomarker 01 | Biomarker 02 | Biomarker 03 | Biomarker 04 | Biomarker 05 | Nº of Fluid type analysed | Nº of Biomarkers detected | State of Development | Tested in Human? | Test model | Patents | CE mark | Innovative Features or Technologies | Company | Country of the company | Software | Software as Medical Device | Publication Year | Author | Title | Publication Title | ISBN | ISSN | Abstract Note | Date | Year
NY6IHN8Q | 10.1038/s41563-024-01918-9 | https://www.nature.com/articles/s41563-024-01918-9 | YES | skin-attached |  | NO | continuously monitor solid-state epidermal biomarkers (SEBs) present in the stratum corneum | electrochemical | Electrochemical | Solid Cholesterol Sensor: The cholesterol sensor demonstrates a high sensitivity of 201.34 nA nmol−1 cm−2. The detection limit for solid-state cholesterol is remarkably low at 0.26 nmol cm−2. Solid Lactate Sensor: The lactate sensor achieves a sensitivity of 93.8 nA nmol−1 cm−2. The lactate sensor boasts a detection limit of 0.51 nmol cm−2 for solid lactate. Experiments using dried human sweat samples revealed a linear relationship between the sensor's measurements and the results obtained from commercial colorimetric assay kits. The Pearson correlation coefficients were 0.82 for cholesterol and 0.90 for lactate, indicating good agreement between the methods. The study involving human volunteers showed Pearson correlation coefficients of 0.89 for lactate and 0.95 (before food intake) and 0.92 (after food intake) for cholesterol. |  | NA | - | - | - | NONE SPECIFICALLY (skin contact) | [5 empty] | yes |  |  |  | yes |  |  | cholesterol, lactate |  |  |  |  | 0 | 2 | Validation | Human | Human studies | NA | NA | The sensor utilizes a novel bilayer hydrogel system, a significant departure from conventional solid electrodes. This bilayer structure comprises an ionic conductive hydrogel (ICH) layer and an electronically conductive hydrogel (ECH) layer. | NO | NA | YES | NA | 2024 | Arwani, Ruth Theresia; Tan, Sherwin Chong Li; Sundarapandi, Archana; Goh, Wei Peng; Liu, Yin; Leong, Fong Yew; Yang, Weifeng; Zheng, Xin Ting; Yu, Yong; Jiang, Changyun; Ang, Yuan Ching; Kong, Lingxuan; Teo, Siew Lang; Chen, Peng; Su, Xinyi; Li, Hongying; Liu, Zhuangjian; Chen, Xiaodong; Yang, Le; Liu, Yuxin | Stretchable ionic–electronic bilayer hydrogel electronics enable in situ detection of solid-state epidermal biomarkers | Nature Materials |  | 1476-1122, 1476-4660 |  | 45505 | 2024
KR22IAJL | 10.1016/j.bios.2019.02.041 | https://linkinghub.elsevier.com/retrieve/pii/S0956566319301551 | YES | skin-attached |  | NO | Real-Time Biomarker Monitoring | electrochemical | Electrochemical | Lactate Sensor: The lactate sensor achieved a sensitivity of 0.395 µA/mM. This value signifies that for every 1 millimolar (mM) increase in lactate concentration, the sensor output increases by 0.395 microamperes (µA). This sensitivity is comparable to previously reported lactate sensors, indicating that the device's performance is in line with existing technologies. Oxygen Sensor: The oxygen sensor exhibited a sensitivity of 21.6 nA/[O2]%. This value represents the change in sensor output (in nanoamperes, nA) for a 1% change in dissolved oxygen concentration. The dynamic range of the oxygen sensor was determined to be between 58.5 and 178 [O2]%, indicating the range of oxygen concentrations over which the sensor can provide reliable measurements. |  | NA | - | - | - | NONE SPECIFICALLY (skin contact) | [5 empty] | yes | [6 empty] | Lactate, Oxygen |  |  |  |  | 0 | 2 | Proof-of-concept |  | Laboratory testing | NA | NA | A skin-inspired, porous, elastic and conformable. biosensor, engineered in deterministic geometries, capable of real-time. biofluid monitoring | NO | NA | NA | NA | 2019 | Ashley, Brandon K.; Brown, Matthew S.; Park, Youjoong; Kuan, Sally; Koh, Ahyeon | Skin-inspired, open mesh electrochemical sensors for lactate and oxygen monitoring | Biosensors and Bioelectronics |  | 9565663 | Research in wearable electronics has paved the way for next-generation technology, sought to create point-ofcare biosensors that combine chemical sensing on a biocompatible platform with a broad range of applications in human health monitoring. Despite significant progress, the microspatial mechanical mismatch and fluid-impermeable interface presented between skin and the electronics create adscititious problems in device lamination, conformality, and long-term monitoring. Herein, we engineered a skin-inspired, deterministically patterned, electrochemical biosensor that can be fully integrated with the curvilinear surface of the human body, while mechanically adapting to the natural stresses applied to the skin and allowing the mass transfer of gas and fluids. In particular, we developed mechanically-compliant lattice-structured biosensors for the continuous evaluation of lactate and oxygen. Systematic studies of the sensor performance were evaluated with variations in polymeric membranes and its ability to withstand commonplace harsh, multiaxial stresses. | 43586 | 2019
P33V9NFL | 10.1002/adhm.201800232 | https://onlinelibrary.wiley.com/doi/10.1002/adhm.201800232 | YES | skin-attached |  | NO | detection of volatile organic compounds (VOCs) emitted from human skin | Chemiresistive | chemiresistive sensing | The sensor exhibits a significantly higher response to polar VOCs compared to nonpolar VOCs, The research evaluated the sensor's response to 11 representative VOCs, including both polar and nonpolar compounds, at concentrations ranging from 25 to 100 ppm. The sensor demonstrated distinct and measurable responses to each of these VOCs, indicating its capability to detect and differentiate between a variety of compounds at relevant concentrations. Limit of Detection: not specified. The modification with oleic acid not only enhanced the sensor's response to nonpolar VOCs but also improved the signal-to-noise ratio, further contributing to its sensitivity. A higher signal-to-noise ratio indicates a stronger signal relative to the background noise, allowing for more accurate and reliable detection of VOCs, even at lower concentrations. |  | NA | - | - | - | Skin Emissions | [12 empty] | 2-hexanone, cyclohexanone, 1-hexanal, 1-hexanol, 1-heptanal, 3-heptanone, 1-octanal, 3-octanone, hexane, toluene, p-xylene |  |  |  |  | 1 | 11 | prototype |  | Laboratory testing | NA | NA | The sensor utilizes PANI as the sensing material due to its advantages such as low cost, ease of mass production, and biocompatibility. The researchers chemically modify the PANI thin film with oleic acid to create a hydrophobic layer. This modification is a key innovation as it addresses the challenge of humidity interference, a common issue with traditional conducting polymer-based sensors. The sensor is fabricated on a polyethylene terephthalate (PET) substrate, which provides flexibility and conformability for wearable applications. Flexible substrates are crucial for creating wearable devices that can adapt to the contours of the human body. The researchers employ a unique fabrication method called WDFA to create the flexible sensor. This technique involves spin-coating PANI on a treated coverslip, detaching the film, doping it with H3PO4, modifying it with oleic acid, and finally transferring it to the PET substrate with interdigitated electrodes. The WDFA method allows for precise control over the sensor's structure and properties. | NO | NA | NO | NA | 2018 | Deng, Yunfeng; Sun, Jiaxing; Jin, Han; Khatib, Muhammad; Li, Xianhui; Wei, Zongsu; Wang, Feng; Horev, Yehu David; Wu, Weiwei; Haick, Hossam | Chemically Modified Polyaniline for the Detection of Volatile Biomarkers of Minimal Sensitivity to Humidity and Bending | Advanced Healthcare Materials |  | 2192-2640, 2192-2659 | Abstract Flexible sensors can be widely used in future wearable devices to monitor people's health states. However, most of the sensors are sensitive to humidity and bending effects, limiting their application in a real‐world environment. A new strategy is proposed for obtaining flexible sensors with good tolerance to humidity. By integrating a hydrophobic layer on the surface of doped polyaniline, a flexible sensor that can resist water response with a concentration up to 350 ppm is developed. Good resilience against mechanical bending is also achieved in this flexible sensor. These results may trigger a renaissance in flexible sensor applications for disease diagnosing by wearable devices. | 43313 | 2018
added later (from the PDF library) | 10.1038/ncomms11650 | https://www.nature.com/articles/ncomms11650 | YES | skin-attached |  | NO | real-time health and fitness monitoring | electrochemical | Electrochemical + electrophysiological | Lactate Biosensor Sensitivity: 96 nA/mM, with a linear response across the physiological lactate range (0–28 mM). ECG Sensor Correlation: Pearson’s correlation coefficient of r = 0.975 when compared to a commercial heart rate monitor, indicating highly accurate heart rate measurements. Validation with ECG Electrodes: ECG signal morphology from the patch's printed electrodes matches that of commercial 3M Red Dot electrodes. Lactate Specificity Test: Control experiments showed negligible response without lactate oxidase (LOx), confirming the biosensor's high selectivity for lactate. Sweat Lactate Dynamics: On-body data aligned with expected sweat lactate trends, validating the biosensor's reliability during exercise. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate |  |  |  |  | 1 | 1 | Prototype/Validation |  | in-vitro + laboratory testing | NA | NA | Flexible and Conformal Design: Utilizes a 50 μm flexible polyester substrate, enabling close skin contact and wearability. Screen-printing fabrication ensures cost-effective, scalable production and precise electrode patterning, aligning with epidermal electronics principles for unobtrusive use. Hybrid Sensor Integration: Combines chemical (sweat lactate) and electrophysiological (ECG) sensing into one compact device. A hydrophobic Ecoflex layer minimizes cross-talk, ensuring independent, accurate measurements. High-Sensitivity Biosensor: Employs enzyme-based lactate detection using lactate oxidase (LOx) and amperometric detection, providing selective, sensitive, and real-time monitoring of lactate levels. Wireless Data Transmission: Features Bluetooth Low-Energy (BLE) for seamless communication with external devices, enabling real-time data access and analysis. | NO | NA | YES | NA | 2016 | Imani S, Bandodkar AJ, Mohan AM, Kumar R, Yu S, Wang J, Mercier PP. | A wearable chemical–electrophysiological hybrid biosensing system for real-time health and fitness monitoring | Nature Communications |  |  | Flexible, wearable sensing devices can yield important information about the underlying physiology of a human subject for applications in real-time health and fitness monitoring. Despite significant progress in the fabrication of flexible biosensors that naturally comply with the epidermis, most designs measure only a small number of physical or electrophysiological parameters, and neglect the rich chemical information available from biomarkers. Here, we introduce a skin-worn wearable hybrid sensing system that offers simultaneous real-time monitoring of a biochemical (lactate) and an electrophysiological signal (electrocardiogram), for more comprehensive fitness monitoring than from physical or electrophysiological sensors alone. The two sensing modalities, comprising a three-electrode amperometric lactate biosensor and a bipolar electrocardiogram sensor, are co-fabricated on a flexible substrate and mounted on the skin. Human experiments reveal that physiochemistry and electrophysiology can be measured simultaneously with negligible cross-talk, enabling a new class of hybrid sensing devices. | 42491 | 2016
M6F7ZS9Q | 10.1016/j.bios.2015.07.039 | https://www.ncbi.nlm.nih.gov/pubmed/26276541 | YES | mouthguard |  | NO | monitoring salivary uric acid | electrochemical | Electrochemical | The biosensor displays attractive electrochemical performance with high selectivity, sensitivity and stability. The resulting calibration plot for UA detection in artificial saliva obtained by wireless transmission for different UA concentrations up to 600 μM with 100 μM increments, covering possible physiological levels of SUA, is presented. The resulting calibration plot, shown in the inset, is highly linear (correlation coefficient R2=0.998) with a sensitivity of 2.45 μA/mM. | Hyperuricemia | Hyperuricemia | - | yes | - | saliva | [7 empty] | yes |  |  |  |  | uric acid (UA) |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | The device utilizes a screen-printed electrode system that has been modified with the enzyme uricase. This enzyme-modified electrode system enables the detection of uric acid in saliva. The mouthguard incorporates anatomically miniaturized instrumentation electronics, including a potentiostat, microcontroller, and a Bluetooth Low Energy (BLE) transceive. This device allows real-time wireless transmission of the sensed information to commonly used devices such as smartphones and laptops. | NO | NA | YES | NA | 2015 | Kim, J.; Imani, S.; de Araujo, W. R.; Warchall, J.; Valdes-Ramirez, G.; Paixao, T. R.; Mercier, P. P.; Wang, J. | Wearable salivary uric acid mouthguard biosensor with integrated wireless electronics | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Print) 0956-5663 (Linking) | This article demonstrates an instrumented mouthguard capable of non-invasively monitoring salivary uric acid (SUA) levels. The enzyme (uricase)-modified screen printed electrode system has been integrated onto a mouthguard platform along with anatomically-miniaturized instrumentation electronics featuring a potentiostat, microcontroller, and a Bluetooth Low Energy (BLE) transceiver. Unlike RFID-based biosensing systems, which require large proximal power sources, the developed platform enables real-time wireless transmission of the sensed information to standard smartphones, laptops, and other consumer electronics for on-demand processing, diagnostics, or storage. The mouthguard biosensor system offers high sensitivity, selectivity, and stability towards uric acid detection in human saliva, covering the concentration ranges for both healthy people and hyperuricemia patients. The new wireless mouthguard biosensor system is able to monitor SUA level in real-time and continuous fashion, and can be readily expanded to an array of sensors for different analytes to enable an attractive wearable monitoring system for diverse health and fitness applications. | 42353 | 2015
ZYDU6GIJ | 10.1038/s41598-022-06434-x | https://www.ncbi.nlm.nih.gov/pubmed/35165316 | YES | Electronic devices (combined with a Machine Learning platform) |  | NO | glucose monitoring | electrochemical | Electrochemical Impedance Spectroscopy (EIS) | The sources provide information regarding the performance of the machine learning algorithm utilized by the sweat sensor platform for glucose measurement. The algorithm achieved a goodness of fit (R²) value of 0.94. This indicates that the model explains 94% of the variance in the glucose readings, suggesting a high level of accuracy. The root mean square error (RMSE) value was reported as 0.1 mg/dL. | Diabetes | Diabetes management | yes | - | - | human eccrine sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation | Human | Human studies | NA | NA | A highly novel and completely non-invasive sweat sensor platform technology that can measure and report glucose concentrations from passively expressed human eccrine sweat using electrochemical impedance spectroscopy and affinity capture probe functionalized sensor surfaces. | EnLiSense LLC | USA | YES | NA | 2022 | Sankhala, D.; Sardesai, A. U.; Pali, M.; Lin, K. C.; Jagannath, B.; Muthukumar, S.; Prasad, S. | A machine learning-based on-demand sweat glucose reporting platform | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | Diabetes is a chronic endocrine disease that occurs due to an imbalance in glucose levels and altering carbohydrate metabolism. It is a leading cause of morbidity, resulting in a reduced quality of life even in developed societies, primarily affected by a sedentary lifestyle and often leading to mortality. Keeping track of blood glucose levels noninvasively has been made possible due to diverse breakthroughs in wearable sensor technology coupled with holistic digital healthcare. Efficient glucose management has been revolutionized by the development of continuous glucose monitoring sensors and wearable, non/minimally invasive devices that measure glucose concentration by exploiting different physical principles, e.g., glucose oxidase, fluorescence, or skin dielectric properties, and provide real-time measurements every 1-5 min. This paper presents a highly novel and completely non-invasive sweat sensor platform technology that can measure and report glucose concentrations from passively expressed human eccrine sweat using electrochemical impedance spectroscopy and affinity capture probe functionalized sensor surfaces. The sensor samples 1-5 microL of sweat from the wearer every 1-5 min and reports sweat glucose from a machine learning algorithm that samples the analytical reference values from the electrochemical sweat sensor. These values are then converted to continuous time-varying signals using the interpolation methodology. Supervised machine learning, the decision tree regression algorithm, shows the goodness of fit R(2) of 0.94 was achieved with an RMSE value of 0.1 mg/dL. The output of the model was tested on three human subject datasets. The results were able to capture the glucose progression trend correctly. Sweet sensor platform technology demonstrates a dynamic response over the physiological sweat glucose range of 1-4 mg/dL measured from 3 human subjects. The technology described in the manuscript shows promise for real-time biomarkers such as glucose reporting from passively expressed human eccrine sweat. | 44606 | 2022
3STEX5X4 | 10.1002/smll.202201861 | https://www.ncbi.nlm.nih.gov/pubmed/35676237 | YES | patch |  | NO | monitoring sweat | electrochemical | Electrochemical | The sensors show excellent selectivity toward the target analyte, low working voltages (<0.5 V), fast (5-15 s), linear response at a wide range of concentrations (10 µm to 100 mm), and sensitivities of 1 µA/decade. |  | NA | yes | - | - | artificial sweat | [12 empty] | Na+, K+ |  |  |  |  | 1 | 2 | Proof-of-concept |  | In-vitro tests | NA | NA | The device integrates several cutting-edge innovations, including the use of novel materials like rGO, advanced sensor designs like LGTs and ISMs, and sophisticated fabrication techniques like laser micromachining. | NO | NA | NA | NA | 2022 | Furlan de Oliveira, R.; Montes-Garcia, V.; Livio, P. A.; Gonzalez-Garcia, M. B.; Fanjul-Bolado, P.; Casalini, S.; Samori, P. | Selective Ion Sensing in Artificial Sweat Using Low-Cost Reduced Graphene Oxide Liquid-Gated Plastic Transistors | Small |  | 1613-6829 (Electronic) 1613-6810 (Linking) | Health monitoring is experiencing a radical shift from clinic-based to point-of-care and wearable technologies, and a variety of nanomaterials and transducers have been employed for this purpose. 2D materials (2DMs) hold enormous potential for novel electronics, yet they struggle to meet the requirements of wearable technologies. Here, aiming to foster the development of 2DM-based wearable technologies, reduced graphene oxide (rGO)-based liquid-gated transistors (LGTs) for cation sensing in artificial sweat endowed with distinguished performance and great potential for scalable manufacturing is reported. Laser micromachining is employed to produce flexible transistor test patterns employing rGO as the electronic transducer. Analyte selectivity is achieved by functionalizing the transistor channel with ion-selective membranes (ISMs) via a simple casting method. Real-time monitoring of K(+) and Na(+) in artificial sweat is carried out employing a gate voltage pulsed stimulus to take advantage of the fast responsivity of rGO. The sensors show excellent selectivity toward the target analyte, low working voltages (<0.5 V), fast (5-15 s), linear response at a wide range of concentrations (10 microm to 100 mm), and sensitivities of 1 microA/decade. The reported strategy is an important step forward toward the development of wearable sensors based on 2DMs for future health monitoring technologies. | 44743 | 2022
2XN8KXFW | 10.1002/adhm.202303511 | https://onlinelibrary.wiley.com/doi/10.1002/adhm.202303511 | YES | patch | microneedle | minimally | detection of multiple biomarkers related to Type 2 Diabetes (T2D) | Colorimetric | Colorimetric | The sources provide information on the accuracy of the wearable microneedle patch for detecting glucose, but they do not provide specific sensitivity values for any of the biomarkers the device is designed to measure. The results showed that the glucose concentration estimated by the patch was 18.63 ± 1.096 mm, which was close to the glucose meter reading of 21.9 ± 0.45 mm. This suggests that the glucose sensor on the patch provides reasonably accurate measurements, although the blood glucose level was slightly higher than the interstitial fluid glucose level, as expected based on previous studies. The sources mention that the researchers compared the results for uric acid, pH, and insulin obtained from the microneedle patch with measurements taken using commercially available meters or an ELISA kit. They state that similar results were obtained for pH and insulin, indicating the accuracy of the device for these biomarkers. However, they do not provide specific values for these measurements. The researchers note that the uric acid concentration measured by the patch was lower than that measured in the blood, attributing this discrepancy to differences in concentration between body fluids. The sources focus on demonstrating the feasibility and proof-of-concept of the wearable microneedle patch for detecting multiple biomarkers. They describe the color changes observed on the test strips integrated into the patch, which allow for semi-quantitative detection of the biomarkers. However, they do not provide quantitative sensitivity data, such as limit of detection (LOD) or limit of quantification (LOQ), for any of the sensors. | Diabetes | diabetes management | - | - | yes | interstitial fluid |  |  |  |  | yes |  |  | yes |  |  | yes |  | Glucose, Uric Acid, pH, Insulin |  |  |  |  | 1 | 4 | Proof-of-concept |  | Animal models | NA | NA | Utilizing methacrylic acid modified gelatin and polyvinyl alcohol in the fabrication of microneedles has improved their mechanical properties for skin penetration and increased swelling capacity for interstitial fluid extraction, thanks to the double crosslinking mechanism. The fabricated microneedles are further integrated with test paper functionalized with enzyme and dye molecules to detect multiple signature biomarkers of diabetes in vivo through a colorimetric reaction. | NO | NA | NA | NA | 2024 | Bao, Ziting; Lu, Sheng; Zhang, Duo; Wang, Guanyu; Cui, Xiaolin; Liu, Guozhen | Wearable Microneedle Patch for Colorimetric Detection of Multiple Signature Biomarkers in vivo Toward Diabetic Diagnosis | Advanced Healthcare Materials |  | 2192-2640, 2192-2659 | Abstract Type 2 diabetes is rapidly emerging as a global public health problem. While blood glucose monitoring has been the primary method of managing diabetes for decades, the increasing global prevalence of the disease suggests that there might be a need to identify additional biomarkers for a more precise early diagnosis. Herein, a microneedle patch based wearable sensor is developed for the purpose of diabetic diagnosis. Utilizing methacrylic acid modified gelatin and polyvinyl alcohol in the fabrication of microneedles has improved their mechanical properties for skin penetration and increased swelling capacity for interstitial fluid extraction, thanks to the double crosslinking mechanism. The fabricated microneedles are further integrated with test paper functionalized with enzyme and dye molecules to detect multiple signature biomarkers of diabetes in vivo through a colorimetric reaction. Such a wearable microneedle patch holds significant promise for the real‐time monitoring of various biomarkers related to chronic diseases and aging. | 45413 | 2024
DJ2GKTZZ | 10.1016/j.bios.2021.113496 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321005339 | YES | patch | microneedle array | minimally | detection and quantification of the ErbB2 breast cancer biomarker. | electrochemical | Electrochemical (electrochemical immunosensor). | The analytical performance of the device was tested in artificial interstitial fluid exhibiting a linear response over a wide concentration range from 10 to 250 ng/mL, with a detection limit of 4.8 ng/mL below the biomarker levels expected in breast cancer patients. As a proof of concept, the immunosensor demonstrated its ability to successfully extract ErbB2 from a phantom gel mimicking the epidermis and dermis layers, and subsequently quantify it showing a linear range from 50 to 250 ng/mL and a detection limit of 25 ng/mL. | Cancer | breast cancer | - | - | yes | Artificial interstitial fluid (aISF) | [12 empty] | ErbB2 |  |  |  |  | 1 | 1 | Proof-of-concept |  | In -vitro + laboratory testing | NA | NA | The ErbB2 immunosensor incorporates a combination of innovative technologies, including a high-density microneedle array for minimally invasive transdermal extraction, an integrated electrochemical transducer for rapid and direct biomarker quantification, and a skin-mimicking phantom gel for realistic performance evaluation | NO | NA | NO | NA | 2021 | Dervisevic, Muamer; Alba, Maria; Adams, Timothy E.; Prieto-Simon, Beatriz; Voelcker, Nicolas H. | Electrochemical immunosensor for breast cancer biomarker detection using high-density silicon microneedle array | Biosensors and Bioelectronics |  | 9565663 | Electrochemical devices for transdermal monitoring of key biomarkers are the potential next frontier of wearable technologies for point-of-care disease diagnosis, including Cancer in which Cancer is the leading cause of death worldwide with estimated 10 million deaths in 2018 according to the World Health Organization and breast cancer is one of the five most common causes of cancer death with over two million cases recorded in 2018. Early diagnosis and prognosis based on monitoring of breast cancer biomarkers is of high importance. In this work, high-density gold coated silicon microneedle arrays (Au–Si-MNA) were simultaneously used as biomarker extraction platform and electrochemical transducer, enabling the selective immunocapture of epidermal growth factor receptor 2 (ErbB2), a key breast cancer biomarker, and its subsequent quantification. The analytical performance of the device was tested in artificial interstitial fluid exhibiting a linear response over a wide concentration range from 10 to 250 ng/mL, with a detection limit of 4.8 ng/mL below the biomarker levels expected in breast cancer patients. As a proof of concept, the immunosensor demonstrated its ability to successfully extract ErbB2 from a phantom gel mimicking the epidermis and dermis layers, and subsequently quantify it showing a linear range from 50 to 250 ng/mL and a detection limit of 25 ng/mL. The uniqueness of this sensing platform combining direct transdermal biomarker extraction and quantification opens up new avenues towards the development of high performing wearable point-of-care devices. | 44501 | 2021
GIFBFKSF | 10.1038/s41598-018-24543-4 | https://www.ncbi.nlm.nih.gov/pubmed/29695724 | YES | not specified |  | NO | glycemic control | electrochemical | Electrochemical | The sensitivity calculated from the calibration curve is 0.2 ± 0.02 µA/mM with the R2 of 0.986 |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose, alcohol |  |  |  |  | 1 | 2 | prototype | Human | Human | NA | NA | Zinc oxide thin films integrated into a nanoporous flexible electrode system. | NO | NA | NA | NA | 2018 | Bhide, A.; Muthukumar, S.; Saini, A.; Prasad, S. | Simultaneous lancet-free monitoring of alcohol and glucose from low-volumes of perspired human sweat | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | A lancet-free, label-free biosensor for simultaneous detection of sweat glucose and alcohol was demonstrated using zinc oxide thin films integrated into a nanoporous flexible electrode system. Sensing was achieved from perspired human sweat at low volumes (1-3 muL), comparable to ambient conditions without external stimulation. Zinc oxide thin film electrodes were surface functionalized with alcohol oxidase enzyme and with glucose oxidase enzyme towards developing an affinity biosensor specific to the physiological relevant range of alcohol comprising of 0-2 drinks (0-50 mg/dl) and physiologically relevant range of glucose ranging from hypo- to hyper-glycaemia (50-130 mg/dl) in perspired human sweat. Sensing was achieved by measuring impedance changes associated with alcohol and glucose binding onto the sensor interface using electrochemical impedance spectroscopy with a dynamic range from 0.01-200 mg/dl and a limit of detection of 0.01 mg/dl for alcohol in human sweat. Sensor calibration in synthetic sweat containing interferents (25-200 mg/dl) and comparison using regression and Bland-Altman analysis of sweat sensor performance was done with BACtrack((R)). Combinatorial detection of glucose and ethanol in perspired human sweat and comparison of sweat sensor performance with Accu-Chek((R)) blood glucose monitoring system that we expect would be relevant for pre-diabetics and diabetics for monitoring their glucose levels and alcohol consumption. | 43215 | 2018
IX2KJQ6R | 10.1038/nature16521 | https://www.ncbi.nlm.nih.gov/pubmed/26819044 | YES | skin-attached | Array | NO | monitoring metabolites and electrolytes in sweat | electrochemical | Electrochemical | Na+ and K+ sensors had a relative standard deviation of ~1% in sensitivity while glucose and lactate sensors had a relative standard deviation of ~5% in sensitivity. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | Glucose, Lactate, Na+, K+ |  |  |  |  | 1 | 4 | prototype | Human | Human studies | NA | NA | FPCB (flexible printed circuit board) technology | NO | NA | YES | NA | 2016 | Gao, W.; Emaminejad, S.; Nyein, H. Y. Y.; Challa, S.; Chen, K.; Peck, A.; Fahad, H. M.; Ota, H.; Shiraki, H.; Kiriya, D.; Lien, D. H.; Brooks, G. A.; Davis, R. W.; Javey, A. | Fully integrated wearable sensor arrays for multiplexed in situ perspiration analysis | Nature |  | 1476-4687 (Electronic) 0028-0836 (Print) 0028-0836 (Linking) | Wearable sensor technologies are essential to the realization of personalized medicine through continuously monitoring an individual's state of health. Sampling human sweat, which is rich in physiological information, could enable non-invasive monitoring. Previously reported sweat-based and other non-invasive biosensors either can only monitor a single analyte at a time or lack on-site signal processing circuitry and sensor calibration mechanisms for accurate analysis of the physiological state. Given the complexity of sweat secretion, simultaneous and multiplexed screening of target biomarkers is critical and requires full system integration to ensure the accuracy of measurements. Here we present a mechanically flexible and fully integrated (that is, no external analysis is needed) sensor array for multiplexed in situ perspiration analysis, which simultaneously and selectively measures sweat metabolites (such as glucose and lactate) and electrolytes (such as sodium and potassium ions), as well as the skin temperature (to calibrate the response of the sensors). Our work bridges the technological gap between signal transduction, conditioning (amplification and filtering), processing and wireless transmission in wearable biosensors by merging plastic-based sensors that interface with the skin with silicon integrated circuits consolidated on a flexible circuit board for complex signal processing. This application could not have been realized using either of these technologies alone owing to their respective inherent limitations. The wearable system is used to measure the detailed sweat profile of human subjects engaged in prolonged indoor and outdoor physical activities, and to make a real-time assessment of the physiological state of the subjects. This platform enables a wide range of personalized diagnostic and physiological monitoring applications. | 42397 | 2016
XXWLAY4V | 10.3390/s19204590 | https://www.ncbi.nlm.nih.gov/pubmed/31652574 | YES | not specified |  | NO | monitoring hydration | electrochemical | Electrochemical | The precision and reliability of the data collected by the device are indicated by the R-squared values obtained from the calibration dose responses: - EIS: R-squared value of 0.9746; CA: R-squared value of 0.9403. These values, ranging from 0 to 1, indicate the goodness of fit of the calibration curve. A value closer to 1 suggests a stronger correlation between the measured signal and the actual chloride concentration. The sensor achieved linear calibration dose responses for a dynamic range of 10-100 mM (millimolar) of chloride concentration. Additionally, the device demonstrated the capability to track sweat chloride levels dynamically over a period of 180 minutes. |  | NA | yes | - | - | sweat | [12 empty] | Chloride ions |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | Non-faradaic electroanalysis using gold microelectrodes deposited on a flexible nanoporous substrate, for high nanoscale surface area to volume enhancement, | NO | NA | NA | NA | 2019 | Ganguly, A.; Prasad, S. | Passively Addressable Ultra-Low Volume Sweat Chloride Sensor | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | This work demonstrates a novel electrochemical biosensor for the detection of chloride ion levels in ultra-low volumes (1-3 microliters) of passively expressed human sweat. We present here a hydration monitor that the pediatric, geriatric, and other immune-compromised or physically inactive/sedentary population cohort can utilize, for whom the current methods of chloride quantification of active stimulation of sweat glands through iontophoresis or treadmill runs are unsuitable. In this work, non-faradaic electroanalysis using gold microelectrodes deposited on a flexible nanoporous substrate, for high nanoscale surface area to volume enhancement, was leveraged to operate in ultra-low sweat volumes of <3 microL eluted at natural rates. The specific chloride ionophore-based affinity of chloride ions resulted in the modulation of charge transfer within the electrical double layer at the electrode-sweat buffer interface, which was transduced using electrochemical impedance spectroscopy (EIS) and chronoamperometry (CA). Linear calibration dose responses with R-squared values of 0.9746 and 0.9403 for EIS and CA respectively were obtained for a dynamic range of 10-100 mM. The surface charge and the binding chemistry of the capture probe were studied using zeta potential studies and UV-Vis. The dynamic sweat chloride-tracking capability of the sensor was evaluated for a duration of 180 min. Studies were conducted to probe the efficacy of the developed sensor for passive ultra-low sweat chloride assessment on human subjects (n = 3). | 43760 | 2019
V627Z7VQ | 10.1021/acsnano.8b07413 | https://www.ncbi.nlm.nih.gov/pubmed/30543395 | YES | skin-attached | chip | NO | analysis of pH level and Na+ and K+ concentrations. | electrochemical | Microfluidics with Ion-Sensitive Field-Effect Transistors | K+ (55 mV/dec); Na+ (62mV/dec); pH (36mV/dec). |  | NA | yes | - | - | sweat | [10 empty] | yes |  | Na+, K+, pH |  |  |  |  | 1 | 3 | Prototype/Validation |  | Laboratory testing | NA | NA | This system heterogeneously integrates, on a single silicon chip, state-of-the-art ultrathin body (UTB) fully depleted silicon-on-insulator (FD-SOI) ISFET sensors with a biocompatible microfluidic interface, to deliver a "lab-on-skin" sensing platform. | NO | NA | NA | NA | 2018 | Garcia-Cordero, E.; Bellando, F.; Zhang, J.; Wildhaber, F.; Longo, J.; Guerin, H.; Ionescu, A. M. | Three-Dimensional Integrated Ultra-Low-Volume Passive Microfluidics with Ion-Sensitive Field-Effect Transistors for Multiparameter Wearable Sweat Analyzers | ACS Nano |  | 1936-086X (Electronic) 1936-0851 (Linking) | Wearable systems could offer noninvasive and real-time solutions for monitoring of biomarkers in human sweat as an alternative to blood testing. Recent studies have demonstrated that the concentration of certain biomarkers in sweat can be directly correlated to their concentrations in blood, making sweat a trusted biofluid candidate for noninvasive diagnostics. We introduce a fully on-chip integrated wearable sweat sensing system to track biochemical information at the surface of the skin in real time. This system heterogeneously integrates, on a single silicon chip, state-of-the-art ultrathin body (UTB) fully depleted silicon-on-insulator (FD-SOI) ISFET sensors with a biocompatible microfluidic interface, to deliver a "lab-on-skin" sensing platform. A full process for the fabrication of this system is proposed in this work and is demonstrated by standard semiconductor fabrication procedures. The system is capable of collecting small volumes of sweat from the skin of a human and posteriorly passively driving the biofluid, by capillary action, to a set of functionalized ISFETs for analysis of pH level and Na(+) and K(+) concentrations. Drop-casted ion-sensing membranes on different sets of sensors on the same substrate enable multiparameter analysis on the same chip, with small and controlled cross-sensitivities, whereas a miniaturized quasireference electrodes set a stable analyte potential, avoiding the use of a cumbersome external reference electrode. The progress of lab-on-skin technology reported here can lead to autonomous wearable systems enabling real-time continuous monitoring of sweat composition, with applications ranging from medicine to lifestyle behavioral engineering and sports. | 43460 | 2018
DXCPHXUC | 10.1038/srep14586 | https://www.ncbi.nlm.nih.gov/pubmed/26420511 | YES | not specified | flexible nanoporous | NO | monitoring cortisol in human sweat | Electrical | Electrical double layer (EDL) modulation and impedance measurements | 1 pg/mL or 2.75 pmol cortisol in synthetic sweat; 1 ng/mL in human sweat. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | Electrical double layer (EDL) modulated biosensor | NO | NA | NA | NA | 2015 | Munje, R. D.; Muthukumar, S.; Panneer Selvam, A.; Prasad, S. | Flexible nanoporous tunable electrical double layer biosensors for sweat diagnostics | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | An ultra-sensitive and highly specific electrical double layer (EDL) modulated biosensor, using nanoporous flexible substrates for wearable diagnostics is demonstrated with the detection of the stress biomarker cortisol in synthetic and human sweat. Zinc oxide thin film was used as active region in contact with the liquid i.e. synthetic and human sweat containing the biomolecules. Cortisol detection in sweat was accomplished by measuring and quantifying impedance changes due to modulation of the double layer capacitance within the electrical double layer through the application of a low orthogonally directed alternating current (AC) electric field. The EDL formed at the liquid-semiconductor interface was amplified in the presence of the nanoporous flexible substrate allowing for measuring the changes in the alternating current impedance signal due to the antibody-hormone interactions at diagnostically relevant concentrations. High sensitivity of detection of 1 pg/mL or 2.75 pmol cortisol in synthetic sweat and 1 ng/mL in human sweat is demonstrated with these novel biosensors. Specificity in synthetic sweat was demonstrated using a cytokine IL-1beta. Cortisol detection in human sweat was demonstrated over a concentration range from 10-200 ng/mL. | 42277 | 2015
6JD63KMV | 10.3390/s21217311 | https://www.ncbi.nlm.nih.gov/pubmed/34770617 | YES | skin-attached | electronic tongue | NO | monitoring sweat | electrochemical | Electrochemical (Voltammetry) | After analysis of the data, principal component analysis (PCA) explained a total of 87.8% of the variance, where five PCs were extracted, all with absolute Eigen values larger than one. The PCs explained a total 47.4%, 21.7%, 8.3%, 6.2%, and 4.2% of the variance, respectively. A detailed breakdown of one of each repeat measurement of different fluids per PC is also listed in Table 2 (see pag.11/17) |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes |  |  | yes |  |  |  | Glucose, lactate, ascorbic acid, urea, sodium bicarbonate (NaCarb), sodium chloride, (+ albumin only for characterisation*) |  |  |  |  | 1 | 6 | Prototype |  | In-vitro tests | NA | NA | A wearable electronic tongue is presented by incorporating metal electrodes on a flexible circuit board and used to non-invasively monitor sweat on the body. The data obtained from the measurements were treated by multivariate data processing. | Covercast AB (deducted by affiliations) | Sweden | NA | NA | 2021 | Falk, M.; Nilsson, E. J.; Cirovic, S.; Tudosoiu, B.; Shleev, S. | Wearable Electronic Tongue for Non-Invasive Assessment of Human Sweat | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | Sweat is a promising biofluid in allowing for non-invasive sampling. Here, we investigate the use of a voltammetric electronic tongue, combining different metal electrodes, for the purpose of non-invasive sample assessment, specifically focusing on sweat. A wearable electronic tongue is presented by incorporating metal electrodes on a flexible circuit board and used to non-invasively monitor sweat on the body. The data obtained from the measurements were treated by multivariate data processing. Using principal component analysis to analyze the data collected by the wearable electronic tongue enabled differentiation of sweat samples of different chemical composition, and when combined with (1)H-NMR sample differentiation could be attributed to changing analyte concentrations. | 44503 | 2021
5WFLQGRX | 10.1016/j.crmeth.2023.100579 | https://www.ncbi.nlm.nih.gov/pubmed/37751686 | YES | not specified |  | NO | real-time monitoring of multiple metabolites | electrochemical | Electrochemical | The 3D-structured biosensor shows wide linear ranges of 400-1,400 μM and 0.1-8 mM and high sensitivities of 460.5 and 283.09 μA/(mM·cm2) for lactate and glucose detection, respectively. |  | NA | - | - | yes | interstitial fluid |  |  |  |  | yes | yes | [6 empty] | lactate, glucose |  |  |  |  | 1 | 2 | Validation |  | Animal models | NA | NA | A key technological innovation incorporated into the wearable biosensor is the use of a 3D-structured nickel foam in the sensor design | NO | NA | NA | NA | 2023 | Yang, M.; Wang, H.; Cheng, J. | Continuous monitoring of multiple biomarkers with an ultrasensitive 3D-structured wearable biosensor | Cell Rep Methods |  | 2667-2375 (Electronic) 2667-2375 (Linking) | Chronic diseases call for routine management of frequent monitoring of specific biomarkers. Traditional in vitro diagnostics technologies suffer from complex sampling processes and long detection intervals, which cannot meet the need of continuous monitoring. Wearable devices taking advantage of compact size, rapid detection process, and small sample consumption are promising to take the place of endpoint detection, providing more comprehensive information about human health. Here, we proposed a fully integrated wearable system with an ultrasensitive 3D-structured biosensor for real-time monitoring of multiple metabolites. The 3D-structured biosensor shows wide linear ranges of 400-1,400 muM and 0.1-8 mM and high sensitivities of 460.5 and 283.09 muA/(mM.cm(2)) for lactate and glucose detection, respectively. We have conducted in vivo animal experiments, and the proposed wearable biosensor demonstrated high consistency with established methods. We envision that this system could provide a real-time wearable detection platform for multiple biomarker detection. | 45194 | 2023
QHFMFSQK | 10.1021/acssensors.7b00961 | https://www.ncbi.nlm.nih.gov/pubmed/29741360 | YES | patch |  | NO | monitoring sweat rate and composition | electrochemical | Electrochemical + electrical impedance | The SD varies between 6 and 11 %. As evidenced from the experimental results, the sensor can accurately quantify the flow rate that is in the physiologically relevant range of the sweat rate. |  | NA | yes | - | - | sweat | [12 empty] | H+, Na+, K+, Cl- |  |  |  |  | 1 | 4 | prototype | Human | Human studies | NA | NA | Flexible microfluidic sweat sensing patch that enhances real-time electrochemical sensing and sweat rate analysis via sweat sampling. | NO | NA | YES | NA | 2018 | Nyein, H. Y. Y.; Tai, L. C.; Ngo, Q. P.; Chao, M.; Zhang, G. B.; Gao, W.; Bariya, M.; Bullock, J.; Kim, H.; Fahad, H. M.; Javey, A. | A Wearable Microfluidic Sensing Patch for Dynamic Sweat Secretion Analysis | ACS Sens |  | 2379-3694 (Electronic) 2379-3694 (Linking) | Wearable sweat sensing is a rapidly rising research area driven by its promising potential in health, fitness, and diagnostic applications. Despite the growth in the field, major challenges in relation to sweat metrics remain to be addressed. These challenges include sweat rate monitoring for its complex relation with sweat compositions and sweat sampling for sweat dynamics studies. In this work, we present a flexible microfluidic sweat sensing patch that enhances real-time electrochemical sensing and sweat rate analysis via sweat sampling. The device contains a spiral-patterned microfluidic component that is embedded with ion-selective sensors and an electrical impedance-based sweat rate sensor on a flexible plastic substrate. The patch is enabled to autonomously perform sweat analysis by interfacing the sensing component with a printed circuit board that is capable of on-site signal conditioning, analysis, and transmission. Progressive sweat flow in the microfluidic device, governed by the pressure induced by the secreted sweat, enhances sweat sampling and electrochemical detection via a defined sweat collection chamber and a directed sweat route. The characteristic of the sweat rate sensor is validated through a theoretical simulation, and the precision and accuracy of the flow rate is verified with a commercial syringe pump and a Macroduct sweat collector. On-body simultaneous monitoring of ion (H(+), Na(+), K(+), Cl(-)) concentration and sweat rate is also demonstrated for sensor functionality. This sweat sensing patch provides an integrated platform for a comprehensive sweat secretion analysis and facilitates physiological and clinical investigations by closely monitoring interrelated sweat parameters. | 43245 | 2018
M2GMIVPN | 10.1002/advs.202103030 | https://www.ncbi.nlm.nih.gov/pubmed/34719884 | YES | not specified | Dermal Tattoo | NO | detection of health-related biomarkers | Optical | Optical | It can be seen that the tattoo clearly formed on the rabbit skin, changing the color under different concentrations of biomarkers injected subcutaneously. First, the glucose sensing area shows green because of hypoglycemia caused by fasting, while pH, uric acid, and temperature sensing areas show green, orange, and blue, respectively, corresponding to the normalphysiological conditions. After subcutaneously injecting glucose solutions under the tattooed skin, we observed obvious color change of the glucose sensing area (i.e., yellow for 6 mm and red for 8 mm), while the colors of other areas remain unchanged. Then, for the uric acid sensing area, it can be seen that the color changes from orange to green with subcutaneously injection of 0.8 mm uric acid. For the temperature sensing area, the color turns from blue to magenta after heating the skin to 39 °Cwitha hot towel. Last, the color of pH sensing area quickly turns to yellowish brown after injecting PBS solution (pH 6.5) into dermis. The Hue values also confirm the relative changes of biomarkers in skinISF. |  | NA | - | - | yes | interstitial fluid |  |  |  |  | yes |  |  | yes |  |  | yes |  | pH, glucose, uric acid (UA) |  |  |  |  | 1 | 3 | Proof-of-concept |  | In-vitro, ex-vivo, in-vivo (animal models) | NA | NA | Dermal tattoo biosensor to simultaneously detect multiple health-related biomarkers in skin interstitial fluid. The biosensor exhibits color changes in response to the change of biomarker concentration (i.e., pH, glucose, uric acid, and temperature), which can be directly read by naked eyes or captured by a camera for semi-quantitative measurement. | NO | NA | NA | NA | 2021 | He, R.; Liu, H.; Fang, T.; Niu, Y.; Zhang, H.; Han, F.; Gao, B.; Li, F.; Xu, F. | A Colorimetric Dermal Tattoo Biosensor Fabricated by Microneedle Patch for Multiplexed Detection of Health-Related Biomarkers | Adv Sci (Weinh) |  | 2198-3844 (Electronic) 2198-3844 (Linking) | Detection of biomarkers associated with body conditions provides in-depth healthcare information and benefits to disease management, where the key challenge is to develop a minimally invasive platform with the ability to directly detect multiple biomarkers in body fluid. Dermal tattoo biosensor holds the potential to simultaneously detect multiple health-related biomarkers in skin interstitial fluid because of the features of minimal invasion, easy operation, and equipment-free result reading. Herein, a colorimetric dermal tattoo biosensor fabricated by a four-area segmented microneedle patch is developed for multiplexed detection of health-related biomarkers. The biosensor exhibits color changes in response to the change of biomarker concentration (i.e., pH, glucose, uric acid, and temperature), which can be directly read by naked eyes or captured by a camera for semi-quantitative measurement. It is demonstrated that the colorimetric dermal tattoo biosensor can simultaneously detect multiple biomarkers in vitro, ex vivo, and in vivo, and monitor the changes of the biomarker concentration for at least 4 days, showing its great potential for long-term health monitoring. | 44531 | 2021
RNFKQPTY | 10.1039/c7lc00192d | https://www.ncbi.nlm.nih.gov/pubmed/28470263 | YES | eyeglasses |  | NO | monitoring of electrolytes and metabolites | electrochemical | Electrochemical | Lactate sensor characterization: detection (LOD) of 0.39 mM (CLk=3) was calculated based on the IUPAC recommendation.The lactate sensor inter-electrode reproducibility was evaluated. A deviation of 5-15% was found for high and low lactate concentrations, respectively. Such reproducibility is satisfactory considering the low cost electrode fabrication route. Potassium sensor characterization: the sensor responds instantaneously in a near-Nernstian fashion to varying potassium concentrations over the 0.1 to 100 mM range. This potential-time trace reveals that the signal reaches a steady-state within 20 s, indicating great promise for real-time monitoring of dynamic changes in the potassium sweat levels. The resulting calibration plot (shown in the inset) is linear (R2=0.99), with a Nernstian slope of 58.0 ± 4.3 mV·decade-1(n=15). The sensor offers a LOD of 10 −3.9 M, calculated in accordance with IUPAC recommendation.44 The slope and LOD are comparable to the values reported previously using similar ion-selective membranes. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | lactate, glucose, K+ |  |  |  |  | 1 | 3 | Prototype | Human | Human studies | NA | NA | Fully-integrated wireless “Lab-on-a-Glass” multiplexed biosensor platform | NO | NA | NA | NA | 2017 | Sempionatto, J. R.; Nakagawa, T.; Pavinatto, A.; Mensah, S. T.; Imani, S.; Mercier, P.; Wang, J. | Eyeglasses based wireless electrolyte and metabolite sensor platform | Lab Chip |  | 1473-0189 (Electronic) 1473-0197 (Print) 1473-0189 (Linking) | The demand for wearable sensors has grown rapidly in recent years, with increasing attention being given to epidermal chemical sensing. Here, we present the first example of a fully integrated eyeglasses wireless multiplexed chemical sensing platform capable of real-time monitoring of sweat electrolytes and metabolites. The new concept has been realized by integrating an amperometric lactate biosensor and a potentiometric potassium ion-selective electrode into the two nose-bridge pads of the glasses and interfacing them with a wireless electronic backbone placed on the glasses' arms. Simultaneous real-time monitoring of sweat lactate and potassium levels with no apparent cross-talk is demonstrated along with wireless signal transduction. The electrochemical sensors were screen-printed on polyethylene terephthalate (PET) stickers and placed on each side of the glasses' nose pads in order to monitor sweat metabolites and electrolytes. The electronic backbone on the arms of the glasses' frame offers control of the amperometric and potentiometric transducers and enables Bluetooth wireless data transmission to the host device. The new eyeglasses system offers an interchangeable-sensor feature in connection with a variety of different nose-bridge amperometric and potentiometric sensor stickers. For example, the lactate bridge-pad sensor was replaced with a glucose one to offer convenient monitoring of sweat glucose. Such a fully integrated wireless "Lab-on-a-Glass" multiplexed biosensor platform can be readily expanded for the simultaneous monitoring of additional sweat electrolytes and metabolites. | 42871 | 2017
BPSFDJSB | 10.3390/bios13030372 | https://www.ncbi.nlm.nih.gov/pubmed/36979585 | YES | patch | epidermal microfluidic sampler | NO | sweat analysis for health monitoring | Colorimetric | Colorimetric | high detection sensitivity" for glucose in the range of 0.1 mM to 1 mM. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Proof-of-concept/validation | Human | In-vitro + in-vivo (human) | NA | NA | The sources emphasize the use of low-cost laser-cutting technology as a key innovation in the fabrication process. A notable innovation lies in the integration of capillary bursting valves (CBVs). | NO | NA | NA | NA | 2023 | He, Y.; Wei, L.; Xu, W.; Wu, H.; Liu, A. | Laser-Cutted Epidermal Microfluidic Patch with Capillary Bursting Valves for Chronological Capture, Storage, and Colorimetric Sensing of Sweat | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | Flexible wearable microfluidic devices show great feasibility and potential development in the collection and analysis of sweat due to their convenience and non-invasive characteristics in health-level feedback and disease prediction. However, the traditional production process of microfluidic patches relies on resource-intensive laboratory and high-cost facilities. In this paper, a low-cost laser-cutting technology is proposed to fabricate epidermal microfluidic patches for the collection, storage and colorimetric analysis of sweat. Two different types of capillary bursting valves are designed and integrated into microchannel layers to produce two-stage bursting pressure for the reliable routing of sweat into microreservoirs in sequential fashion, avoiding the mixing of old and new sweat. Additionally, an enzyme-based reagent is embedded into the microreservoirs to quantify the glucose level in sweat by using colorimetric methods, demonstrating a high detection sensitivity at the glucose concentration from 0.1 mM to 1 mM in sweat and an excellent anti-interference performance that prevents interference from substances probably existent in sweat. In vitro and on-body experiments demonstrate the validity of the low-cost, laser-cut epidermal microfluidic patch for the chronological analysis of sweat glucose concentration and its potential application in the monitoring of human physiological information. | 44997 | 2023
A34E4MWM | 10.1016/j.bios.2023.115097 | https://www.ncbi.nlm.nih.gov/pubmed/36858023 | YES | skin-worn | microfluidic sampler | NO | real-time cortisol monitoring in sweat | electrochemical | Electrochemical | Device operation was demonstrated invitro with a broad linear dynamic range (1 pM - 1 μM) covering the physiological range and a sub-picomolar (0.2 pM) limit of detection in sweat. Real-time, on-body measurements were collected from human subjects using an induced stress protocol, demonstrating in-situ signal regeneration and the ability to detect dynamic cortisol fluctuations continuously for up to 90 min. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | Prototype/Validation | Human | In-vitro + human studies | NA | NA | Cortisol is detected using an electrochemical sensor functionalized with a pseudoknot-assisted aptamer and a flexible microfluidic sweat sampling system. The skin-worn microfluidic sampler provides rapid sweat collection while separating old and new sweat. The conformation-switching aptamer provides high specificity towards cortisol while being regenerable, allowing it to monitor temporal changes continuously. The aptamer was engineered to add a pseudoknot, restricting it to only two states, thus minimizing the background signal and enabling high sensitivity. | NO | NA | NA | NA | 2023 | Singh, N. K.; Chung, S.; Chang, A. Y.; Wang, J.; Hall, D. A. | A non-invasive wearable stress patch for real-time cortisol monitoring using a pseudoknot-assisted aptamer | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Linking) | Stress is part of everyone's life and is exacerbated by traumatic events such as pandemics, disasters, violence, lifestyle changes, and health disorders. Chronic stress has many detrimental health effects and can even be life-threatening. Long-term stress monitoring outside of a hospital is often accomplished by measuring heart rate variability. While easy to measure, this digital biomarker has low specificity, greatly limiting its utility. To address this shortcoming, we report a non-invasive, wearable biomolecular sensor to monitor cortisol levels in sweat. Cortisol is a neuroendocrine hormone that regulates homeostasis as part of the stress pathway. Cortisol is detected using an electrochemical sensor functionalized with a pseudoknot-assisted aptamer and a flexible microfluidic sweat sampling system. The skin-worn microfluidic sampler provides rapid sweat collection while separating old and new sweat. The conformation-switching aptamer provides high specificity towards cortisol while being regenerable, allowing it to monitor temporal changes continuously. The aptamer was engineered to add a pseudoknot, restricting it to only two states, thus minimizing the background signal and enabling high sensitivity. An electrochemical pH sensor allows pH-corrected amperometric measurements. Device operation was demonstrated invitro with a broad linear dynamic range (1 pM - 1 muM) covering the physiological range and a sub-picomolar (0.2 pM) limit of detection in sweat. Real-time, on-body measurements were collected from human subjects using an induced stress protocol, demonstrating in-situ signal regeneration and the ability to detect dynamic cortisol fluctuations continuously for up to 90 min. The reported device has the potential to improve prognosis and enable personalized treatments. | 45047 | 2023
F6FGVLCA | 10.1021/acssensors.1c01009 | https://www.ncbi.nlm.nih.gov/pubmed/34228919 | YES | patch |  | NO | real-time monitoring of lactate in sweat | electrochemical | Electrochemical | Correlation with a Standardized Method - comparison of the biosensor's performance against ion chromatography - Pearson correlation coefficient of 0.97, strongly supports the accuracy of the biosensor's measurement. The variations in sensitivity were reported to be less than 5%, indicating good repeatability, reproducibility and reversibility. |  | NA | yes | - | - | sweat | [5 empty] | yes |  |  |  |  | yes |  | lactate, pH |  |  |  |  | 1 | 2 | Validation | Human | Human studies | NA | NA | A novel analytical strategy based on the restriction of the lactate flux that reaches the enzyme lactate oxidase, which is immobilized in the biosensor core. This is accomplished by means of an outer plasticized polymeric layer containing the quaternary salt tetradodecylammonium tetrakis(4-chlorophenyl) borate (traditionally known as ETH500). | NO | NA | NA | NA | 2021 | Xuan, X.; Perez-Rafols, C.; Chen, C.; Cuartero, M.; Crespo, G. A. | Lactate Biosensing for Reliable On-Body Sweat Analysis | ACS Sens |  | 2379-3694 (Electronic) 2379-3694 (Linking) | Wearable lactate sensors for sweat analysis are highly appealing for both the sports and healthcare fields. Electrochemical biosensing is the approach most widely used for lactate determination, and this technology generally demonstrates a linear range of response far below the expected lactate levels in sweat together with a high influence of pH and temperature. In this work, we present a novel analytical strategy based on the restriction of the lactate flux that reaches the enzyme lactate oxidase, which is immobilized in the biosensor core. This is accomplished by means of an outer plasticized polymeric layer containing the quaternary salt tetradodecylammonium tetrakis(4-chlorophenyl) borate (traditionally known as ETH500). Also, this layer prevents the enzyme from being in direct contact with the sample, and hence, any influence with the pH and temperature is dramatically reduced. An expanded limit of detection in the millimolar range (from 1 to 50 mM) is demonstrated with this new biosensor, in addition to an acceptable response time; appropriate repeatability, reproducibility, and reversibility (variations lower than 5% for the sensitivity); good resiliency; excellent selectivity; low drift; negligible influence of the flow rate; and extraordinary correlation (Pearson coefficient of 0.97) with a standardized method for lactate detection such as ion chromatography (through analysis of 22 sweat samples collected from 6 different subjects performing cycling or running). The developed lactate biosensor is suitable for on-body sweat lactate monitoring via a microfluidic epidermal patch additionally containing pH and temperature sensors. This applicability was demonstrated in three different body locations (forehead, thigh, and back) in a total of five on-body tests while cycling, achieving appropriate performance and validation. Moreover, the epidermal patch for lactate sensing is convenient for the analysis of sweat stimulated by iontophoresis in the subjects' arm, which is of great potential toward healthcare applications. | 44400 | 2021
K4UVSAA5 | 10.1126/sciadv.aax0649 | https://www.ncbi.nlm.nih.gov/pubmed/31723600 | YES | patch |  | NO | simultaneous detection of six health-related biomarkers in sweat | electrochemical | Electrochemical | The relative standard deviations (RSDs) are 3.6% for glucose, 4.2% for lactate, 2.8% for AA, 3.9% for UA, 8.2% for Na+, and 7.5% for K+, demonstrating the good reproducibility of these sensors. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes |  | yes |  |  |  |  | glucose, lactate, ascorbic acid, uric acid, Na+, K+ |  |  |  |  | 1 | 6 | prototype | Human | Human studies | NA | NA | The device incorporates a flexible sweat analysis patch fabricated using a silk fabric-derived carbon textile. This textile exhibits a unique combination of properties: an intrinsically N-doped graphitic structure and a hierarchical woven, porous structure. | NO | NA | NA | NA | 2019 | He, W.; Wang, C.; Wang, H.; Jian, M.; Lu, W.; Liang, X.; Zhang, X.; Yang, F.; Zhang, Y. | Integrated textile sensor patch for real-time and multiplex sweat analysis | Sci Adv |  | 2375-2548 (Electronic) 2375-2548 (Linking) | Wearable sweat analysis devices for monitoring of multiple health-related biomarkers with high sensitivity are highly desired for noninvasive and real-time monitoring of human health. Here, we report a flexible sweat analysis patch based on a silk fabric-derived carbon textile for simultaneous detection of six health-related biomarkers. The intrinsically N-doped graphitic structure and the hierarchical woven, porous structure provided the carbon textile good electrical conductivity, rich active sites, and good water wettability for efficient electron transmission and abundant access to reactants, enabling it to serve as an excellent working electrode in electrochemical sensors. On the basis of the above, we fabricated a multiplex sweat analysis patch that is capable of simultaneous detection of glucose, lactate, ascorbic acid, uric acid, Na(+), and K(+). The integration of selective detectors with signal collection and transmission components in this device has enabled us to realize real-time analysis of sweat. | 43770 | 2019
FEE26R9V | 10.1016/j.bios.2023.115504 | https://linkinghub.elsevier.com/retrieve/pii/S0956566323004463 | YES | patch | 2 forms: Skin-attached sensor/Textile-embedded | NO | continuously monitor and analyze sweat in real-time | electrochemical | Electrochemical | Accurate detection within 0.2–4.0 μL/min, with a strong linear correlation (R² = 0.999) between injected flow rates and measured values. Consistent performance across devices from the same batch due to FPCB-based fabrication. Linear relationship (R² = 0.992) between conductance difference (ΔG) and NaCl concentrations ranging from 10–200 mmol/L. Measurements validated using dedicated electrodes, with close agreement confirming accuracy. Insensitivity to flow rate variations (0.5–2.0 μL/min) and specific ion types ensures reliability in diverse conditions. |  | NA | yes | - | - | sweat | [12 empty] | Total Ionic Charge Concentration, Sweat Rate and Total Sweat Loss (Sweat Volume)* |  |  |  |  | 1 | NA* | Validation | Human | in-vitro tests + laboratory testing + human studies | NA | NA | Dual-Mode Signal Decoupling: A hollowed-out serpentine microchannel with interdigitated electrodes (IDSs) allows simultaneous measurement of sweat rate and ionic charge concentration, enhanced by a decoupling algorithm for precise data analysis without complex calibration. Janus Textile Material: Features a tri-layer structure with directional sweat permeability—hydrophobic, diffusion, and transfer layers—to ensure efficient sweat collection while keeping the skin dry. Enhanced Breathability: The porous, breathable design prevents compensatory sweating and supports long-term wearability. Reusable and Washable: Durable textile and sensor design enable cost-effective, practical usage. | NO | NA | NA | NA | 2023 | Liu, Mengyuan; Wang, Shuqi; Xiong, Zuoping; Zheng, Zhuo; Ma, Nan; Li, Lianhui; Gao, Qiang; Ge, Changlei; Wang, Yongfeng; Zhang, Ting | Perspiration permeable, textile embeddable microfluidic sweat sensor | Biosensors and Bioelectronics |  | 9565663 | Epidermal microfluidic devices are continuously being developed for efficient sweat collection and sweat rate detection. However, most microfluidic designs ignore the use of airtight/adhesive substrate will block the natural perspiration of the covered sweat pores, which will seriously affect normal sweat production and long-term wearable comfort. Herein, we present a Janus textile-embedded microfluidic sensor platform with high breathability and directional sweat permeability for synchronous sweat rate and total electrolyte concentration detection. The device consists of a hollowed-out serpentine microchannel with interdigital electrodes and Janus textile. The dual-mode signal of the sweat rate (0.2–4.0 μL min 1) and total ionic charge concentration (10–200 mmol L 1) can be obtained synchronously by decoupli]ng conductance step signals generated when sweat flows through alternating interdigitated spokes at equal intervals in the microchannel. Meanwhile, the hollowed-out microchannel structure significantly reduces the coverage area of the sensor on the skin, and the Janus textile-embedded device ensures a comfortable skin/device interface (fewer sweat pores are blocked) and improves breathability (503.15 g m 2 d 1) and sweat permeability (directional liquid transportation) during longterm monitoring. This device is washable and reusable, which shows the potential to integrate with clothing and smart textile, and thus facilitate the practicality of wearable sweat sensors for personalized healthcare. | 45200 | 2023
CFR6UIMY | 10.1016/j.talanta.2021.123187 | https://www.ncbi.nlm.nih.gov/pubmed/35030501 | YES | patch | hydrogel | NO | noninvasive glucose monitoring using natural sweat | electrochemical | Electrochemical | We characterized the sweat glucose sensor response to different deposition charges. The sensitivity was estimated from the slope of the calibration curve, and the deposition charge of 0.2 mA reveals good sensitivity compared to others with a similar conclusion about the electrochemical characteristics. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation | Human | Human studies | NA | NA | A noninvasive sweat glucose sensor with convenient hydrogel patches for rapid sampling of natural perspiration without external activities that stimulate sweating. The wearable hydrogel patch rapidly takes up natural sweat from the hand and serves as a medium for electrochemical sensing. A Prussian blue-doped poly(3,4-ethylenedioxythiophene nanocomposite (PB-PEDOT NC) electrode provides cost-effective, stable and excellent electrocatalitic activity in sweat glucose measurements. | NO | NA | NA | NA | 2022 | Lin, P. H.; Sheu, S. C.; Chen, C. W.; Huang, S. C.; Li, B. R. | Wearable hydrogel patch with noninvasive, electrochemical glucose sensor for natural sweat detection | Talanta |  | 1873-3573 (Electronic) 0039-9140 (Linking) | Recent advances in microelectronics and electrochemical sensing platforms have preceded the development of devices for personal monitoring and managing physiological and metabolic information that exploit sweat as a noninvasive, convenient approach for providing information about underlying health conditions, such as glucose level monitoring. Although most sweat glucose sensors have targeted applications during exercise and other active stimulation induced-sweat, natural sweating offers an attractive alternative with minimal effect on users that can be accessed during routine and sedentary activities without impeding personal lifestyle and preserves the correlation between blood and sweat glucose. Here, we present a noninvasive sweat glucose sensor with convenient hydrogel patches for rapid sampling of natural perspiration without external activities that stimulate sweating. The wearable hydrogel patch rapidly takes up natural sweat from the hand and serves as a medium for electrochemical sensing. A prussian blue-doped poly(3,4-ethylenedioxythiophene nanocomposite (PB-PEDOT NC) electrode provides cost-effective, stable and excellent electrocatalytic activity in sweat glucose measurements. We demonstrated sweat glucose sensor functionality by long-term measurements of glucose in sweat from human subjects consuming food and drinks. By enabling the analysis of sweat glucose during routine and sedentary activities, the sweat glucose sensor shows great promise for clinical-grade glucose management and enlarges the scope of next-generation noninvasive sensing systems. | 44682 | 2022
TMEAC3VV | 10.1038/s41551-022-00887-1 | https://www.nature.com/articles/s41551-022-00887-1 | YES | microneedle | Array | minimally | continuous, real-time, and simultaneous measurement of multiple biomarkers in interstitial fluid (ISF) | electrochemical | Electrochemical | Accuracy: Correlation with Reference Methods: The device's measurements of ISF lactate, glucose, and alcohol levels were consistently correlated with those obtained using standard reference methods (blood lactate meter, blood glucose meter, and breathalyzer). Lactate: Pearson correlation coefficient (Pearson's r) of 0.94 Glucose: Mean absolute relative difference (MARD) of 8.83% Alcohol: Pearson's r of 0.94 Accuracy Metrics: The researchers calculated specific accuracy metrics to evaluate the device's performance: Lactate: Overall MARD of 16.1% Glucose: Overall MARD of 9.6% Alcohol: Overall MARD of 15.7% Sensitivity: Limits of Detection: The microneedle biosensors exhibited the following limits of detection (calculated based on a signal-to-noise ratio of 3): Glucose: 0.32 mM Lactate: 0.15 mM Alcohol: 0.50 mM Dynamic Ranges: The sensors demonstrated the ability to detect the target analytes within physiologically relevant ranges: Glucose: 0–40 mM Lactate: 0–28 mM Alcohol: 0–100 mM |  | NA | - | - | yes | interstitial fluid |  |  |  |  | yes | yes | [6 empty] | Glucose, Lactate, Alcohol |  |  |  |  | 1 | 3 | validation | Human | in-vitro tests + human studies | NA | NA | Pain-Free Microneedle Design: a cost-effective micromachining process ensures precise, reproducible microneedles for painless skin penetration and continuous interstitial fluid (ISF) access. Multiplexed Sensing: multiple regionally separated electrodes enable simultaneous real-time detection of biomarkers (e.g., glucose, lactate, alcohol) with minimal crosstalk. Advanced Biosensing Chemistry: multilayered coatings enhance selectivity, analyte recognition, and biofouling resistance, ensuring stable, accurate measurements. Wireless Integration: low-power electronics, Bluetooth connectivity, and a smartphone app provide real-time data visualization, analysis, and tracking for user empowerment. Durability and Biocompatibility: robust, biocompatible materials withstand daily wear, passing cytotoxicity tests for safe on-body use. Biofouling Prevention: a surfactant-enhanced outer layer minimizes biofouling, maintaining long-term sensor sensitivity and performance. | NO | NA | YES | NA | 2022 | Tehrani, Farshad; Teymourian, Hazhir; Wuerstle, Brian; Kavner, Jonathan; Patel, Ravi; Furmidge, Allison; Aghavali, Reza; Hosseini-Toudeshki, Hamed; Brown, Christopher; Zhang, Fangyu; Mahato, Kuldeep; Li, Zhengxing; Barfidokht, Abbas; Yin, Lu; Warren, Paul; Huang, Nickey; Patel, Zina; Mercier, Patrick P.; Wang, Joseph | An integrated wearable microneedle array for the continuous monitoring of multiple biomarkers in interstitial fluid | Nature Biomedical Engineering |  | 2157-846X |  | 44690 | 2022
F5T4UQEW | 10.1016/j.talanta.2023.124507 | https://linkinghub.elsevier.com/retrieve/pii/S0039914023002588 | YES | textile |  | NO | non-invasive, continuous sweat collection and analysis for health monitoring. | Colorimetric | Colorimetric | Here is a summary of the LOD for each analyte: Chloride: 1.06 mM Urea: 3.05 mM The researchers note that the normal ranges for the analytes in healthy individuals are: Chloride: 10–30 mM pH: 4.2–7.5 Urea: 5–40 mM The source does not provide a specific value for the LOD of pH. |  | NA | yes | - | - | sweat | [8 empty] | yes |  | yes |  | chloride, pH, urea |  |  |  |  | 1 | 3 | Proof-of-concept | Human | In-vitro tests + human studies | NA | NA | The sources highlight the use of a Janus fabric as a key technological innovation in the wearable colorimetric sweat sensing system. This fabric features opposite wettability on its two sides—hydrophobic and hydrophilic—enabling unidirectional sweat transport. | NO | NA | NA | NA | 2023 | Xi, Pengyu; He, Xuecheng; Fan, Chuan; Zhu, Qinglin; Li, Zehua; Yang, Yuemeng; Du, Xin; Xu, Tailin | Smart Janus fabrics for one-way sweat sampling and skin-friendly colorimetric detection | Talanta |  | 399140 | Functionalized textiles with biofluid management capability have attracted tremendous attention in recent years due to their significant roles in health monitoring and dehydration prevention. Here we propose a one-way colorimetric sweat sampling and sensing system based on a Janus fabric using interfacial modification techniques. The opposite wettability of Janus fabric enables sweat to be quickly moved from the skin surface to the hydrophilic side and colorimetric patches. The unidirectional sweat-wicking performance of Janus fabric not only facilitates adequate sweat sampling but also inhibits the backflow of hydrated colorimetric regent from the assay patch toward the skin, eliminating potential epidermal contaminations. On this basis, visual and portable detection of sweat biomarkers including chloride, pH, and urea is also achieved. The results show that the true concentrations of chloride, pH, and urea in sweat are ~10 mM, ~7.2, and ~10 mM, respectively. The detection limits of chloride and urea are 1.06 mM and 3.05 mM. This work bridges the gap between sweat sampling and a friendly epidermal microenvironment, providing a promising way for multifunctional textiles. | 45108 | 2023
AHH5I82G | 10.1039/c6an02571d | https://www.ncbi.nlm.nih.gov/pubmed/28300229 | YES | patch | skin-attached | NO | sugar monitoring | Impedance spectroscopy | Electrical impedance spectroscopy | Experimental results showed that impedance and capacitance variations of the device, in the measurement frequency range, responded consistently to sugar concentrations from 0 to 5 mM, in aqueous phosphate buffer solutions. Impedance measured at a frequency of 1 kHz shows a decrease with increasing sugar concentration. Additionally, the relative capacitance changes observed are markedly different for fructose and glucose, as the hydrogel exhibits greater swelling in the presence of fructose than glucose over the same concentration range. However, with this simple sensor, quantification of different types of sugars present in one sample is not possible, and for this scope an array of differently functionalised interdigitated electrodes and principal component analysis (PCA) is envisioned. | Diabetes | diabetes management | yes | - | - | PBS solution added with monosaccharides (to simulate human sweat).) |  |  |  |  | yes | yes |  |  |  |  | yes |  | Glucose, Fructose, Chloride, pH, Lactate |  |  |  |  | 1 | 5 | Proof-of-concept |  | Laboratory testing | NA | NA | Carbon interdigitated electrodes printed on paper, modified with a thin layer of responsive hydrogel | NO | NA | NA | NA | 2017 | Daikuzono, C. M.; Delaney, C.; Tesfay, H.; Florea, L.; Oliveira, O. N.; Morrin, A.; Diamond, D. | Impedance spectroscopy for monosaccharides detection using responsive hydrogel modified paper-based electrodes | Analyst |  | 1364-5528 (Electronic) 0003-2654 (Linking) | Herein we present a novel sensor for the detection of monosaccharides (e.g. glucose, fructose) in solution, using electrical impedance spectroscopy. The sensor is based on carbon interdigitated electrodes, printed on paper using screen printing. The surface of the electrodes was modified with a thin layer of hydrogel containing acrylamide copolymerised with 20 mol% 3-(Acrylamido)phenylboronic acid (PBA). It was observed that the hydrogel layers containing 20 mol% PBA swell considerably in the presence of glucose and fructose. This in turn changes the measured impedance across the electrodes, making it a suitable sensor for the quantitative detection of saccharides. We investigated the impedance and capacitance variations with different concentrations of glucose and fructose (0-5 mM) in aqueous phosphate buffer solutions. Variations in impedance were attributed to changes in the dielectric properties of the hydrogel under an applied electric field, due to swelling of the hydrogel layer induced by uptake and binding of sugar molecules to the boronate species within the gel. Impedance measurements at 1 kHz demonstrated that hydrogel swelling leads to an increased mobility of ions within the swollen hydrogel layer. The impedance decreased with increasing sugar concentration and the relative capacitance curves are markedly different for fructose and glucose, as the hydrogel exhibits greater swelling in the presence of fructose than glucose over the same concentration range. As the proposed sensor was shown to be suitable for the detection of glucose at concentration levels found in human sweat, future work will focus on the incorporation of these modified paper-based electrodes into wearable skin patches for non-invasive sugar monitoring in sweat. | 42821 | 2017
R6N6CZZS | 10.1021/acssensors.2c01827 | https://www.ncbi.nlm.nih.gov/pubmed/36455259 | YES | textile |  | NO | quantification of biomarkers in sweat for health monitoring | electrochemical | Electrochemical | The limit of detection (LOD) was determined to be 5.8 μM, consistent with most published works focusing on Ca2+ measurements, which were in the single digit μM range.This is sufficient for wearable Ca2+ detection because the physiological range in sweat is mainly around 0.2 mM to 2 mM. |  | NA | yes | - | - | sweat | [12 empty] | Ca2+ |  |  |  |  | 1 | 1 | Prototype/Validation | Human | Human studies | NA | NA | Integration of microfluidic channels directly onto fabric, creating a flexible and wearable platform for sweat analysis. The device utilizes electrodes that are screen-printed directly onto the fabric substrate. | NO | NA | YES | NA | 2022 | Zhang, T.; Ratajczak, A. M.; Chen, H.; Terrell, J. A.; Chen, C. | A Step Forward for Smart Clothes horizontal line Fabric-Based Microfluidic Sensors for Wearable Health Monitoring | ACS Sens |  | 2379-3694 (Electronic) 2379-3694 (Linking) | We report the first demonstration of fabric-based microfluidics for wearable sensing. A new technology to develop microfluidics on fabrics, as a part of an undergarment, is described here. Compared to conventional microfluidics from polydimethylsiloxane, fabric-based microfluidics are simple to make, robust, and suitable for efficient sweat delivery. Specifically, acrylonitrile butadiene styrene (ABS) films with precut microfluidic patterns were infused through fabrics to form hydrophobic areas in a specially controlled sandwich structure. Experimental tests and simulations confirmed the sweat delivery efficiency of the microfluidics. Electrodes were screen-printed onto the fabric-based microfluidic. A novel wearable potentiometer based on Arduino was also developed as the transducer and signal readouts, which was low-cost, standardized, open-source, and capable of wireless data transfer. We applied the sensor system as a standalone or as a module of a T-shirt to quantify [Ca(2+)] in a wearer's sweat, with physiological and accurate results generated. Overall, this work represents a critical step in turning regular undergarments into biochemically smart platforms for health monitoring, which will broadly benefit human healthcare. | 44918 | 2022
UC8J35PX | 10.1038/s41598-019-53677-2 | https://www.ncbi.nlm.nih.gov/pubmed/31754149 | YES | wrist-worn | clothing | NO | sweat monitoring | electrochemical | Electrochemical | The conductivity of a PEDOT:PSS fiber changes linearly according to the concentration of sodium chloride in liquid. |  | NA | yes | - | - | sweat | [12 empty] | NaCl |  |  |  |  | 1 | 1 | prototype |  | Laboratory testing | NA | NA | The device features a novel material: a flexible, hair-like conducting polymer fiber composed of PEDOT:PSS. | NO | NA | NA | NA | 2019 | Lim, T.; Kim, Y.; Jeong, S. M.; Kim, C. H.; Kim, S. M.; Park, S. Y.; Yoon, M. H.; Ju, S. | Human sweat monitoring using polymer-based fiber | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | Lightweight nano/microscale wearable devices that are directly attached to or worn on the human body require enhanced flexibility so that they can facilitate body movement and overall improved wearability. In the present study, a flexible poly(3,4-ethylenedioxythiophene):poly(styrenesulfonate) (PEDOT:PSS) fiber-based sensor is proposed, which can accurately measure the amount of salt (i.e., sodium chloride) ions in sweat released from the human body or in specific solutions. This can be performed using one single strand of hair-like conducting polymer fiber. The fabrication process involves the introduction of an aqueous PEDOT:PSS solution into a sulfuric acid coagulation bath. This is a repeatable and inexpensive process for producing monolithic fibers, with a simple geometry and tunable electrical characteristics, easily woven into clothing fabrics or wristbands. The conductivity of the PEDOT:PSS fiber increases in pure water, whereas it decreases in sweat. In particular, the conductivity of a PEDOT:PSS fiber changes linearly according to the concentration of sodium chloride in liquid. The results of our study suggest the possibility of PEDOT:PSS fiber-based wearable sensors serving as the foundation of future research and development in skin-attachable next-generation healthcare devices, which can reproducibly determine the physiological condition of a human subject by measuring the sodium chloride concentration in sweat. | 43790 | 2019
TULBVKWJ | 10.1126/sciadv.aar2904 | https://www.ncbi.nlm.nih.gov/pubmed/30035216 | YES | patch | skin-attached | NO | monitoring stress | electrochemical | Electrochemical | RSD of 5% (for two measurements). |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | prototype | Human | Human | NA | NA | Artificial receptor as a biomimetic polymeric membrane for stable and selective molecular recognition using OECTs to produce a wearable sweat diagnostics platform | NO | NA | NA | NA | 2018 | Parlak, O.; Keene, S. T.; Marais, A.; Curto, V. F.; Salleo, A. | Molecularly selective nanoporous membrane-based wearable organic electrochemical device for noninvasive cortisol sensing | Sci Adv |  | 2375-2548 (Electronic) 2375-2548 (Linking) | Wearable biosensors have emerged as an alternative evolutionary development in the field of healthcare technology due to their potential to change conventional medical diagnostics and health monitoring. However, a number of critical technological challenges including selectivity, stability of (bio)recognition, efficient sample handling, invasiveness, and mechanical compliance to increase user comfort must still be overcome to successfully bring devices closer to commercial applications. We introduce the integration of an electrochemical transistor and a tailor-made synthetic and biomimetic polymeric membrane, which acts as a molecular memory layer facilitating the stable and selective molecular recognition of the human stress hormone cortisol. The sensor and a laser-patterned microcapillary channel array are integrated in a wearable sweat diagnostics platform, providing accurate sweat acquisition and precise sample delivery to the sensor interface. The integrated devices were successfully used with both ex situ methods using skin-like microfluidics and on human subjects with on-body real-sample analysis using a wearable sensor assembly. | 43282 | 2018
3DJCTG3D | 10.1039/d2tb02142k | https://www.ncbi.nlm.nih.gov/pubmed/37254923 | YES | skin-worn | microneedle | Minimally | monitoring of cholesterol levels in interstitial fluid | electrochemical | Electrochemical | Under optimal conditions, the enzymatic microneedle sensor exhibited high sensitivity (0.201 μA μM−1) towards cholesterol in buffer solution, with good linearity over the 1–20 μM range and a correlation coefficient of 0.9910. |  | NA | - | - | yes | interstitial fluid | [9 empty] | yes |  |  | cholesterol |  |  |  |  | 1 | 1 | Prototype/Validation |  | In-vitro tests | NA | NA | The microneedle sensor was developed by incorporating platinum (Pt) and silver (Ag) wires within pyramidal microneedles containing a microcavity opening; cholesterol oxidase (ChOx) was coupled on the Pt transducer surface using bovine serum albumin and Nafion. | NO | NA | NA | NA | 2023 | Li, Z.; Kadian, S.; Mishra, R. K.; Huang, T.; Zhou, C.; Liu, S.; Wang, Z.; Narayan, R.; Zhu, Z. | Electrochemical detection of cholesterol in human biofluid using microneedle sensor | J Mater Chem B |  | 2050-7518 (Electronic) 2050-750X (Linking) | The development of a straightforward, economical, portable, and highly sensitive sensing platform for the rapid detection of cholesterol is desirable for the early diagnosis of several pathologic conditions. In this work, we present a fascinating skin-worn microneedle sensor for monitoring cholesterol in interstitial fluid samples. The microneedle sensor was developed by incorporating platinum (Pt) and silver (Ag) wires within pyramidal microneedles containing a microcavity opening; cholesterol oxidase (ChOx) was coupled on the Pt transducer surface using bovine serum albumin and Nafion. Under optimal conditions, the enzymatic microneedle sensor exhibited high sensitivity (0.201 muA muM(-1)) towards cholesterol in buffer solution, with good linearity over the 1-20 muM range and a correlation coefficient of 0.9910. The analytical performance of the microneedle sensor was also investigated in artificial interstitial fluid and a skin-mimicking phantom gel; the sensor showed great potential for skin-worn/wearable applications with excellent linearity and a low detection limit. In addition, the developed microneedle sensor showed satisfactory stability and good selectivity towards cholesterol in the presence of potential interfering biomolecules, including glucose, lactic acid, uric acid, and ascorbic acid. This sensor exhibits enormous promise for straightforward, sensitive, and minimally invasive monitoring of cholesterol. | 45112 | 2023
RHV9F42R | 10.1016/j.bios.2023.115301 | https://linkinghub.elsevier.com/retrieve/pii/S0956566323002439 | YES | textile |  | NO | real-time, non-invasive monitoring of biomarkers in sweat. | electrochemical | Electrochemical | The sources report a remarkable LOD of 280 fg mL−1 for IL-6 using the electrochemical fabric. The sensing fibers were tested using artificial sweat spiked with known concentrations of IL-6. The fabric exhibited good recovery with relative standard deviation (RSD) values of less than 10%. While the primary focus is on IL-6 detection, the sources also mention the fabric's capability to detect cortisol, a biomarker associated with stress. The LOD for cortisol using the fabric system is reported as 380 fg mL−1, demonstrating its sensitivity to this biomarker as well. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | IL-6, cortisol |  |  |  |  | 1 | 2 | Prototype/Validation | Human | In-vitro experiments, in-situ (human) | NA | NA | A novel electrochemical fabric based on aptamer-functionalized carbon nanotube/graphene fibers for real-time and in situ monitoring of IL-6, a paramount cytokine biomarker for inflammation and cancer. This fabric system possesses flexibility, anti-fatigue ability and breathability for wearable applications and can apply to different body parts in various forms. | NO | NA | YES | NA | 2023 | Chu, Hongwei; Hu, Xiaokang; Lee, Cheng-Yu; Zhang, Anning; Ye, Yang; Wang, Yuxin; Chen, Yangyang; Yan, Xiao; Wang, Xinzhong; Wei, Jun; He, Sisi; Li, Yingchun | A wearable electrochemical fabric for cytokine monitoring | Biosensors and Bioelectronics |  | 9565663 | Wearable biosensors monitoring various biomarkers in sweat provide comprehensive and prompt profiling of health states at molecular levels. Cytokines existed in sweat with trace amounts play an important role in cellular activity modulation. Unfortunately, flexible and wearable biosensors for cytokine monitoring have not yet been achieved due to the limitation of membrane-based structure and sensing strategy. Herein, we develop a novel electrochemical fabric based on aptamer-functionalized carbon nanotube/graphene fibers for real-time and in situ monitoring of IL-6, a paramount cytokine biomarker for inflammation and cancer. This fabric system possesses flexibility, anti-fatigue ability and breathability for wearable applications and can apply to different body parts in various forms. Moreover, the electrochemical fabric can track other biomarkers by replacing the coupling aptamer, serving as a universal platform for sweat analysis. This fabric-based platform holds the potential to facilitate an intelligent and personalized health monitoring approach. | 45108 | 2023
L5KQYQR9 | 10.1016/j.talanta.2024.125967 | https://linkinghub.elsevier.com/retrieve/pii/S0039914024003461 | YES | skin-attached |  | NO | real-time collection and analysis of sweat | electrochemical | Electrochemical | Na+ and K+ sensors: Repeatability: RSD < 1%, ensuring consistent measurements. Consistency: RSD < 2% across multiple sensors, highlighting reliable fabrication. Stability: RSD < 5% over half a month, indicating durable performance. Glucose sensor: Sensitivity: 981.3 nA/mM, reflecting high responsiveness to glucose concentration changes. Linearity: Accurate quantification in the 0–400 μM range. Stability: RSD < 4% over half a month, ensuring reliable long-term functionality. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | Na+, K+, Glucose |  |  |  |  | 1 | 3 | Validation | Human | in-vitro tests + laboratory testing + human studies. | NA | NA | Advanced Microfluidic System: sophisticated design with optimized channels and chambers ensures efficient, continuous sweat collection while minimizing evaporation and contamination. Screen-Printed Sensor Array: enables precise, reproducible, and cost-effective fabrication of compact electrochemical sensors for Na+, K+, and glucose detection. PEDOT:PSS for Stability: conductive polymer enhances ion-selective electrode performance by minimizing potential drift. Prussian Blue and PtNPs: prussian blue boosts glucose selectivity, while platinum nanoparticles increase sensitivity and active surface area. Flexible Substrate: polyimide material ensures wearability and comfort during physical activity. | NO | NA | NO | NA | 2024 | Yin, Yingda; Tan, Zhiguang; Zhu, Wangwang; Pu, Zhihua; Yu, Haixia; Wang, Ridong; Li, Dachao | A wearable microfluidic system for efficient sweat collection and real-time detection | Talanta |  | 399140 | Sweat is an important biofluid with rich physiological information that can evaluate human health condition. Wearable sweat sensors have received widespread attention in recent years due to the benefits of non-invasive, continuous, and real-time monitoring. Currently, an efficient device integrating sweat collection and detection is still needed. Here, a wearable sweat microfluidic system was fabricated for real-time collection and analysis of sweat. The fabricated microfluidic system consisted of four layers, including a skin adhesive layer, a microfluidic layer, an electrode layer, and a capping layer. The sweat collection rate was around 0.79 μL/min, which demonstrated efficient sweat sampling, storage, and refreshing capabilities. Simultaneous detection of multiple sweat biomarkers was achieved with a screen-printed sweat sensing array, which could realize high-precision detection of Na+, K+, and glucose. Moreover, the sensing array also showed good repeatability and stability, with a relative standard deviation of sensitivity of less than 5%. Additionally, human testing was conducted to demonstrate that this microfluidic system can continuously monitor Na+, K+, and glucose in subjects’ sweat during exercise, which showed high potential for non-invasive human health monitoring. | 45474 | 2024
9NW94WCS | 10.1016/j.bios.2024.116502 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324005074 | YES | skin-attached |  | NO | non-invasively monitor biomarkers related to oxidative stress in human sweat | electrochemical | Electrochemical | For H2O2, the sensor demonstrated a detection limit of 4.80 μM. The sensor showed a linear response to H2O2 concentrations ranging from 0.01 to 1 mM. For phosphorylated proteins, the sensor achieved a detection limit of 0.917 μg/mL. The sensor's linear range for phosphorylated proteins spanned from 0.01 to 1 mg/mL. This range encompasses the typical concentration of proteins found in sweat, ensuring accurate measurements within the relevant physiological range. The research also emphasizes the sensor's high selectivity for both H2O2 and phosphorylated proteins, meaning it can effectively distinguish these target biomarkers from other potentially interfering substances present in sweat. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | Hydrogen peroxide (H2O2), Phosphorylated proteins* (types or numbers NA) |  |  |  |  | 1 | 2* | Prototype/Validation | Human | laboratory testing + in-vitro tests + animal models + human studies (preliminary) | NA | NA | Morphology and Defect Modulation: a novel coaxial layering of CNTs/MoS2-X with sulfur-rich vacancies boosts electroactive sites, improving sensitivity for H2O2 detection. Phosphoprotein Detection: integration of TiO2 nanoparticles enables selective binding of phosphorylated proteins, broadening biomarker analysis. Microfluidic System: features a concave design for efficient sweat collection and controlled flow for real-time analysis. Flexibility and Wearability: a flexible printed circuit ensures comfort, enabling non-invasive, continuous health monitoring. Wireless Connectivity: Bluetooth integration supports real-time data transmission to smartphones, offering convenient health insights. | NO | NA | YES | NA | 2024 | Ying, Zhiye; Qiao, Lijuan; Liu, Bingxin; Gao, Li; Zhang, Peng | Development of a microfluidic wearable electrochemical sensor for the non-invasive monitoring of oxidative stress biomarkers in human sweat | Biosensors and Bioelectronics |  | 9565663 | Oxidative stress is widely recognized as a pivotal factor contributing to numerous Central Nervous System (CNS) ailments. The concentrations of hydrogen peroxide (H2O2) and phosphorylated proteins within the human body serve as crucial indicators of oxidative stress. As such, the real-time monitoring of H2O2 and phosphorylated proteins in sweat is vital for the early identification, diagnosis, and management of diseases linked to oxidative stress. In this context, we present a novel microfluidic wearable electrochemical sensor by modifying the electrode with Prussian blue (PB) and loading sulfur-rich vacancy-containing molybdenum disulfide (MoS2-X) onto Multi-walled carbon nanotube (CNTs) to form coaxially layered CNTs/MoS2-X, which was then synthesized with highly dispersed titanium dioxide nanoparticles (TiO2) to synthesize CNTs/MoS2-X/TiO2 composites for the detection of human sweat H2O2 and phosphorylated proteins, respectively. This structure, with its sulfur vacancies and coaxial layering, significantly improved sensitivity of electrochemical sensors, allowing it to detect H2O2 in a range of 0.01–1 mM with a detection limit of 4.80 μM, and phosphoproteins in a range of 0.01–1 mg/ mL with a threshold of 0.917 μg/mL. Furthermore, the miniature sensor demonstrates outstanding performance in detecting analytes in both simulated and real sweat. Comprehensive biosafety assessments have validated the compatibility of the electrode material, underscoring the potential of sensor as a reliable and non-invasive method for tracking biomarkers linked to CNS disorders. This microfluidic wearable electrochemical biosensor with high performance and biosafety features shows great promise for the development of cutting-edge wearable technology devices for tracking CNS disease indicators. | 45566 | 2024
JZH2EGNN | 10.1016/j.bios.2024.116712 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324007188 | YES | skin-attached |  | NO | monitoring amino acid levels and pH in sweat | electrochemical | Electrochemical | The device's performance metrics include: Regression Analysis (R-value): the BPNN model achieved R-values >0.99 for tyrosine, tryptophan, and pH, indicating strong prediction accuracy in controlled settings. Mean Absolute Error (MAE): prediction errors averaged 2.79 for tryptophan and 5.41 for tyrosine. Root Mean Square Error (RMSE): RMSE values were 3.59 for tryptophan and 6.90 for tyrosine, reflecting low deviations from actual values. Relative Standard Deviation (RSD): <1% RSD for modified electrodes during bending/twisting tests. RSD of 0.67% (MWCNT/LIGE) and 0.41% (CB/LIGE) during flow rate tests. Predicted pH values showed RSDs of 0.86% in PBS and 1.26% in sweat. HPLC Validation: high Pearson correlation coefficients (0.980 for tyrosine, 0.988 for tryptophan) confirm strong agreement with HPLC standards. These results demonstrate high precision and reliability, with further real-world testing needed for comprehensive validation. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | Tyrosine, Tryptophan, pH |  |  |  |  | 1 | 3 | Validation | Human | laboratory testing + in-vitro tests + human studies (preliminary) | NA | NA | Laser-Induced Graphene (LIG) Electrodes: Flexible, biocompatible, and highly sensitive electrodes with high electron mobility and large surface area, directly patterned on a polyimide substrate. Dual-Working Electrode System: Multi-walled carbon nanotube (MWCNT) and carbon black (CB) electrodes enable selective detection of overlapping sweat analytes using machine learning. Microfluidic Integration: Continuous sweat sampling improves stability and prevents direct skin-electrode contact. Machine Learning Analysis: A backpropagation neural network (BPNN) model enhances selectivity and accuracy in detecting tyrosine, tryptophan, and pH. Miniaturized Electrochemical Workstation: Compact design with signal processing and wireless communication for real-time monitoring. | NO | NA | YES | NA | 2024 | Zhou, Zhongzeng; He, Xuecheng; Xiao, Jingyu; Pan, Jiuxiang; Li, Mengmeng; Xu, Tailin; Zhang, Xueji | Machine learning-powered wearable interface for distinguishable and predictable sweat sensing | Biosensors and Bioelectronics |  | 9565663 | The constrained resources on wearable devices pose a challenge in meeting the demands for comprehensive sensing information, and current wearable non-enzymatic sensors face difficulties in achieving specific detection in biofluids. To address this issue, we have developed a highly selective non-enzymatic sweat sensor that seamlessly integrates with machine learning, ensuring reliable sensing and physiological monitoring of sweat biomarkers during exercise. The sensor consists of two electrodes supported by a microsystem that incorporates signal processing and wireless communication. The device generates four explainable features that can be used to accurately predict tyrosine and tryptophan concentrations, as well as sweat pH. The reliability of this device has been validated through rigorous statistical analysis, and its performance has been tested in subjects with and without supplemental amino acid intake during cycling trials. Notably, a robust linear relationship has been identified between tryptophan and tyrosine concentrations in the collected samples, irrespective of the pH dimension. This innovative sensing platform is highly portable and has significant potential to advance the biomedical applications of non-enzymatic sensors. It can markedly improve accuracy while decreasing costs. | 45627 | 2024
TBGFRBWU | 10.1016/j.bios.2019.04.058 | https://www.ncbi.nlm.nih.gov/pubmed/31096082 | YES | eyeglasses |  | NO | monitoring of tear biomarkers | electrochemical | Electrochemical | The correlation between tear alcohol readings and BAC measurements from a breathalyzer showed a Pearson's r value of 0.852 in in-vitro studies and an average R<sup>2</sup> of 0.912 in on-body experiments. A general correlation with a Pearson's r value of ~0.7 was observed between blood glucose levels from a glucometer and the device's readings in on-body experiments |  | NA | - | - | - | tear |  |  |  |  | yes | [7 empty] | Alcohol, glucose, vitamin B2, vitamin C, vitamin B6. |  |  |  |  | 1 | 5 | Proof-of-concept | Human | Human studies | NA | NA | The device integrates a microfluidic electrochemical detector into an eyeglasses nose-bridge pad, for non-invasive monitoring of key tear biomarkers. The alcohol-oxidase (AOx) biosensing fluidic system allowed real-time tear collection and direct alcohol measurements in stimulated tears, leading to the first wearable platform for tear alcohol monitoring. | NO | NA | YES | NA | 2019 | Sempionatto, J. R.; Brazaca, L. C.; Garcia-Carmona, L.; Bolat, G.; Campbell, A. S.; Martin, A.; Tang, G.; Shah, R.; Mishra, R. K.; Kim, J.; Zucolotto, V.; Escarpa, A.; Wang, J. | Eyeglasses-based tear biosensing system: Non-invasive detection of alcohol, vitamins and glucose | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Print) 0956-5663 (Linking) | We report on a wearable tear bioelectronic platform, integrating a microfluidic electrochemical detector into an eyeglasses nose-bridge pad, for non-invasive monitoring of key tear biomarkers. The alcohol-oxidase (AOx) biosensing fluidic system allowed real-time tear collection and direct alcohol measurements in stimulated tears, leading to the first wearable platform for tear alcohol monitoring. Placed outside the eye region this fully wearable tear-sensing platform addresses drawbacks of sensor systems involving direct contact with the eye as the contact lenses platform. Integrating the wireless electronic circuitry into the eyeglasses frame thus yielded a fully portable, convenient-to-use fashionable sensing device. The tear alcohol sensing concept was demonstrated for monitoring of alcohol intake in human subjects over multiple drinking courses, displaying good correlation to parallel BAC measurements. We also demonstrate for the first time the ability to monitor tear glucose outside the eye and the utility of wearable devices for monitoring vitamin nutrients in connection to enzymatic flow detector and rapid voltammetric scanning, respectively. These developments pave the way to build an effective eyeglasses system capable of chemical tear analysis. | 43661 | 2019
98P5WNTU | 10.1016/j.snb.2021.131204 | https://linkinghub.elsevier.com/retrieve/pii/S092540052101772X | YES | skin-attached |  | NO | monitoring of glucose levels in sweat. | electrochemical | Photoelectrochemistry + enzymatic reactions | The wearable PEC sensor patch exhibits high sensitivity for glucose detection, with a reported limit of detection (LOD) of 22.2 pM. Correlation with Commercial Assay Kit: sweat glucose measurements obtained using the PEC sensor patch showed strong agreement with results from a commercial glucose assay kit. This indicates the device's reliability in quantifying glucose levels. Relationship with Blood Glucose Levels: while sweat glucose concentrations are generally lower than blood glucose, the researchers found a correlation between the two. This suggests the device can provide valuable insights into blood glucose fluctuations by tracking sweat glucose levels. | Diabetes | diabetic diagnosis and management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype | Human | Human studies | NA | NA | Unbiased Detection Mode: Operates at zero potential, eliminating the need for batteries, simplifying the design, and reducing costs. High Sensitivity: Achieves a detection limit of 22.2 pM using SiO2 nanoparticles for noise reduction and a BiVO4-MWNT composite for superior photocatalytic properties. Flexible and Durable Design: Features flexible three-electrode arrays and a PDMS sweat collector, ensuring consistent performance under mechanical deformation. Integrated Sweat Collection: Includes a laser-engraved hydrophilic PDMS channel for efficient sweat collection and delivery to the sensor's surface. These innovations enhance sensitivity, durability, and usability in wearable biosensing. | NO | NA | NA | NA | 2022 | Liu, Yanwen; Zhong, Lin; Zhang, Shubo; Wang, Juan; Liu, Zhihong | An ultrasensitive and wearable photoelectrochemical sensor for unbiased and accurate monitoring of sweat glucose | Sensors and Actuators B: Chemical |  | 9254005 | Wearable electrochemical sensors have been extensively used for the non-invasive and quantitative on-body detection of biomarkers in sweat. However, the photoelectrochemical (PEC) sensor, which is an evolutionary generation of the electrochemical method and holds higher sensitivity and the ability to unbiased analysis, is rarely applied to construct flexible wearable sweat devices. Herein, we demonstrated a flexible PEC patch comprised of an ultrasensitive PEC sensor and a hydrophilic poly-(dimethylsiloxane) sweat collector for unbiased and high-performance sweat glucose detection. The proposed PEC sensor offered excellent durability against mechanical deformations, low detection limit (22.2 pM), desirable selectivity, and high storage stability. Moreover, the feasibility and accuracy of this wearable glucose sensor patch were validated through human subjects studies, and the obtained sweat glucose test results closely matched with the commercial kit and correlated well with the blood glucose level. Taken together, we could confirm the considerable promise of the unbiased wearable PEC sensor for patient-friendly diabetic diagnosis. | 44621 | 2022
ZYYXPBLE | 10.1016/j.bios.2020.112038 | https://www.ncbi.nlm.nih.gov/pubmed/31989942 | YES | wrist-worn | smartwatch-like | NO | metabolic health and physiological monitoring | electrochemical | Electrochemical. | The sensors demonstrated sensitivities of: 26.31 μA mM-1·cm-2 for glucose; 1.49 μA mM-1·cm-2 for lactate; 54 mV·pH-1 for pH. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes |  |  |  |  | yes |  | glucose, lactate, pH |  |  |  |  | 1 | 3 | Prototype/Validation |  | In-vitro tests | NA | NA | A wearable system with a flexible sensor array for non-invasive and continuous monitoring of human biochemistry. The system includes signal conditioning, processing, and transmission parts for continuous measurement of glucose, lactate, pH, and temperature. The system can operate three discrete electrochemical cells.Utilization of redundant working electrodes improves the measurement accuracy of sensors by averaging chronoamperometric responses across the array. | NO | NA | NA | NA | 2020 | Yokus, M. A.; Songkakul, T.; Pozdin, V. A.; Bozkurt, A.; Daniele, M. A. | Wearable multiplexed biosensor system toward continuous monitoring of metabolites | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Linking) | Comprehensive metabolic panels are the most reliable and common methods for monitoring general physiology in clinical healthcare. Translation of this clinical practice to personal health and wellness tracking requires reliable, non-invasive, miniaturized, ambulatory, and inexpensive systems for continuous measurement of biochemical analytes. We report the design and characterization of a wearable system with a flexible sensor array for non-invasive and continuous monitoring of human biochemistry. The system includes signal conditioning, processing, and transmission parts for continuous measurement of glucose, lactate, pH, and temperature. The system can operate three discrete electrochemical cells. The system draws 15 mA under continuous operation when powered by a 3.7 V 150 mAh battery. The analog front-end of the electrochemical cells has four potentiostats and three multiplexers for multiplexed and parallel readout from twelve working electrodes. Utilization of redundant working electrodes improves the measurement accuracy of sensors by averaging chronoamperometric responses across the array. The operation of the system is demonstrated in vitro by simultaneous measurement of glucose and lactate, pH, and skin temperature. In benchtop measurements, the sensors are shown to have sensitivities of 26.31 muA mM(-1).cm(-2) for glucose, 1.49 muA mM(-1).cm(-2) for lactate, 54 mV.pH(-1) for pH, and 0.002 degrees C(-1) for temperature. With the custom wearable system, these values were 0.84 +/- 0.03 mV muM(-1).cm(-2) or glucose, 31.87 +/- 9.03 mV mM(-1).cm(-2) for lactate, 57.18 +/- 1.43 mV.pH(-1) for pH, and 63.4 muV. degrees C(-1) for temperature. This miniaturized wearable system enables future evaluation of temporal changes of the sweat biomarkers. | 43922 | 2020
JQS2327V | 10.1039/c8nr04315a | https://www.ncbi.nlm.nih.gov/pubmed/30431030 | YES | not specified | nanosensor | NO | measurement of cytokine biomarkers | Other | Aptamer-based affinity | 26 pM in a repeatable manner. |  | NA | yes | - | - | sweat | [12 empty] | TNF-α |  |  |  |  | 1 | 1 | prototype |  | Laboratory testing | NA | NA | Aptamer functionalized, graphene-based field effect transistor (GFET) nanosensor on a flexible, SiO2 coated substrate of the polymer polyethylene naphthalate (PEN). | NO | NA | YES | NA | 2018 | Hao, Z.; Wang, Z.; Li, Y.; Zhu, Y.; Wang, X.; De Moraes, C. G.; Pan, Y.; Zhao, X.; Lin, Q. | Measurement of cytokine biomarkers using an aptamer-based affinity graphene nanosensor on a flexible substrate toward wearable applications | Nanoscale |  | 2040-3372 (Electronic) 2040-3364 (Print) 2040-3364 (Linking) | We present an approach for the label-free detection of cytokine biomarkers using an aptamer-functionalized, graphene-based field effect transistor (GFET) nanosensor on a flexible, SiO2-coated substrate of the polymer polyethylene naphthalate (PEN). The nanosensor conforms to the underlying nonplanar surface and performs GFET-based rapid transduction of the aptamer-biomarker binding, thereby potentially allowing the detection of cytokine biomarkers that are sampled reliably from human bodily fluids (e.g., sweat) in wearable sensing applications. In characterizing the suitability of the nanosensor for wearable applications, we investigate the effects of substrate bending on the equilibrium dissociation constant between the aptamer and the biomarker as well as the graphene transconductance. The utility of the nanosensor is demonstrated by the detection of tumor necrosis factor-alpha (TNF-alpha), an inflammatory cytokine biomarker. Experimental results show that the flexible nanosensor can specifically respond to changes in the TNF-alpha concentration within 5 minutes with a limit of detection as low as 26 pM in a repeatable manner. | 43433 | 2018
2KEDTNX7 | 10.1016/j.talanta.2020.120922 | https://linkinghub.elsevier.com/retrieve/pii/S0039914020302137 | YES | patch |  | NO | non-invasive monitoring of specific biomarkers in sweat | electrochemical | Electrochemical | Pearson correlation for glucose variation rates between sweat and blood: 0.75 Pearson correlation for lactate variation rates between sweat and blood: >0.8 High agreement between flow-through biosensor readout and independent sweat measurements, with a Pearson correlation coefficient of >0.98. The device successfully tracked blood glucose dynamics during a glucose tolerance test, showing close correspondence with invasive measurements. | Diabetes | diabetes and hypoxia | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | glucose, lactate |  |  |  |  | 1 | 2 | Validation | Human | Human studies | NA | NA | The wearable sweat-monitoring device incorporates a flow-through biosensor system for continuous sweat analysis, addressing limitations of earlier "tattoo" biosensors that could not accurately monitor successive sweat samples. The flow-through design allows freshly secreted sweat to flow over the sensor surface, enabling real-time monitoring of glucose and lactate levels without mixing previous samples. This innovation improves accuracy and responsiveness, making it ideal for tracking dynamic changes in biomarkers, particularly in diabetes management and hypoxia detection. | NO | NA | NA | NA | 2020 | Karpova, Elena V.; Karyakina, Elena E.; Karyakin, Arkady A. | Wearable non-invasive monitors of diabetes and hypoxia through continuous analysis of sweat | Talanta |  | 399140 | We present here wearable devices for continuous monitoring of diabetes and hypoxia based on continuous analysis of sweat. To induce sweating the clinically relevant procedure (pilocarpine electrophoresis) is used. Being a sufficient requirement for diagnostics, positive correlations in variation rates between glucose and lactate concentrations in sweat and the corresponding values in blood are shown. Continuous monitoring of human condition is possible only with the use of flow-through wearable devices providing a delivery of sweat to the biosensor almost immediately after secretion. Evaluating blood glucose through continuous sweat analysis upon glucose tolerance test, we clearly show that diabetics can actually be monitored reliably via non-invasive approach. | 44044 | 2020
K2F8QUM5 | 10.1038/s41598-020-64406-5 | https://www.ncbi.nlm.nih.gov/pubmed/32382047 | YES | wrist and ventral forearm | band | NO | sweat monitoring | electrochemical | Combined: Electrochemical + Ion-Selective Electrodes (ISEs) +Conductometry | On increasing the exercise load from 100 W to 200 W the sweat chloride concentration increased from 12.0 ± 5.9 to 31.4 ± 16 mM (mean ± SD). On decreasing the load from 200 W to 100 W, the sweat chloride concentration decreased from 27.7 ± 10.5 to 14.8 ± 8.1 mM. The half-time associated with the change in sweat chloride was about 6 minutes. While the changes in sweat chloride were statistically significant, there was no correlation with changes in sweat rate or other physiological parameters, which we attribute to intra-individual variation (SD = 1.68.1 mM). |  | NA | yes | - | - | sweat | [12 empty] | chloride |  |  |  |  | 1 | 1 | Validation | Human | Human studies | NA | NA | The sensor likely incorporates advanced signal processing techniques to ensure accurate and reliable data interpretation.The ability to provide real-time feedback on sweat chloride levels requires efficient algorithms capable of rapidly processing and analyzing the sensor's output. | NO | NA | NA | NA | 2020 | Choi, D. H.; Kitchen, G. B.; Stewart, K. J.; Searson, P. C. | The Dynamic Response of Sweat Chloride to Changes in Exercise Load Measured by a Wearable Sweat Sensor | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | Wearable sensors enable the monitoring of an individual's sweat composition in real time. In this work, we recorded real-time sweat chloride concentration for 12 healthy subjects in three different protocols involving step changes in exercise load and compared the results to laboratory-based analysis. The sensor results reflected the changes in exercise load in real time. On increasing the exercise load from 100 W to 200 W the sweat chloride concentration increased from 12.0 +/- 5.9 to 31.4 +/- 16 mM (mean +/- SD). On decreasing the load from 200 W to 100 W, the sweat chloride concentration decreased from 27.7 +/- 10.5 to 14.8 +/- 8.1 mM. The half-time associated with the change in sweat chloride, defined as the time at which the concentration reached half of the overall change, was about 6 minutes. While the changes in sweat chloride were statistically significant, there was no correlation with changes in sweat rate or other physiological parameters, which we attribute to intra-individual variation (SD = 1.6-8.1 mM). The response to exercise-induced sweating was significantly different to chemically-induced sweating where the sweat chloride concentration was almost independent of sweat rate. We speculate that this difference is related to changes in the open probability of the CFTR channel during exercise, resulting in a decrease in reabsorption efficiency at higher sweat rates. | 43958 | 2020
JFQZ2JTD | 10.1016/j.bios.2021.113685 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321007223 | YES | skin-attached |  | NO | continuous, real-time, and non-invasive monitoring of various physiological and electrochemical parameters | electrochemical | Electrochemical + electrical | The fabricated electrochemical glucose biosensor exhibited excellent linearity (R2 = 0.9986) and sensitivity (29.10 μA mM-1 cm-2), covering the normal range of human sweat. The potentiometric pH sensor displayed a good response with an excellent sensitivity of -77.81 mV/pH and high linearity (R2 = 0.991), indicating that it can distinguish variations in the typical pH range for human sweat. Furthermore, the P, QRS complex, and T peaks in the measured ECG waveforms could be clearly distinguished, indicating the reliability of the fabricated flexible dry electrodes for continuous monitoring. | Diabetes | monitoring and management of diabetes and cardiovascular health | yes | - | - | sweat |  |  |  |  | yes | [5 empty] | yes |  | glucose, pH |  |  |  |  | 1 | 2 | Prototype/Validation | Human | laboratory testing + human studies (preliminary) | NA | NA | Flexible Substrate: a polyimide base ensures conformal skin attachment, enhancing comfort and reliability for wearable use. Integrated Multi-Sensor Platform: combines glucose, pH, temperature, and ECG sensors in a miniaturized design, enabling comprehensive health monitoring. High-Sensitivity Glucose Sensor: features rGO and PtNP-enhanced electrodes for precise glucose detection at low concentrations in sweat. Advanced pH Sensor: uses PANI and PVB-coated Ag/AgCl electrodes for stable and selective pH measurement. Dry ECG Electrodes: flexible gold electrodes allow gel-free, comfortable heart monitoring. Temperature Sensor: resistance-based design provides accurate, continuous temperature tracking. | NO | NA | NA | NA | 2022 | Yoon, Sanghyuk; Yoon, Hyosang; Zahed, Md Abu; Park, Chani; Kim, Dongkyun; Park, Jae Yeong | Multifunctional hybrid skin patch for wearable smart healthcare applications | Biosensors and Bioelectronics |  | 9565663 | Recent advances in wearable patches have included various sensors to monitor either physiological signs, such as the heart rate and respiration rate, or metabolites. Nevertheless, most of these have focused only on a single physiological measurement at a time, which significantly inhibits the calibration of various biological signals and diagnostic facilities. In this study, a novel multifunctional hybrid skin patch was developed for the electrochemical analysis of sweat glucose levels and simultaneous monitoring of electrocardiograms (ECGs). Furthermore, pH and temperature sensors were co-integrated onto the same patch for the calibration of the glucose biosensor to prevent inevitable inhibition and weakening of enzyme activity due to changes in the sweat pH and temperature levels. The fabricated electrochemical glucose biosensor exhibited excellent linearity (R2 = 0.9986) and sensitivity (29.10 μA mM 1 cm 2), covering the normal range of human sweat. The potentiometric pH sensor displayed a good response with an excellent sensitivity of 77.81 mV/pH and high linearity (R2 = 0.991), indicating that it can distinguish variations in the typical pH range for human sweat. Furthermore, the P, QRS complex, and T peaks in the measured ECG waveforms could be clearly distinguished, indicating the reliability of the fabricated flexible dry electrodes for continuous monitoring. The fabricated skin patch overcomes the inconvenience of the mandatory attachment of multiple patches on the human body by fully integrating all the electrochemical and electrophysiological sensors on a single patch, thus facilitating advanced glycemic control and continuous ECG monitoring for smart management of chronic diseases and healthcare applications. | 44562 | 2022
D29U6F8K | 10.1126/sciadv.aaz0007 | https://www.ncbi.nlm.nih.gov/pubmed/32219164 | YES | wrist-worn | smartwatch-like | NO | to monitor the sweat metabolite profiles of individuals in sedentary and high-intensity | electrochemical | Electrochemical. | The measured sensor responses across all conditions exhibit negligible fluctuations (<8%), indicating the suitability of the FESS-based smartwatch to render high-fidelity biomarker data recording in the presence of unconstrained body motion. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | glucose, lactate |  |  |  |  | 1 | 2 | Prototype/Validation | Human | Human studies | NA | NA | Freestanding Electrochemical Sensing System (FESS) into a smartwatch platform, creating a self-contained unit capable of sweat sampling, analysis, and data display/transmission. | NO | NA | YES | NA | 2020 | Zhao, Y.; Wang, B.; Hojaiji, H.; Wang, Z.; Lin, S.; Yeung, C.; Lin, H.; Nguyen, P.; Chiu, K.; Salahi, K.; Cheng, X.; Tan, J.; Cerrillos, B. A.; Emaminejad, S. | A wearable freestanding electrochemical sensing system | Sci Adv |  | 2375-2548 (Electronic) 2375-2548 (Linking) | To render high-fidelity wearable biomarker data, understanding and engineering the information delivery pathway from epidermally retrieved biofluid to a readout unit are critical. By examining the biomarker information delivery pathway and recognizing near-zero strained regions within a microfluidic device, a strain-isolated pathway to preserve biomarker data fidelity is engineered. Accordingly, a generalizable and disposable freestanding electrochemical sensing system (FESS) is devised, which simultaneously facilitates sensing and out-of-plane signal interconnection with the aid of double-sided adhesion. The FESS serves as a foundation to realize a system-level design strategy, addressing the challenges of wearable biosensing, in the presence of motion, and integration with consumer electronics. To this end, a FESS-enabled smartwatch was developed, featuring sweat sampling, electrochemical sensing, and data display/transmission, all within a self-contained wearable platform. The FESS-enabled smartwatch was used to monitor the sweat metabolite profiles of individuals in sedentary and high-intensity exercise settings. | 43891 | 2020
GXRIEIX5 | 10.1002/adma.202209300 | https://www.ncbi.nlm.nih.gov/pubmed/36576895 | YES | electronic skin (E-skin) platform |  | NO | monitoring multiple physiological and metabolic biomarkers. | electrochemical | Electrochemical | NA |  | NA | yes | - | yes | sweat | interstitial fluid |  |  |  | yes | yes | [6 empty] | Glucose, Lactate, alcohol |  |  |  |  | 1 | 2 | Prototype/Validation | Human | In-vitro + human studies | NA | NA | A biocompatible and biodegradable E-skin patch based on flexible gelatin methacryloyl aerogel (FGA) for non-invasive and continuous monitoring of multiple biomarkers of interest is engineered and demonstrated. Taking advantage of cryogenic temperature treatment and slow polymerization, FGA is fabricated with a highly interconnected porous structure that displays good flexibility, passive-cooling capabilities, and ultra-lightweight properties that make it comfortable to wear for long periods of time. | NO | NA | NA | NA | 2023 | Zhu, Y.; Haghniaz, R.; Hartel, M. C.; Guan, S.; Bahari, J.; Li, Z.; Baidya, A.; Cao, K.; Gao, X.; Li, J.; Wu, Z.; Cheng, X.; Li, B.; Emaminejad, S.; Weiss, P. S.; Khademhosseini, A. | A Breathable, Passive-Cooling, Non-Inflammatory, and Biodegradable Aerogel Electronic Skin for Wearable Physical-Electrophysiological-Chemical Analysis | Adv Mater |  | 1521-4095 (Electronic) 0935-9648 (Print) 0935-9648 (Linking) | Real-time monitoring of human health can be significantly improved by designing novel electronic skin (E-skin) platforms that mimic the characteristics and sensitivity of human skin. A high-quality E-skin platform that can simultaneously monitor multiple physiological and metabolic biomarkers without introducing skin discomfort or irritation is an unmet medical need. Conventional E-skins are either monofunctional or made from elastomeric films that do not include key synergistic features of natural skin, such as multi-sensing, breathability, and thermal management capabilities in a single patch. Herein, a biocompatible and biodegradable E-skin patch based on flexible gelatin methacryloyl aerogel (FGA) for non-invasive and continuous monitoring of multiple biomarkers of interest is engineered and demonstrated. Taking advantage of cryogenic temperature treatment and slow polymerization, FGA is fabricated with a highly interconnected porous structure that displays good flexibility, passive-cooling capabilities, and ultra-lightweight properties that make it comfortable to wear for long periods of time. It also provides numerous permeable capillary channels for thermal-moisture transfer, ensuring its excellent breathability. Therefore, the engineered FGA-based E-skin can simultaneously monitor body temperature, hydration, and biopotentials via electrophysiological sensors and detect glucose, lactate, and alcohol levels via electrochemical sensors. This work offers a previously unexplored materials strategy for next-generation E-skin platforms with superior practicality. | 44986 | 2023
7XVRRYBE | 10.1002/adma.202102740 | https://onlinelibrary.wiley.com/doi/10.1002/adma.202102740 | YES | skin-attached |  | NO | efficient collection and transport of sweat to a sensor for analysis. | Other | Cactus-spine-inspired, wedge-shaped wettability-patterned channel to collect sweat and transport it to a sweat reservoir where it can be analyzed by a sensor. | The lactate and glucose sensors used in the wearable patch have a highly linear concentration dependence, meaning the sensor output is directly proportional to the concentration of the analyte. The source reports an R2 value of 0.99 for both sensors, indicating a strong correlation between the sensor response and the actual concentration of lactate and glucose. The source doesn't provide specific sensitivity values for the sensors, such as the limit of detection or the minimum detectable concentration. However, it does state that the sensors have reliable sensing performance |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | glucose, lactate |  |  |  |  | 1 | 2 | prototype | Human | laboratory testing + in-vitro tests + human studies (preliminary) | NA | NA | Wedge-Shaped Wettability-Patterned Channels: these microfluidic channels mimic cactus spines, using a unidirectional Laplace pressure gradient to rapidly transport sweat to sensors, even against gravity. Superhydrophobic/Superhydrophilic Surface Patterning: hierarchical surface engineering amplifies sweat transport efficiency through precise spatial control of surface energy. Enclosed Channel Design: the sealed system prevents evaporation and ensures consistent performance regardless of body orientation. | NO | NA | NA | NA | 2021 | Son, Jonghyun; Bae, Geun Yeol; Lee, Siyoung; Lee, Giwon; Kim, Seong Won; Kim, Daegun; Chung, Sein; Cho, Kilwon | Cactus‐Spine‐Inspired Sweat‐Collecting Patch for Fast and Continuous Monitoring of Sweat | Advanced Materials |  | 0935-9648, 1521-4095 | Abstract A sweat sensor is expected to be the most appropriate wearable device for noninvasive healthcare monitoring. However, the practical use of sweat sensors is impeded by irregular and low sweat secretion rates. Here, a sweat‐collecting patch that can collect sweat efficiently for fast and continuous healthcare monitoring is demonstrated. The patch uses cactus‐spine‐inspired wedge‐shaped wettability‐patterned channels on a hierarchical microstructured/nanostructured surface. The channel shape, in combination with the superhydrophobic/superhydrophilic surface materials, induces a unidirectional Laplace pressure that transports the sweat to the sensing area spontaneously even when the patch is aligned vertically. The patch demonstrates superior sweat‐collecting efficiency and reduces the time required to fill the sensing area by transporting sweat almost without leaving it inside the channel. Therefore, a sensor based on the patch responds quickly to biochemicals in sweat, and the patch enables the continuous monitoring of changes in sweat biochemicals according to their changes in the wearer's blood. | 44470 | 2021
7NI55BGL | 10.1021/acsami.8b17419 | https://www.ncbi.nlm.nih.gov/pubmed/30516364 | YES | not specified |  | NO | sweat monitoring | electrochemical | Electrochemical | detection limit of 6 μM for H2O2 sensing. |  | NA | yes | - | - | sweat | [12 empty] | H2O2 |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | Graphene films as working electrodes and metal pins as counter electrodes. | NO | NA | NA | NA | 2019 | Li, G.; Mo, X.; Law, W. C.; Chan, K. C. | Wearable Fluid Capture Devices for Electrochemical Sensing of Sweat | ACS Appl Mater Interfaces |  | 1944-8252 (Electronic) 1944-8244 (Linking) | Wearable sensing technologies are vital for realizing personalized health monitoring. Noninvasive human sweat sampling is essential for monitoring an individual's physical state using rich physiological data. However, existing wearable sensing technologies lack the controlled capture of body sweat and in performing on-device measurement without inflammatory contact. Herein, we report the development of a wearable sweat-capture device using patterned graphene arrays with controlled superwettability and electrical conductivity for simultaneously capturing and electrochemically measuring sweat droplets. The sweat droplets exhibited strong attachment on the superhydrophilic graphene patterns, even during moderate exercising. The captured sweat droplets present strong electrochemical signals using graphene films as the working electrode and metal pins as the counter electrode arrays assembled on 3D printed holders, at the detection limit of 6 muM for H(2)O(2) sensing. This research enables full-body spatiotemporal mapping of sweat, which is beneficial for a broad range of personalized monitoring applications, such as drug abuse detection, athletics performance optimization, and physiological wellness tracking. | 43474 | 2019
3G3K5PD7 | 10.1021/acssensors.9b00891 | https://www.ncbi.nlm.nih.gov/pubmed/31271034 | YES | wrist-worn | smartwatch-like | NO | monitoring of sweat glucose levels | electrochemical | Electrochemical techniques | Reproducibility of the sensors shows an average sensitivity of 3.29 nA μM−1, with a relative standard deviation (RSD) of 8.8%. Sensors maintained a sensitivity of 3.18 nA μM−1 after 2 h long-term measurements, which is within a variation of 3.3% compared to the average sensitivities of the as-fabricated sensors. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | Self-powered and fully integrated smartwatch that consists of flexible photovoltaic cells and rechargeable batteries in the forms of a "watch strap", electrochemical glucose sensors, customized circuits, and display units integrated into a "dial" platform is successfully fabricated for real-time and continuous monitoring of sweat glucose levels | NO | NA | NA | NA | 2019 | Zhao, J.; Lin, Y.; Wu, J.; Nyein, H. Y. Y.; Bariya, M.; Tai, L. C.; Chao, M.; Ji, W.; Zhang, G.; Fan, Z.; Javey, A. | A Fully Integrated and Self-Powered Smartwatch for Continuous Sweat Glucose Monitoring | ACS Sens |  | 2379-3694 (Electronic) 2379-3694 (Linking) | Wearable devices for health monitoring and fitness management have foreseen a rapidly expanding market, especially those for noninvasive and continuous measurements with real-time display that provide practical convenience and eliminated safety/infection risks. Herein, a self-powered and fully integrated smartwatch that consists of flexible photovoltaic cells and rechargeable batteries in the forms of a "watch strap", electrochemical glucose sensors, customized circuits, and display units integrated into a "dial" platform is successfully fabricated for real-time and continuous monitoring of sweat glucose levels. The functionality of the smartwatch, including sweat glucose sensing, signal processing, and display, can be supported with the harvested/converted solar energy without external charging devices. The Zn-MnO(2) batteries serve as intermediate energy storage units and the utilization of aqueous electrolytes eliminated safety concerns for batteries, which is critical for wearable devices. Such a wearable system in a smartwatch fashion realizes integration of energy modules with self-powered capability, electrochemical sensors for noninvasive glucose monitoring, and in situ and real-time signal processing/display in a single platform for the first time. The as-fabricated fully integrated and self-powered smartwatch also provides a promising protocol for statistical study and clinical investigation to reveal correlations between sweat compositions and human body dynamics. | 43672 | 2019
AGN4RZSZ | 10.1002/adhm.202302173 | https://onlinelibrary.wiley.com/doi/10.1002/adhm.202302173 | YES | skin-attached |  | NO | colorimetric analysis of sweat biomarkers, enabling non-invasive monitoring of various health parameters | Colorimetric | Colorimetric | These tests demonstrate the patch's accuracy in recovering known concentrations of metabolites. The reported recovery rates range from 99.4% to 103.7% for pH, glucose, lactate, and uric acid. This indicates a high level of precision in measuring these biomarkers. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes |  | yes |  |  | yes |  | pH, Glucose, Lactate, Uric Acid |  |  |  |  | 1 | 4 | Validation | Human | laboratory testing + human studies | NA | NA | Ceramic-Ink Microfluidic Channels: provides higher resolution (450 μm vs. 850 μm), enhanced thermal stability, leakage prevention, and resistance to organic solvents, improving durability and precision. Nanoparticle-Enhanced Detection Zones: incorporates SiO2, ZnO, and Al2O3 nanoparticles to improve color intensity, sensitivity, and dynamic range for colorimetric analysis. Moisture-Sensitive Channel: uses Co²⁺ ions for visual sweat rate tracking with color transitions from blue to pink, enabling quantitative measurements via image analysis. Multiplexed Sensor Integration: combines pH, glucose, lactate, and uric acid detection in one device for comprehensive health assessment. | NO | NA | NA | NA | 2024 | Zheng, Xin Ting; Goh, Wei Peng; Yu, Yong; Sutarlie, Laura; Chen, Der Ying; Tan, Sherwin Chong Li; Jiang, Changyun; Zhao, Meng; Ba, Te; Li, Hongying; Su, Xiaodi; Yang, Le | Skin‐Attachable Ink‐Dispenser‐Printed Paper Fluidic Sensor Patch for Colorimetric Sweat Analysis | Advanced Healthcare Materials |  | 2192-2640, 2192-2659 | Abstract In situ analysis of sweat biomarkers potentially provides noninvasive lifestyle monitoring and early diagnosis. Quantitative detection of sweat rate is crucial for thermoregulation and preventing heat injuries. Here, a skin‐attachable paper fluidic patch is reported for in situ colorimetric sensing of multiple sweat markers (pH, glucose, lactate, and uric acid) with concurrent sweat rate tracking. Two sets of fluidic patterns—multiplexed detection zones and a longitudinal sweat rate channel—are directly printed by an automated ink dispenser from a specially developed ceramic‐based ink. The ceramic ink thermal‐cures into an impervious barrier, confining sweat within the channels. The ceramic‐ink‐printed boundary achieves higher pattern resolution, prevents fluid leakage, attains pattern thermal stability, and resistant to organic solvents. The cellulose matrix of the detection zones is modified with nanoparticles to improve the color homogeneity and sweat sensor sensitivity. The sweat rate channel is made moisture sensitive by incorporating a metal‐salt‐based dye. The change in saturation/color of the detection zones and/or channels upon sweat addition can be visually detected or quantified by a smartphone camera. A cost‐effective way is provided to fabricate paper fluidic sensor patches, successfully demonstrating on‐body multiplexed evaluation of sweat analytes. Such skin wearables offer on‐site analysis, meaningful to an increasingly health‐conscious population. | 45292 | 2024
I5C5IKGR | 10.1021/acssensors.0c00791 | https://www.ncbi.nlm.nih.gov/pubmed/32429661 | YES | sweat band |  | NO | monitoring nicotine level | electrochemical | Electrochemical detection. | The sensors’ sensitivities to nicotine (10 sensors) show these values: The mean, median, and standard deviation are 4.4, 4.6, and 1.4 nA/μM, respectively. |  | NA | yes | - | - | sweat | [12 empty] | nicotine |  |  |  |  | 1 | 1 | Prototype/Validation | Human | Human studies | NA | NA | The development of a wearable sweatband itself represents a significant advancement in health monitoring technology. | NO | NA | NA | NA | 2020 | Tai, L. C.; Ahn, C. H.; Nyein, H. Y. Y.; Ji, W.; Bariya, M.; Lin, Y.; Li, L.; Javey, A. | Nicotine Monitoring with a Wearable Sweat Band | ACS Sens |  | 2379-3694 (Electronic) 2379-3694 (Linking) | The tobacco epidemic is a public health threat that has taken a heavy toll of lives around the globe each year. Smoking affects both the smokers and those who are exposed to secondhand smoke, and careful tracking of exposure can be key to mitigating the potential hazards. For smokers, the variation of chemical compositions between commercial cigarettes has led to ambiguity in estimating the health risks, both for active smokers and others involuntarily exposed to tobacco smoke and byproducts. In this regard, sweat possesses an attractive opportunity to monitor smoke exposure due to sweat's abundance in biomolecules and its great accessibility. Here, we present a wearable sweat band to monitor nicotine, a prominent ingredient in cigarettes, as a viable way to quantitatively assess a wearer's exposure to smoking. Both smokers and normal subjects are tested to demonstrate the use of this device for smoke-related health monitoring. Our results exhibit confirmable and elevated nicotine levels in sweat for subjects inhaling cigarette smoke. This continuous and personalized sweat sensing device is leverage to monitor smoke pollution for a potentially broad population. | 44008 | 2020
NCRESJCU | 10.1021/acs.nanolett.9b02478 | https://www.ncbi.nlm.nih.gov/pubmed/31381353 | YES | wrist-worn | band | NO | monitoring levodopa dynamics in the body | Other | Iontophoresis | The inset displays the calibration curve with a sensitivity of 15 nA/μM, which is on par with the best levodopa sensors reported. This is notable considering the simplicity of electrode functionalization and the electrode’s excellent stability for long-term usage | Neuro | Parkinson’s disease | yes | - | - | sweat | [12 empty] | Levodopa |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | Nanodendritic Platform | NO | NA | NA | NA | 2019 | Tai, L. C.; Liaw, T. S.; Lin, Y.; Nyein, H. Y. Y.; Bariya, M.; Ji, W.; Hettick, M.; Zhao, C.; Zhao, J.; Hou, L.; Yuan, Z.; Fan, Z.; Javey, A. | Wearable Sweat Band for Noninvasive Levodopa Monitoring | Nano Lett |  | 1530-6992 (Electronic) 1530-6984 (Linking) | Levodopa is the standard medication clinically prescribed to patients afflicted with Parkinson's disease. In particular, the monitoring and optimization of levodopa dosage are critical to mitigate the onset of undesired fluctuations in the patients' physical and emotional conditions such as speech function, motor behavior, and mood stability. The traditional approach to optimize levodopa dosage involves evaluating the subjects' motor function, which has many shortcomings due to its subjective and limited quantifiable nature. Here, we present a wearable sweat band on a nanodendritic platform that quantitatively monitors levodopa dynamics in the body. Both stationary iontophoretic induction and physical exercise are utilized as our methods of sweat extraction. The sweat band measures real-time pharmacokinetic profiles of levodopa to track the dynamic response of the drug metabolism. We demonstrated the sweat band's functionalities on multiple subjects with implications toward the systematic administering of levodopa and routine management of Parkinson's disease. | 43719 | 2019
82MHTNK6 | 10.1007/s00216-024-05135-w | https://link.springer.com/10.1007/s00216-024-05135-w | YES | skin-attached |  | NO | sweat glucose detection | electrochemical | Electrochemical | The wearable biosensor exhibits a sensitivity of 14.64 µA mM−1 cm−2 and a limit of detection (LOD) of 4.83 µM. The LOD is a measure of the smallest amount of glucose the sensor can reliably detect. The researchers tested the biosensor's performance in both a controlled buffer solution (PBS) and artificial sweat, achieving good linearity and accuracy in both scenarios1. The linear range of the sensor, meaning the range of glucose concentrations where the output signal is directly proportional to the concentration, spans from 5 µM to 1250 µM. To assess real-world accuracy, the researchers compared the biosensor's measurements of glucose levels in human sweat to those obtained using a commercial glucose test kit. The results showed a close correlation between the two methods, with the biosensor accurately detecting glucose levels both before and after meals. | Diabetes | diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation | Human | laboratory testing + human studies | NA | NA | Cost-Effective and Scalable: CBNPs are less expensive and simpler to synthesize compared to graphene or nanotubes, enabling practical large-scale production. Superior Conductivity and Stability: high electrical conductivity and excellent solvent dispersibility ensure uniform, stable electrode coatings for efficient electron transfer. Enhanced Surface Area and Sensitivity: their large surface area and electrocatalytic properties improve enzyme immobilization, boosting interaction with glucose and achieving higher sensitivity. Performance Boost: experiments confirm reduced electron transfer resistance, heightened sensitivity, and a low detection limit (4.83 µM), enabling accurate sweat glucose monitoring. | NO | NA | NA | NA | 2024 | Tang, Chaoli; Zhou, Kai; Wang, Rujing; Li, Mengya; Liu, Wenlong; Li, Chengpan; Chen, Xiangyu; Lu, Qinwen; Chang, Yongjia | Wearable biosensors for human sweat glucose detection based on carbon black nanoparticles | Analytical and Bioanalytical Chemistry |  | 1618-2642, 1618-2650 | Wearable glucose biosensors enable noninvasive glucose monitoring, thereby enhancing blood glucose management. In this work, we present a wearable biosensor based on carbon black nanoparticles (CBNPs) for glucose detection in human sweat. The biosensor consists of CBNPs, Prussian blue (PB), glucose oxidase, chitosan, and Nafion. The fabricated biosensor has a linear range of 5 µM to 1250 µM, sensitivity of 14.64 µA mM−1 cm−2, and a low detection potential (−0.05 V, vs. Ag/ AgCl). The detection limit for glucose was calculated as 4.83 µM. This reusable biosensor has good selectivity and stability and exhibits a good response to glucose in real sweat. These results demonstrate the potential of our CBNP-based biosensor for monitoring blood glucose in human sweat. | 45352 | 2024
273QL2UT | 10.1016/j.talanta.2020.121153 | https://linkinghub.elsevier.com/retrieve/pii/S0039914020304446 | YES | skin-attached |  | NO |  | electrochemical | Electrochemical | The LOD values achieved using the wearable sensor are: Cd2+: 1.01 μM Pb2+: 0.43 μM Uric acid: 1.8 μM (using anodically pre-treated electrodes) 17β-estradiol: 0.58 μM (using anodically pre-treated electrodes) |  | NA | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | Cadmium (Cd2+), Lead (Pb2+), Uric acid, 17β-estradiol |  |  |  |  | 1 | 4 | proof-of-concept/exploratory |  | in-vitro tests (with artificial sweat) | NA | NA | This wearable sensor leverages microbial nanocellulose (MNC) as a breakthrough substrate, offering innovative features: Biocompatibility and Comfort: MNC adheres to skin without adhesives and minimizes irritation, ensuring a comfortable, skin-friendly design. Breathability and Sweat Management: its porous structure enables gas and liquid permeability, preventing sweat accumulation. Sustainable and Durable: derived from bacterial fermentation, MNC is renewable, biodegradable, and highly durable despite its flexibility. Advanced Detection: screen-printed carbon electrodes (SPCEs) on MNC enable detection of toxic metals, uric acid, and hormones, creating lightweight and disposable sensors. Enhanced Performance: electrochemical pre-treatment of SPCEs boosts conductivity, sensitivity, and lowers detection limits, elevating device performance. Future-Ready Design: the combination of MNC and optimized SPCEs establishes a path toward sustainable, high-performance wearable biosensors. | NO | NA | NO | NA | 2020 | Silva, Robson R.; Raymundo-Pereira, Paulo A.; Campos, Anderson M.; Wilson, Deivy; Otoni, Caio G.; Barud, Hernane S.; Costa, Carlos A.R.; Domeneguetti, Rafael R.; Balogh, Debora T.; Ribeiro, Sidney J.L.; Oliveira Jr., Osvaldo N. | Microbial nanocellulose adherent to human skin used in electrochemical sensors to detect metal ions and biomarkers in sweat | Talanta |  | 399140 | The pursuit of biocompatible, breathable and skin-conformable wearable sensors has predominantly focused on synthetic stretchable hydrophobic polymers. Microbial nanocellulose (MNC) is an exceptional skin-substitute natural polymer routinely used for wound dressing and offers unprecedented potential as substrate for wearable sensors. A versatile strategy for engineering wearable sensing platforms is reported, with sensing units made of screen-printed carbon electrodes (SPCEs) on MNC. As-prepared SPCEs were used to detect the toxic metals cadmium (Cd2+) and lead (Pb2+) with limits of detection of 1.01 and 0.43 μM, respectively, which are sufficient to detect these metal ions in human sweat and urine. SPCEs functionalized through anodic pre-treatments were used for detecting uric acid and 17β-estradiol in artificial sweat, with detection limits of 1.8 μM and 0.58 μM, respectively. The electrochemical treatment created oxygen groups on the carbon surfaces, thus improving wettability and hydrophilicity. MNC was herein exploited as an adhesive-free, yet highly skin-adherent platform for wearable sensing devices that also benefit from the semi-permeable, non-allergenic, and renewable features that make MNC unique within the pool of materials that have been used for such a purpose. Our findings have clear implications for the developments on greener and more biocompatible but still efficient substrates and may pave the route for combining immunosensing devices with drug delivery therapies. | 44105 | 2020
IK247MYV | 10.1016/j.talanta.2020.121145 | https://www.ncbi.nlm.nih.gov/pubmed/32887090 | YES | wrist-worn | smartwatch-like | NO | monitoring of electrolyte concentrations | electrochemical | Electrochemical (solid-contact ion-selective electrodes (ISEs)) | The resulting electrodes showed good sensitivity (52.4 ± 6.3 mV/decade (PEDOT) and 56.4 ± 2.2 mV/decade (POT) for Na+ ISEs, and 45.7 ± 7.4 mV/decade (PEDOT) and 54.3 ± 1.5 mV/ decade (POT) for K+) and excellent selectivity towards potential interferents present in human sweat (H+, Na+, K+, Mg2+, Ca2+). |  | NA | yes | - | - | sweat | [12 empty] | Na+, K+. |  |  |  |  | 1 | 2 | Prototype/Validation | Human | Human studies | NA | NA | SwEatch device incorporates several cutting-edge innovations, including solid-contact ISEs, conductive polymer transducers, a 3D-printed platform with a mirrored fluidic unit, integrated electronics with wireless data transmission, and automated fabrication processes. | NO | NA | YES | NA | 2020 | Pirovano, P.; Dorrian, M.; Shinde, A.; Donohoe, A.; Brady, A. J.; Moyna, N. M.; Wallace, G.; Diamond, D.; McCaul, M. | A wearable sensor for the detection of sodium and potassium in human sweat during exercise | Talanta |  | 1873-3573 (Electronic) 0039-9140 (Linking) | The SwEatch platform, a wearable sensor for sampling and measuring the concentration of electrolytes in human sweat in real time, has been improved in order to allow the sensing of two analytes. The solid contact ion-sensitive electrodes (ISEs) for the detection of Na(+) and K(+) have been developed in two alternative formulations, containing either poly(3,4-ethylenedioxythiophene) (PEDOT) or poly(3-octylthiophene-2,5-diyl) (POT) as a conductive polymer transducing component. The solution-processable POT formulation simplifies the fabrication process, and sensor to sensor reproducibility has been improved via partial automation using an Opentron(R) automated pipetting robot. The resulting electrodes showed good sensitivity (52.4 +/- 6.3 mV/decade (PEDOT) and 56.4 +/- 2.2 mV/decade (POT) for Na(+) ISEs, and 45.7 +/- 7.4 mV/decade (PEDOT) and 54.3 +/- 1.5 mV/decade (POT) for K(+)) and excellent selectivity towards potential interferents present in human sweat (H(+), Na(+), K(+), Mg(2+), Ca(2+)). The 3D printed SwEatch platform has been redesigned to incorporate a double, mirrored fluidic unit which is capable of drawing sweat from the skin through passive capillary action and bring it in contact with two independent electrodes. The potentiometric signal generated by the electrodes is measured by an integrated electronics board, digitised and transmitted via Bluetooth to a laptop. The results obtained from on-body trials on athletes during cycling show a relatively small increase in sodium (1.89 mM-2.97 mM) and potassium (3.31 mM-7.25 mM) concentrations during the exercise period of up to 90 min. | 44136 | 2020
3YKGTKMD | 10.1016/j.bios.2020.112844 | https://linkinghub.elsevier.com/retrieve/pii/S0956566320308307 | YES | skin-attached |  | NO | real-time, on-site monitoring of potassium ion (K+) concentration in human sweat | electrochemical | Electrochemical | High sensitivity: the fabricated sensor demonstrated a high sensitivity of 63 mV/dec, further amplified to 173 mV/dec using the integrated NFC amplification system. Excellent selectivity: the sensor exhibited strong selectivity for K+ ions, with minimal interference from other ions commonly found in sweat (Ca2+, Na+, Zn2+). Good repeatability and reproducibility: multiple tests demonstrated consistent performance across different sensors and repeated measurements. Validation against a reference method: the accuracy of on-body measurements was validated by comparing the sensor's readings with those obtained using a standard phosphate-buffered saline (PBS) solution with known K+ concentrations. Additionally, the feasibility was further confirmed using the standard addition method with diluted human perspiration samples |  | NA | yes | - | - | sweat | [12 empty] | K+ |  |  |  |  | 1 | 1 | Validation | Human | laboratory testing + human studies | NA | NA | Hybrid MXene-Ti₃C₂Tₓ/MWCNT Electrode: combines high-conductivity MXene with MWCNTs for a stable, noise-resistant electrode with enhanced sensitivity. Flexible Substrate with Integrated Microfluidics: PET substrate and PDMS microfluidics enable efficient sweat collection and contamination prevention, ensuring real-time K⁺ monitoring. Battery-Free NFC Operation: uses smartphone-powered NFC for wireless data transmission, eliminating batteries and enhancing wearability. Integrated Signal Amplification: A programmable gain amplifier boosts sensitivity from 63 mV/dec to 173 mV/dec for precise K⁺ detection. | NO | NA | YES | NA | 2021 | Zhang, Shipeng; Zahed, Md Abu; Sharifuzzaman, Md.; Yoon, Sanghyuk; Hui, Xue; Chandra Barman, Sharat; Sharma, Sudeep; Yoon, Hyo Sang; Park, Chani; Park, Jae Yeong | A wearable battery-free wireless and skin-interfaced microfluidics integrated electrochemical sensing patch for on-site biomarkers monitoring in human perspiration | Biosensors and Bioelectronics |  | 9565663 | In this study, an ultra-high sensitive, flexible, wireless, battery-free, and fully integrated (no external analysis equipment) electrochemical sensing patch system, including a microfluidic-sweat collecting unit, was newly developed for the on-site monitoring of the [K+] concentration in human sweat. Multiwalled carbon nanotube (MWCNT) and MXene-Ti3C2TX based hybrid multi-dimensional networks were applied to obtain a high surface activation area and faster charge transfer rate, strongly adsorbing the valinomycin membrane to protect the ionophore for effective transshipment and immobilization of the [K+]. Furthermore, the controllable porosity of carbon-based materials can accelerate the kinetic process of ion diffusion. This hybrid nanonetwork structure effectively enhanced electrochemical stability and sensitivity, addressing the noise and signal drifting problems experienced with low concentration detection. The fabricated sensor exhibited a high ion concentration sensitivity of 63 mV/dec with excellent selectivity, amplified to 173 mV/dec with the integrated amplification system. The Near Field Communication (NFC) is used to transmit measurements to a smartphone wirelessly. A microfluidic channel was integrated with the electrochemical sensor patch to efficiently collect sweat on the human skin surface and mitigate the sensor surface contamination problem. Furthermore, the developed sensing patch can also be applied to other biomarkers on-site detection after modifying the working electrode with the corresponding selective membranes. | 44256 | 2021
3LKATQNE | 10.1016/j.bios.2021.113606 | https://www.ncbi.nlm.nih.gov/pubmed/34507206 | YES | skin-interfaced |  | NO | non-enzimatic glucose detection | electrochemical | Electrochemical | The sensor with the porous LIG foam exhibits a high sensitivity of 1080 μA mM1 cm2, whereas a further increased sensitivity of 3500 μA mM1 cm2 is obtained with LIG fibers (LIGF). Impressively, a large linear range (0–30 mM) can be achieved by changing the bias voltage from 0.5 to 0.1 V due to the Au coating. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation |  | In-vitro tests | NA | NA | 3D porous LIG foams or fibers on a f lexible, thin-film substrate and electroless plating of Ni and Au for high- performance non-enzymatic glucose sensors | NO | NA | NA | NA | 2021 | Zhu, J.; Liu, S.; Hu, Z.; Zhang, X.; Yi, N.; Tang, K.; Dexheimer, M. G.; Lian, X.; Wang, Q.; Yang, J.; Gray, J.; Cheng, H. | Laser-induced graphene non-enzymatic glucose sensors for on-body measurements | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Print) 0956-5663 (Linking) | Non-enzymatic glucose sensors outperform enzymatic ones in terms of cost, sensitivity, stability, and operating duration. Though highly sensitive, it is still desirable to further improve the sensitivity of non-enzymatic glucose sensors to detect a trace amount of glucose in sweat and other biofluids. Among the demonstrated effective approaches using bimetals or 3D porous structures, the porous laser-induced graphene (LIG) on flexible polymers showcases good conductivity and a simple fabrication process for the integration of sensing materials. The uniform electroless plating of the nickel and gold layer on LIG electrodes demonstrates significantly enhanced sensitivity and a large linear range for glucose sensing. The sensor with the porous LIG foam exhibits a high sensitivity of 1080 muA mM(-1) cm(-2), whereas a further increased sensitivity of 3500 muA mM(-1) cm(-2) is obtained with LIG fibers (LIGF). Impressively, a large linear range (0-30 mM) can be achieved by changing the bias voltage from 0.5 to 0.1 V due to the Au coating. Because the existing non-enzymatic glucose sensors are limited to use in basic solutions, their application in wearable electronics is elusive. In addition to the reduced requirement for the basic solution, this work integrates a porous encapsulating reaction cavity containing alkali solutions with a soft, skin-interfaced microfluidic component to provide integrated microfluidic non-enzymatic glucose sensors for sweat sampling and glucose sensing. The accurate glucose measurements from the human sweat and cell culture media showcase the practical utility, which opens up opportunities for the non-enzymatic glucose sensors in wearable electronics. | 44531 | 2021
LCQAV858 | 10.3390/bios12060397 | https://www.ncbi.nlm.nih.gov/pubmed/35735545 | YES | not specified | Epidermal | NO | monitoring sweat | electrochemical | Electrochemical | The resulting LIG devices can simultaneously monitor the pH, Na+, and K+ levels in sweat with the sensitivities of 51.5 mV/decade (pH), 45.4 mV/decade (Na+), and 43.3 mV/decade (K+), respectively. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | pH, Na+, K+ |  |  |  |  | 1 | 3 | Proof-of-concept | Human | In-vitro + human studies | NA | NA | The device employs a sophisticated electrochemical detection method based on a liquid-gated transistor architecture using rGO as the sensing material and ion-selective membranes for specificity. | NO | NA | NA | NA | 2022 | Liao, J.; Zhang, X.; Sun, Z.; Chen, H.; Fu, J.; Si, H.; Ge, C.; Lin, S. | Laser-Induced Graphene-Based Wearable Epidermal Ion-Selective Sensors for Noninvasive Multiplexed Sweat Analysis | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | Wearable sweat sensors are a rapidly rising research area owing to their convenience for personal healthcare and disease diagnosis in a real-time and noninvasive manner. However, the fast and scalable fabrication of flexible electrodes remains a major challenge. Here, we develop a wearable epidermal sensor for multiplexed sweat analysis based on the laser-induced graphene (LIG) technique. This simple and mask-free technique allows the direct manufacturing of graphene electrode patterns on commercial polyimide foils. The resulting LIG devices can simultaneously monitor the pH, Na(+), and K(+) levels in sweat with the sensitivities of 51.5 mV/decade (pH), 45.4 mV/decade (Na(+)), and 43.3 mV/decade (K(+)), respectively. Good reproducibility, stability, and selectivity are also observed. On-body testing of the LIG-based sensor integrated with a flexible printed circuit board during stationary cycling demonstrates its capability for real-time sweat analysis. The concentrations of ions can be remotely and wirelessly transmitted to a custom-developed smartphone application during the period in which the sensor user performs physical activities. Owing to the unique advantages of LIG technique, including facile fabrication, mass production, and versatile, more physiological signals (glucose, uric acid, tyrosine, etc.) could be easily expanded into the LIG-based wearable sensors to reflect the health status or clinical needs of individuals. | 44721 | 2022
VBL3EPW6 | 10.3390/s21051847 | https://www.ncbi.nlm.nih.gov/pubmed/33800949 | YES | foot-worn |  | NO | monitoring and managing the health of the diabetic foot | electrochemical | Electrochemical (+ Piezoelectric Pressure Sensing and Temperature Sensing) | The disposable glucose sensors exhibit a 0–6 mM (0–108 mg/dL) glucose concentration optimized linear response (for sweat-sensing), with a LOD of 3.27 µM (0.058 mg/dL) and a sensitivity of 21 µA/mM cm2 in the PBS solution. | Diabetes | Diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | Glucose |  |  |  |  | 1 | 1 | Prototype |  | Laboratory testing | NA | NA | A smart multisensory hardware-software monitoring platform for diabetic patients was presented, in which an antibacterial polyurethane-based insole was proposed as a reliable platform for integration of an array of minimally invasive and low power commercial and proprietary sensors distributed on the plantar surface. | NO | NA | YES | NA | 2021 | De Pascali, C.; Francioso, L.; Giampetruzzi, L.; Rescio, G.; Signore, M. A.; Leone, A.; Siciliano, P. | Modeling, Fabrication and Integration of Wearable Smart Sensors in a Monitoring Platform for Diabetic Patients | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | The monitoring of some parameters, such as pressure loads, temperature, and glucose level in sweat on the plantar surface, is one of the most promising approaches for evaluating the health state of the diabetic foot and for preventing the onset of inflammatory events later degenerating in ulcerative lesions. This work presents the results of sensors microfabrication, experimental characterization and FEA-based thermal analysis of a 3D foot-insole model, aimed to advance in the development of a fully custom smart multisensory hardware-software monitoring platform for the diabetic foot. In this system, the simultaneous detection of temperature-, pressure- and sweat-based glucose level by means of full custom microfabricated sensors distributed on eight reading points of a smart insole will be possible, and the unit for data acquisition and wireless transmission will be fully integrated into the platform. Finite element analysis simulations, based on an accurate bioheat transfer model of the metabolic response of the foot tissue, demonstrated that subcutaneous inflamed lesions located up to the muscle layer, and ischemic damage located not below the reticular/fat layer, can be successfully detected. The microfabrication processes and preliminary results of functional characterization of flexible piezoelectric pressure sensors and glucose sensors are presented. Full custom pressure sensors generate an electric charge in the range 0-20 pC, proportional to the applied load in the range 0-4 N, with a figure of merit of 4.7 +/- 1 GPa. The disposable glucose sensors exhibit a 0-6 mM (0-108 mg/dL) glucose concentration optimized linear response (for sweat-sensing), with a LOD of 3.27 microM (0.058 mg/dL) and a sensitivity of 21 microA/mM cm(2) in the PBS solution. The technical prerequisites and experimental sensing performances were assessed, as preliminary step before future integration into a second prototype, based on a full custom smart insole with enhanced sensing functionalities. | 44261 | 2021
VIUIBGAX | 10.1038/srep23111 | https://www.ncbi.nlm.nih.gov/pubmed/26996103 | YES | skin-attached |  | NO | monitoring alcohol consumption | Immuneassay | Affinity based immunoassay | EtG detection over a dose concentration of 0.001–100 μg/L was demonstrated on both glass and polyimide substrates. Detection sensitivity was lower at 1 μg/L with gold electrodes as compared to ZnO, which had detection sensitivity of 0.001 μg/L. |  | NA | yes | - | - | sweat | [12 empty] | ethyl glucuronide (EtG) |  |  |  |  | 1 | 1 | Proof-of-concept | Human | Human studies | NA | NA | Two co-planar sensors with gold and zinc oxide as sensing electrodes + a LED based reporting | NO | NA | NA | NA | 2016 | Selvam, A. P.; Muthukumar, S.; Kamakoti, V.; Prasad, S. | A wearable biochemical sensor for monitoring alcohol consumption lifestyle through Ethyl glucuronide (EtG) detection in human sweat | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | We demonstrate for the first time a wearable biochemical sensor for monitoring alcohol consumption through the detection and quantification of a metabolite of ethanol, ethyl glucuronide (EtG). We designed and fabricated two co-planar sensors with gold and zinc oxide as sensing electrodes. We also designed a LED based reporting for the presence of EtG in the human sweat samples. The sensor functions on affinity based immunoassay principles whereby monoclonal antibodies for EtG were immobilized on the electrodes using thiol based chemistry. Detection of EtG from human sweat was achieved through chemiresistive sensing mechanism. In this method, an AC voltage was applied across the two coplanar electrodes and the impedance across the sensor electrodes was measured and calibrated for physiologically relevant doses of EtG in human sweat. EtG detection over a dose concentration of 0.001-100 mug/L was demonstrated on both glass and polyimide substrates. Detection sensitivity was lower at 1 mug/L with gold electrodes as compared to ZnO, which had detection sensitivity of 0.001 mug/L. Based on the detection range the wearable sensor has the ability to detect alcohol consumption of up to 11 standard drinks in the US over a period of 4 to 9 hours. | 42450 | 2016
5FNXVCDA | 10.3390/s24144643 | https://www.ncbi.nlm.nih.gov/pubmed/39066041 | YES | garments |  | NO | continuous, non-invasive monitoring of human biomarkers in sweat | Electrical | interdigitated electrodes (IDTs)- resonant frequency, | The sensors operate at frequencies between 1.8 GHz to 2 GHz. They exhibit a sensitivity of 0.17 MHz/(mg/dL) in response to changes in glucose concentration |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Proof-of-concept |  | Laboratory testing | NA | NA | The use of embroidery techniques to create flexible, fabric-based sensors with gold-coated IDTs represents a significant innovation. The incorporation of a wireless sensing mechanism enables real-time data transmission without the need for wired connections. | NO | NA | NA | NA | 2024 | Amers, E. L.; Orme, B. V.; Shi, Y.; Torun, H.; Dodd, L. E. | Embroidered Interdigitated Electrodes (IDTs) with Wireless Readout for Continuous Biomarker Monitoring | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | Non-invasive continuous health monitoring has become feasible with the advancement of biosensors. While monitoring certain biomarkers such as heart rate or skin temperature are now at a certain maturity, monitoring molecular biomarkers is still challenging. Progress has been shown in sampling, measurement, and interpretation of data toward non-invasive molecular sensors that can be integrated into daily wearable items. Toward this goal, this paper explores the potential of embroidered interdigitated transducer (IDT)-based sensors for non-invasive, continuous monitoring of human biomarkers, particularly glucose levels, in human sweat. The study employs innovative embroidery techniques to create flexible fabric-based sensors with gold-coated IDTs. In controlled experiments, we have shown the variation of glucose concentration in water can be wirelessly detected by tracking the resonant frequency of the embroidered sensors. The current sensors operate at 1.8 GHz to 2 GHz and respond to the change in glucose concentration with a sensitivity of 0.17 MHz/(mg/dL). The embroidered IDT-based sensors with wireless sensing will be a new measurement modality for molecular wearable sensors. The establishment of a wireless sensing mechanism for embroidered IDT-based sensors will be followed by an investigation of sweat for molecular detection. This will require adding functionalities for sampling and interpretation of acquired data. We envisage the embroidered IDT-based sensors offer a unique approach for seamless integration into clothing, paving the way for personalised, continuous health data capture. | 45490 | 2024
D3J6Z5ZD | 10.1016/j.bios.2019.111637 | https://linkinghub.elsevier.com/retrieve/pii/S095656631930716X | YES | not specified | Gloves | NO | detection of hydrogen peroxide (H2O2) and glucose. | electrochemical | Electrochemical | H2O2 Sensor Sensitivity: PET-based sensor achieves 27.25 μA·mM⁻¹·cm⁻²; the glove-based version has slightly lower sensitivity due to substrate diffusion differences. Glucose Sensor Sensitivity: PET-based sensor reaches 20.22 μA·mM⁻¹·cm⁻², with the glove-based sensor showing consistent linear performance. Low Detection Limits: H2O2 sensor detects as low as 0.65 μM, while the glucose sensor detects 11.6 μM. Stability and Reproducibility: Both sensors maintain stable performance over multiple cycles and across different electrodes. Specificity: High selectivity ensures accurate detection with minimal interference from other substances. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | Hydrogen peroxide (H2O2), glucose |  |  |  |  | 1 | 2 | Prototype/Validation |  | in-vitro + laboratory testing | NA | NA | One-Step Berlin Green Printing: utilizes a unique Berlin green carbon ink for cost-effective, scalable fabrication of sensor arrays via screen-printing. Wearable and Flexible Design: Robust sensors withstand stretching and bending, enabling integration into platforms like nitrile gloves and PET substrates. Wireless Integration: Features wireless data transmission to smartphones for real-time monitoring and analysis. High Sensitivity: Demonstrates exceptional sensitivity (27.25 μA·mM⁻¹·cm⁻² for H2O2, 20.22 μA·mM⁻¹·cm⁻² for glucose) with advanced electrochemical detection. Enzyme-Based Catalysis: Uses glucose oxidase (GOD) to enhance specificity and catalytic efficiency for accurate glucose detection. | NO | NA | YES | NA | 2019 | Ma, Junlin; Jiang, Yu; Shen, Liuxue; Ma, Hongting; Sun, Tongrui; Lv, Fengjuan; Kiran, Almas; Zhu, Nan | Wearable biomolecule smartsensors based on one-step fabricated berlin green printed arrays | Biosensors and Bioelectronics |  | 9565663 | The wearable smart detection of body biomolecules and biomarkers is being of signiﬁcance in the practical ﬁelds. Hydrogen peroxide (H2O2) is a product of some enzyme-catalyzed biomolecular reactions. The detection of H2O2 could reﬂect the concentration information of the enzyme reaction biomolecule substrate such as glucose. A high-performance berlin green (BG) carbon ink for monitoring H2O2 was prepared in this work. And we have successfully developed the wearable smartsensors for detecting H2O2 and glucose based on one-step fabricated BG arrays by screen-printing technology. Comparing with other detection methods, these sensors are wearable, movable, ﬂexible and biocompatible for monitoring biomolecules. As a result, the sensors exhibited good sensitivity, speciﬁcity, stability and reproductivity towards H2O2 and glucose. Additionally, there also received stable response after near one hundred times stretching and thousands of bending. Moreover, the wearable sensors could be easily remotely controlled by a smart phone, when integrated with wireless into the device. In prospective studies, the one-step fabricated wearable smartsensors is of great signiﬁcance in developing a straightforward, highly-eﬃcient and low-cost method for actual detection of biomolecules reﬂecting body health status, and would potentially be applied in the artiﬁcial intelligence (AI) ﬁelds. | 43770 | 2019
M4L3IETC | 10.1126/scitranslmed.aaf2593 | https://www.ncbi.nlm.nih.gov/pubmed/27881826 | YES | skin-attached |  | NO | evaluating athletic performance and monitoring health and disease status (+ sweat volume and rate ). | Colorimetric | Colorimetric | Sensitivity in the reported devices could diagnose hyperglycemia, for example, limit of detection (LOD) = ~200 μM. The device utilizes colorimetric analysis for markers like chloride, hydronium ions, glucose, and lactate. Colorimetric methods rely on color changes to indicate the presence and concentration of analytes. The accuracy of these methods depends on factors like the sensitivity of the color-changing reagents, the stability of the color signal, and the ability to accurately quantify color changes using image capture and analysis |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [5 empty] | yes | lactate, chloride, glucose, creatinine |  |  |  |  | 1 | 4 | prototype | Human | Human studies | NA | NA | Soft microfluidics, fully integrated. | NO | NA | YES | NA | 2016 | Koh, A.; Kang, D.; Xue, Y.; Lee, S.; Pielak, R. M.; Kim, J.; Hwang, T.; Min, S.; Banks, A.; Bastien, P.; Manco, M. C.; Wang, L.; Ammann, K. R.; Jang, K. I.; Won, P.; Han, S.; Ghaffari, R.; Paik, U.; Slepian, M. J.; Balooch, G.; Huang, Y.; Rogers, J. A. | A soft, wearable microfluidic device for the capture, storage, and colorimetric sensing of sweat | Sci Transl Med |  | 1946-6242 (Electronic) 1946-6234 (Print) 1946-6234 (Linking) | Capabilities in health monitoring enabled by capture and quantitative chemical analysis of sweat could complement, or potentially obviate the need for, approaches based on sporadic assessment of blood samples. Established sweat monitoring technologies use simple fabric swatches and are limited to basic analysis in controlled laboratory or hospital settings. We present a collection of materials and device designs for soft, flexible, and stretchable microfluidic systems, including embodiments that integrate wireless communication electronics, which can intimately and robustly bond to the surface of the skin without chemical and mechanical irritation. This integration defines access points for a small set of sweat glands such that perspiration spontaneously initiates routing of sweat through a microfluidic network and set of reservoirs. Embedded chemical analyses respond in colorimetric fashion to markers such as chloride and hydronium ions, glucose, and lactate. Wireless interfaces to digital image capture hardware serve as a means for quantitation. Human studies demonstrated the functionality of this microfluidic device during fitness cycling in a controlled environment and during long-distance bicycle racing in arid, outdoor conditions. The results include quantitative values for sweat rate, total sweat loss, pH, and concentration of chloride and lactate. | 42697 | 2016
VFD5NGQ2 | 10.1016/j.talanta.2022.123208 | https://linkinghub.elsevier.com/retrieve/pii/S0039914022000042 | YES | skin-attached |  | NO | non-invasively collect, store, and analyze sweat for health monitoring. | Colorimetric | Colorimetric | NA (the authors acknowledge that the current method of visual assessment of color change using the naked eye or smartphone images limits the accuracy of the device.They suggest that incorporating photoelectric sensors and AI-based analysis of color changes in future developments could significantly improve the accuracy and enable quantitative measurements of sweat components). |  | NA | yes | - | - | sweat |  |  |  |  | yes | [5 empty] | yes |  | glucose, pH |  |  |  |  | 1 | 2 | prototype |  | in-vitro tests + laboratory testing | NA | NA | Tesla Valve Integration: adapting Tesla valves for a microfluidic chip eliminates moving parts, offering a robust and efficient alternative to traditional microvalves. This is the standout innovation of the device. Novel Microfluidic Design: the device employs a network of microchannels and sweat collection chambers designed for optimized sweat capture and storage. While the design is novel, it uses standard microfluidic materials and fabrication techniques. Future Potential for Advanced Sensing: the current reliance on conventional colorimetric sensing and smartphone-based visual analysis limits precision. Future enhancements could include photoelectric sensors and AI-driven algorithms, significantly improving data accuracy and processing capabilities. | NO | NA | NA | NA | 2022 | Shi, Huanhuan; Cao, Yu; Zeng, Yining; Zhou, Yanuo; Wen, Weihua; Zhang, Congxuan; Zhao, Yali; Chen, Zhen | Wearable tesla valve-based sweat collection device for sweat colorimetric analysis | Talanta |  | 399140 | Sweat sensor has become one of the most important developing directions of in vitro wearable diagnostic device in recent years. Stable sweat collecting device is the key to realize sweat component analysis. In order to ensure that the collected sweat is not subject to component analysis errors caused by evaporation or environmental pollution, mechanical micro-valves were adopted for microfluidic sweat collection devices to realize sealed storage of sweat. However, this poses a challenge to the stability of machining and reusability of the acquisition device. In this work, the Tesla valve without any mechanical structure were introduced into the design of sweat collection chip. And made full use of its diodicity to improve the collection to a certain extent, prevent backflow at the entrance, and restrain the flow at the exit to contact with the outside world. In addition, through optimizing the shunt angle, branch channel parameters of Tesla valve, boosted its diodicity under low flow rate. Furthermore, a sweat storage chamber with baffle structure that can achieve maximum static storage area was adopted to form a whole sweat collection chip. The design was verified through the flow experiment of methylene blue and methyl red indicators on the chip. Through modification of the filter paper fixed in the collection chamber, the colorimetric analysis of glucose and pH was realized. This device may provide new inspirations for the development of wearable sweat sensor. | 44652 | 2022
6UAK7XV8 | 10.1016/j.bios.2024.116430 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324004354 | YES | patch |  | NO | efficient collection and transport of small amounts of sweat for analysis. |  | NA (the sources primarily describe the device's microfluidic design and sweat collection capabilities) | With only 0.5 μl of the sweat secreted at the collection site, it can converge on the detection medium located within the center reservoir. |  | NA | yes | - | - | sweat | [12 empty] | NONE |  |  |  |  | 1 | NONE | Proof-of-concept/prototype | Human | laboratory testing + human studies | NA | NA | The most significant cutting-edge innovation in the device is its bio-inspired hierarchical microfluidic design for enhanced sweat collection. This design, inspired by the structure of Sarracenia trichomes, incorporates microchannels with rib structures at the bottom to facilitate rapid and efficient transport of small volumes of sweat. | NO | NA | NO | NA | 2024 | Wu, Tianjie; Yang, Pufan; Xie, Xintong; Cao, Xi; Deng, Yujun; Ding, Xianting; Zhang, Zhinan | Bio-inspired hierarchical wearable patch for fast sweat collection | Biosensors and Bioelectronics |  | 9565663 | Sweat contains abundant physiological and metabolic data to evaluate an individual’s physical health. Since the non-exercise sweat secretion rate is low, with an average value of 1–10 μl h 1 cm 2, sweat is generally collected during exercise for existing wearable sweat sensors. To expand their applications to include daily scenarios, these sensors developed for sports and fitness are challenged by the difficulty of collecting trace amounts of sweat. This study proposes a wearable patch inspired by the hierarchical structure of Sarracenia trichomes, allowing for the spontaneous and fast collection of a small amount of secreted sweat. The patch contains microfluidic channels featuring a 20 μm-wide rib structure, fully utilizing the capillary force, thereby eliminating the issue of sweat hysteresis. Furthermore, with only 0.5 μl of the sweat secreted at the collection site, it can converge on the detection medium located within the center reservoir. Volunteer verification demonstrated a twofold increase in sweat collection efficiency compared to traditional wearable patches. This patch serves as an efficient sweatcollection configuration, promising potential for diverse in situ sweat colorimetric analyses. | 45536 | 2024
2J68ABSE | 10.1038/s41598-018-33565-x | https://www.ncbi.nlm.nih.gov/pubmed/30367078 | YES | patch |  | NO | detection of analytes | electrochemical | Organic electrochemical transistor (OECT) | Sensing range is limited to lactate concentrations below approximately 1 mM. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate |  |  |  |  | 1 | 1 | prototype |  | Laboratory testing | NA | NA | Electronics patch OECT functionalized with lactate oxidase. | NO | NA | NA | NA | 2018 | Currano, L. J.; Sage, F. C.; Hagedon, M.; Hamilton, L.; Patrone, J.; Gerasopoulos, K. | Wearable Sensor System for Detection of Lactate in Sweat | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | Increased development of wearable sensors for physiological monitoring has spurred complementary interest in the detection of biochemical indicators of health and performance. We report a wearable sensor system for non-invasive detection of excreted human biomarkers in sweat. The system consists of a thin, flexible, kapton patch (2.5 x 7.5 cm) that can be coated with adhesive and affixed to the skin. The system can be controlled by a cell phone via a near-field communications protocol, charged wirelessly, and the data can be downloaded and displayed in a smart phone app. The system is designed such that the sensing element plugs into a low-profile socket, and can easily be removed and replaced as needed due to saturation or aging effects. As a demonstration case, we examined using an organic electrochemical transistor (OECT) within this system to monitor lactate concentration. Several different methods for optimizing the sensor performance were compared, including altering electrode materials, employing various immobilization techniques, and tailoring operating voltages. Resulting functional response of the lactate oxidase enzyme was compared as a function of the sensor variables. The OECT sensor was shown to have high sensitivity to lactate, however the sensing range is limited to lactate concentrations below approximately 1 mM. | 43399 | 2018
WCDCR8PV | 10.1016/j.bios.2018.02.025 | https://www.ncbi.nlm.nih.gov/pubmed/29471280 | YES | patch |  | NO | sweat pH monitoring | electrochemical | Electrochemical | pH sensitivity of 11.13±5.8 mV/pH with a maximum response time of 8 s. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | pH |  |  |  |  | 1 | 1 | prototype |  | Laboratory testing + in-vitro studies | NA | NA | Stretchable wireless system provided by a pair of serpentine-shaped stretchable interconnects. pH sensing electrode made of graphite-polyurethane composite | NO | NA | YES | NA | 2018 | Dang, W.; Manjakkal, L.; Navaraj, W. T.; Lorenzelli, L.; Vinciguerra, V.; Dahiya, R. | Stretchable wireless system for sweat pH monitoring | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Linking) | Sensor-laden wearable systems that are capable of providing continuous measurement of key physiological parameters coupled with data storage, drug delivery and feedback therapy have attracted huge interest. Here we report a stretchable wireless system for sweat pH monitoring, which is able to withstand up to 53% uniaxial strain and more than 500 cycles to 30% strain. The stretchability of the pH sensor patch is provided by a pair of serpentine-shaped stretchable interconnects. The pH sensing electrode is made of graphite-polyurethane composite, which is suitable for biosensor application. The sensing patch validated through in-depth electrochemical studies, exhibits a pH sensitivity of 11.13 +/- 5.8 mV/pH with a maximum response time of 8 s. Interference study of ions and analyte (Na(+), K(+) and glucose) in test solutions shows negligible influence on the pH sensor performance. The pH data can be wirelessly and continuously transmitted to smartphone through a stretchable radio-frequency-identification antenna, of which the radiating performance is stable under 20% strain, as proved by vector network analyzer measurement. To evaluate the full system, the pH value of a human sweat equivalent solution has been measured and wirelessly transmitted to a custom-developed smart phone App. | 43252 | 2018
VEQ5M8AF | 10.1002/advs.202203597 | https://www.ncbi.nlm.nih.gov/pubmed/35975449 | YES | contact lenses |  | NO | monitoring of cholesterol levels in tear fluid | electrochemical | Electrochemical | This cholesterol biosensor detected the free cholesterol from 0.01 to 1.2 mm (corresponding to 0.4 to 46.4 mg dL–1).This sensor showed high sensitivity with 1% current change per 0.043 mm. The wide linear detection range can cover the physiological tear cholesterol ranges of both healthy humans (≈0.4 mm) and hyperlipidemia patients.The signal-to-noise ratio (SNR) was 4.29 at 0.01 mm (0.4 mg dL–1) and LOD was 9.91 μm (0.38 mg dL–1). | Metabolic Syndrome | management of hyperlipidemia | - | - | - | tear | [9 empty] | yes |  |  | cholesterol |  |  |  |  | 1 | 1 | Validation | Human | Animal models + human studies | NA | NA | A wireless and soft smart contact lens that enables real-time quantitative recording of cholesterol in tear fluids for the monitoring of patients with hyperlipidemia using a smartphone is reported. This contact lens incorporates an electrochemical biosensor for the continuous detection of cholesterol concentrations, stretchable antenna, and integrated circuits for wireless communication, which makes a smartphone the only device required to operate this lens remotely without obstructing the wearer's vision. | NO | NA | YES | NA | 2022 | Song, H.; Shin, H.; Seo, H.; Park, W.; Joo, B. J.; Kim, J.; Kim, J.; Kim, H. K.; Kim, J.; Park, J. U. | Wireless Non-Invasive Monitoring of Cholesterol Using a Smart Contact Lens | Adv Sci (Weinh) |  | 2198-3844 (Electronic) 2198-3844 (Linking) | Herein, a wireless and soft smart contact lens that enables real-time quantitative recording of cholesterol in tear fluids for the monitoring of patients with hyperlipidemia using a smartphone is reported. This contact lens incorporates an electrochemical biosensor for the continuous detection of cholesterol concentrations, stretchable antenna, and integrated circuits for wireless communication, which makes a smartphone the only device required to operate this lens remotely without obstructing the wearer's vision. The hyperlipidemia rabbit model is utilized to confirm the correlation between cholesterol levels in tear fluid and blood and to confirm the feasibility of this smart contact lens for diagnostic application of cholesterol-related diseases. Further in vivo tests with human subjects demonstrated its good biocompatibility, wearability, and reliability as a non-invasive healthcare device. | 44835 | 2022
LVZCVZA2 | 10.1021/acssensors.8b00168 | https://pubs.acs.org/doi/10.1021/acssensors.8b00168 | YES | wrist-worn |  | NO | non-enzymatic electrochemical analysis of perspiration glucose | electrochemical | Electrochemical | Reproducibility: the researchers evaluated the sensor's reproducibility by performing replicated measurements on the same perspiration sample. They found a relative standard deviation (RSD) of approximately 7.5% for seven replicate analyses. Long-Term Stability: to assess the sensor's stability over time, they conducted repeated analyses of perspiration glucose over two weeks. The RSD for these measurements was about 4.6%. | Diabetes | diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype | Human | Laboratory testing + human studies | NA | NA | The most significant innovation is the use of non-enzymatic electrochemical detection for glucose. This approach overcomes several limitations associated with traditional enzyme-based glucose sensors, particularly in the context of wearable devices. The device employs a novel multipotential step method applied to the gold working electrode to achieve accurate and reliable glucose measurements. The researchers incorporated a dual-layer membrane coating on the gold electrode to enhance the sensor's selectivity and robustness The sensor system, including the three electrodes, miniaturized potentiostat, and Bluetooth communication module, is integrated into a 3D-printed wristband, demonstrating the device's practical feasibility as a wearable technology | NO | NA | YES | NA | 2018 | Zhu, Xiaofei; Ju, Yinhui; Chen, Jian; Liu, Deye; Liu, Hong | Nonenzymatic Wearable Sensor for Electrochemical Analysis of Perspiration Glucose | ACS sensors |  |  | We report a nonenzymatic wearable sensor for electrochemical analysis of perspiration glucose. Multipotential steps are applied on a Au electrode, including a high negative pretreatment potential step for proton reduction which produces a localized alkaline condition, a moderate potential step for electrocatalytic oxidation of glucose under the alkaline condition, and a positive potential step to clean and reactivate the electrode surface for the next detection. Fluorocarbon-based materials were coated on the Au electrode for improving the selectivity and robustness of the sensor. A fully integrated wristband is developed for continuous real-time monitoring of perspiration glucose during physical activities, and uploading the test result to a smartphone app via Bluetooth. | 43252 | 2018
5K54N2YN | 10.1016/j.bios.2020.112450 | https://linkinghub.elsevier.com/retrieve/pii/S0956566320304449 | YES | patch |  | NO | non-invasive monitoring of specific biomarkers found in sweat. | Optical | Optical | Chloride Sensor: Linear range: 10–100 mM; Correlation coefficient (r): 0.994; Relative standard deviation: 3.41% for 60 mM chloride; Limit of detection (S/N = 3): 5 mM. pH Sensor: Linear range: 4.1–5.1 and 5.1–6.6; Correlation coefficients (r): 0.982 and 0.988, respectively; Relative standard deviation: 6.28% for pH 5.0; Limit of detection (S/N = 3): Not provided. Volume Sensor: Linear range: 15–65 μL; Correlation coefficient (r): 0.994; Relative standard deviation: 2.90% for 45 μL; Limit of detection (S/N = 3): 8 μL. Glucose Sensor: Linear range: 10–250 μM; Correlation coefficient (r): 0.990; Relative standard deviation: 6.53% for 100 μM glucose; Limit of detection (S/N = 3): 7 μM. Lactate Sensor: Linear range: 1.0–12.5 mM; Correlation coefficient (r): 0.988; Relative standard deviation: 4.86% for 7.5 mM lactate; Limit of detection (S/N = 3): 0.4 mM. The sensors exhibit high correlation coefficients, indicating a strong linear relationship between their responses and the respective analyte concentrations. The relative standard deviations are generally low, suggesting good precision in the measurements. The limits of detection vary depending on the analyte, with glucose having the lowest LOD (7 μM) and chloride having the highest (5 mM). |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes |  |  |  |  | yes |  | Glucose, Lactate, pH, Chloride, Sweat Volume |  |  |  |  | 1 | 5 | Validation | Human | Human studies | NA | NA | The SWSP doesn't introduce entirely new sensor designs or materials, its strength lies in the creative combination of existing technologies to address the limitations of traditional sweat analysis methods. This integration, coupled with the use of a smartphone platform and custom-developed software, results in a device that is low-cost, portable, user-friendly, and capable of multi-sensing five important biomarkers. | NO | NA | YES | NA | 2020 | Ardalan, Sina; Hosseinifard, Mohammad; Vosough, Maryam; Golmohammadi, Hamed | Towards smart personalized perspiration analysis: An IoT-integrated cellulose-based microfluidic wearable patch for smartphone fluorimetric multi-sensing of sweat biomarkers | Biosensors and Bioelectronics |  | 9565663 | Practical obstacles, such as intricate designs and expensive equipment/materials, in the fabrication of wearable sweat sensors, have limited their feasibility as a personalized healthcare device. Herein, we have fabricated a cellulose-based wearable patch, which further paired with a smartphone-based fluorescence imaging module and a self-developed smartphone app for non-invasive and in situ multi-sensing of sweat biomarkers including glucose, lactate, pH, chloride, and volume. The developed Smart Wearable Sweat Patch (SWSP) sensor comprises highly fluorescent sensing probes embedded in paper substrates, and microfluidic channels consisted of cotton threads to harvest sweat from the skin surface and to transport it to the paper-based sensing probes. The imaging module was fabricated by a 3D printer, equipped with UV-LED lamps and an optical filter to provide the in situ capability of capturing digital images of the sensors via a smartphone. A smartphone app was also designed to quantify the concentration of the biomarkers via a detection algorithm. Additionally, we have recommended an Internet of Things (IoT)-based model for our developed SWSP sensor to promote its potential application for the future. The field studies on human subjects were also conducted to investigate the feasibility of our developed SWSP sensor for the analysis of sweat biomarkers. Our findings convincingly demonstrated the applicability of our developed SWSP sensor as a smart, user-friendly, ultra-low-cost (~0.03 $ per sweat patch), portable, selective, rapid, and non-invasive healthcare monitoring device for immense applications in health personalization, sports performance monitoring, and medical diagnostics. | 44136 | 2020
TW8YWWQJ | 10.1007/s00216-021-03865-9 | https://www.ncbi.nlm.nih.gov/pubmed/35015101 | YES | skin-attached |  | NO | monitor and analyze biomarkers in sweat for healthcare and wellness applications. | electrochemical | Electrochemical | The sensing performance of the device was demonstrated by monitoring sweat glucose, presenting good correlation with blood glucose measurements. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Proof-of-concept/validation | Human | Human studies | NA | NA | The device represents the frst system integrating an iontophoretic pilocarpine delivery system around the inlet channels of epidermal polydimethylsiloxane (PDMS) microfuidic device for sweat collection and analysis. | NO | NA | YES | NA | 2022 | Bolat, G.; De la Paz, E.; Azeredo, N. F.; Kartolo, M.; Kim, J.; de Loyola, E. Silva A. N.; Rueda, R.; Brown, C.; Angnes, L.; Wang, J.; Sempionatto, J. R. | Wearable soft electrochemical microfluidic device integrated with iontophoresis for sweat biosensing | Anal Bioanal Chem |  | 1618-2650 (Electronic) 1618-2642 (Linking) | A soft and flexible wearable sweat epidermal microfluidic device capable of simultaneously stimulating, collecting, and electrochemically analyzing sweat is demonstrated. The device represents the first system integrating an iontophoretic pilocarpine delivery system around the inlet channels of epidermal polydimethylsiloxane (PDMS) microfluidic device for sweat collection and analysis. The freshly generated sweat is naturally pumped into the fluidic inlet without the need of exercising. Soft skin-mounted systems, incorporating non-invasive, on-demand sweat sampling/analysis interfaces for tracking target biomarkers, are in urgent need. Existing skin conformal microfluidic-based sensors for continuous monitoring of target sweat biomarkers rely on assays during intense physical exercising. This work demonstrates the first example of combining sweat stimulation, through transdermal pilocarpine delivery, with sample collection through a microfluidic channel for real-time electrochemical monitoring of sweat glucose, in a fully integrated soft and flexible multiplexed device which eliminates the need of exercising. The on-body operational performance and layout of the device were optimized considering the fluid dynamics and evaluated for detecting sweat glucose in several volunteers. Furthermore, the microfluidic monitoring device was integrated with a real-time wireless data transmission system using a flexible electronic board PCB conformal with the body. The new microfluidic platform paves the way to real-time non-invasive monitoring of biomarkers in stimulated sweat samples for diverse healthcare and wellness applications. | 44743 | 2022
LYPA3K6T | 10.1038/s41598-021-86931-7 | https://www.ncbi.nlm.nih.gov/pubmed/33828144 | YES | skin-attached |  | NO | glucose monitoring | Other | magnetohydrodynamics (MHD) | The results showed that the efficiency of the MHD extraction method is superior to that of reverse iontophoresis. Using a 300 mT magnetic field we achieved a 13-fold increase in active glucose extraction. Furthermore, the amount of extracted glucose was proportional to the glucose concentration in the sample, which is important when considering potential applications in glucose sensing where the extracted amount of glucose in ISF is expected to correlate with the concentration of glucose in blood. | Diabetes | diabetes management | - | - | yes | interstitial fluid |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Proof-of-concept |  | Ex-vivo testing | NA | NA | Use of magnetohydrodynamics (MHD) for extracting interstitial fluid through the skin | GlucoModicum Ltd | Helsinki, Finland. | NA | NA | 2021 | Hakala, T. A.; Garcia Perez, A.; Wardale, M.; Ruuth, I. A.; Vanska, R. T.; Nurminen, T. A.; Kemp, E.; Boeva, Z. A.; Alakoskela, J. M.; Pettersson-Fernholm, K.; Haeggstrom, E.; Bobacka, J. | Sampling of fluid through skin with magnetohydrodynamics for noninvasive glucose monitoring | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | Out of 463 million people currently with diabetes, 232 million remain undiagnosed. Diabetes is a threat to human health, which could be mitigated via continuous self-monitoring of glucose. In addition to blood, interstitial fluid is considered to be a representative sample for glucose monitoring, which makes it highly attractive for wearable on-body sensing. However, new technologies are needed for efficient and noninvasive sampling of interstitial fluid through the skin. In this report, we introduce the use of Lorentz force and magnetohydrodynamics to noninvasively extract dermal interstitial fluid. Using porcine skin as an ex-vivo model, we demonstrate that the extraction rate of magnetohydrodynamics is superior to that of reverse iontophoresis. This work seeks to provide a safe, effective, and noninvasive sampling method to unlock the potential of wearable sensors in needle-free continuous glucose monitoring devices that can benefit people living with diabetes. | 44293 | 2021
66IIGTSV | 10.1016/j.bios.2021.113530 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321005674 | YES | patch |  | NO | "turn-on" detection of specific biomarkers in sweat | Optical | Optical | Limit of Detection (LOD): The source mentions that for uric acid detection, the LOD was determined to be 136 nM. The LOD represents the lowest concentration of an analyte that can be reliably distinguished from the background noise. This value provides an indication of the sensor's sensitivity towards uric acid. Linear Ranges: The source provides linear ranges for each analyte, indicating the concentration ranges where the sensor's response is proportional to the analyte concentration: Uric acid: 0 μM to 50 μM Glucose: 0 μM to 1 mM Alcohol: 0 mM to 15 mM | Metabolic Syndrome | monitoring metabolic conditions | yes | - | - | sweat |  |  |  |  | yes |  |  | yes |  |  |  |  | Uric acid, Glucose, Alcohol |  |  |  |  | 1 | 3 | Proof-of-concept | Human | laboratory testing + human studies (preliminary) | NA | NA | Enzyme-Embedded AuNC Assemblies: Combines enzymes for high specificity with gold nanoclusters (AuNCs) acting as luminophores, enabling selective, color-based detection of uric acid, glucose, and alcohol in sweat. MnO2 Nanosheets as a Luminescence Switch: Employs MnO2 nanosheets to quench AuNC luminescence, which is restored upon enzymatic degradation, enhancing sensitivity at low analyte concentrations. Microwell-Patterned Design: Incorporates microwells for efficient sweat collection and multiplexed detection, enabling localized and simultaneous analyte analysis. Flexible Wearable Substrate: Built on a biocompatible, flexible tape for comfortable, real-world usability. Smartphone-Based Detection: Uses a smartphone app for image-based luminescence analysis, providing a portable, user-friendly monitoring solution. | NO | NA | YES | NA | 2021 | Zhou, Ziping; Shu, Tong; Sun, Yafang; Si, Hongxin; Peng, Peiwen; Su, Lei; Zhang, Xueji | Luminescent wearable biosensors based on gold nanocluster networks for “turn-on” detection of Uric acid, glucose and alcohol in sweat | Biosensors and Bioelectronics |  | 9565663 | From the difficulty of awareness of abnormal concentrations of biochemical indexes in people’s daily life come wearable sensing technologies. Recently, luminescent wearable biosensors are emerging with simple fabrication, easy use, cost-effectivity and reliability. But several challenges should be taken up, such as availability of varied analytes, high sensitivity, stability of enzymes, photostability, low signal noises and recyclability of sensors. Here, the Luminescent Wearable Sweat Tape (LWST) biosensor is developed via embedding multi-component nanoprobes onto microwell-patterned paper substrates of hollowed-out double-side tapes. The nanoprobes consist of responsive luminophores, enzyme-loaded gold nanocluster (AuNCs) nano-networks, which are wrapped by the switch, MnO2 nanosheets. The responsive luminophores are constructed by 3 substitutable components: enzymes (uricase, GOx and alcohol dehydrogenase) for molecular target recognition, glutathioneprotected AuNCs (yellow, red and green) for luminescent signal output and polycations PAH for integration. MnO2 NSs as the switch can quench the emission of the AuNCs but degraded by the reductive product of incorporated enzymes. Thus, targeting analysts (uric acid, glucose and alcohol) can be dose-dependently detected through “turn-on” luminescence approach. After incorporating the nanoprobes into hollow-out tapes, the formed LWST biosensors can detect uric acid, glucose and alcohol in sweat with the help of a smartphone. Subsequently, we primarily apply them into human daily life scenario, sampling from dine parties, and the positive relationships of analyte intakes and the increase of analytes in sweat are significant with individual difference. | 44501 | 2021
XZKVMMFJ | 10.1002/adma.202304704 | https://www.ncbi.nlm.nih.gov/pubmed/37709513 | YES | patch | microneedle skin-attached | YES | continuous sampling and sensing of dermal interstitial fluid (ISF) | electrochemical | Electrochemical | The sensitivity and limit of detection (LOD= 3.3𝜎/S, where 𝜎 is the standard deviation of sensor response, and S is the slope of the calibration curve) of the electrode at low glucose concentrations (20-1000 μM) were 0.254 μA mM−1 and 5.6 μM, respectively. Subsequently, the sensor response was examined for glucose concertation ranging from 3 to 15 mM, relevant to the physiological glucose concentrations in ISF. The measured sensor response increased with the addition of glucose, showing a linear trend with a correlation co-eﬃcient of 0.98, which could be further used to obtain the calibration curve. The corresponding sensitivity and LOD were calculated as 7.1 nA mM−1 and 0.2 mM, respectively. The results suggest that the glucose sensor could fulﬁll the requirements for glucose monitoring in a detection range that can cover normal glucose levels in human blood (≈4.4 to 6.6 mM), which is practically the same as ISF glucose level. The relative standard deviation (RSD) of the measured current at glucose concentrations of 100 μM and 5 mM were 1.22% and 2.57%, respectively, indicating the reliable detection consistency of the sensors. Additionally, a negligible deviation (RSD of 2.9%) was evidenced in the functionality of the sensors over a storage time of seven days, reﬂect-ing robust enzyme immobilization. The pH sensor exhibited a linear response of 50.38 mV pH−1 with a correlation coeﬃcient and maximum RSD value of 0.98 and 9.07%, respectively, in three complete cycles of pH between 7.17 and 8.19. Because of the importance of the reversibility in potentiometric sensors, the cyclic response of the pH sensor was also examined for a broader range of pH values between 6.94 and 9.23. The results revealed good reversibility behaviour with an average RSD value of 7.54%. Besides, the sensors were exposed to various pH levels of between 6.94 and 9.23 in a staircase manner to explore the eﬀect of abrupt pH change on the sensor performance. The large pH variations had an eﬀect on the continuous sensor response with a minimal deviation of 8.04%, and the pH of ISF can be reliably monitored using the sensor. |  | NA | - | - | yes | dermal interstitial fluid |  |  |  |  | yes | [5 empty] | yes |  | glucose, pH |  |  |  |  | 1 | 2 | Proof-of-concept |  | In-vitro, ex-vivo, in-vivo (animal models) | NA | NA | a fully integrated touch-activated wearable device based on a laser-drilled hollow microneedle (HMN) patch for continuous sampling and sensing of dermal ISF is introduced. The developed platform can produce and maintain the required vacuum pressure (as low as ≈ -53 kPa) to collect adequate volumes of ISF (≈2 µL needle-1 h-1 ) for medical applications. The vacuum system can be activated through a one-touch finger operation. | NO | NA | NA | NA | 2024 | Abbasiasl, T.; Mirlou, F.; Mirzajani, H.; Bathaei, M. J.; Istif, E.; Shomalizadeh, N.; Cebecioglu, R. E.; Ozkahraman, E. E.; Yener, U. C.; Beker, L. | A Wearable Touch-Activated Device Integrated with Hollow Microneedles for Continuous Sampling and Sensing of Dermal Interstitial Fluid | Adv Mater |  | 1521-4095 (Electronic) 0935-9648 (Linking) | Dermal interstitial fluid (ISF) is emerging as a rich source of biomarkers that complements conventional biofluids such as blood and urine. However, the impact of ISF sampling in clinical applications has been limited owing to the challenges associated with extraction. The implementation of microneedle-based wearable devices that can extract dermal ISF in a pain-free and easy-to-use manner has attracted growing attention in recent years. Here, a fully integrated touch-activated wearable device based on a laser-drilled hollow microneedle (HMN) patch for continuous sampling and sensing of dermal ISF is introduced. The developed platform can produce and maintain the required vacuum pressure (as low as approximately -53 kPa) to collect adequate volumes of ISF ( approximately 2 microL needle(-1) h(-1) ) for medical applications. The vacuum system can be activated through a one-touch finger operation. A parametric study is performed to investigate the effect of microneedle array size, vacuum pressure, and extraction duration on collected ISF. The capability of the proposed platform for continuous health monitoring is further demonstrated by the electrochemical detection of glucose and pH levels of ISF in animal models. This HMN-based system provides an alternative tool to the existing invasive techniques for ISF collection and sensing for medical diagnosis and treatment. | 45292 | 2024
7ZN7U6CB | 10.1002/advs.202104635 | https://www.ncbi.nlm.nih.gov/pubmed/35088587 | YES | skin-attached |  | NO | monitoring sweat | electrochemical | Electrochemical | the voltages of the glucose sensors are linearly proportional to the glucose concentrations with a determination coefficient (R2) of 0.998. An excellent linear relationship between the OCV (versusAg/AgCl) of the Na+ sensor and the value of the logarithm of Na+ concentration, with an excellent R2 of 0.994 that corresponds to a sensitivity of 131.95 mV per decade of concentration for Na+. A near-Nernstian sensitivity o f58.27mV per pH is observed for the pH sensor based on polyaniline (PANI), with an R2 of 0.995. This sensor also shows good selectivity and anti-interference ability, as the sensor is only responsive to H+ concentration. In conclusion, all these soft biosensors exhibit good performance, including linearity, sensitivity, and selectivity, which would be extremely important for real-time and continuous analysis of perspiration contents and reflecting the health status of the testing subjects. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [5 empty] | yes |  | pH, glucose, Na+ |  |  |  |  | 1 | 3 | Proof-of-concept | Human | Human studies | NA | NA | Sweat-activated batteries (SABs) for powering microfluidic system based sweat sensors are realized in this work, allowing real-time monitoring of pH, glucose, and Na+ in sweat. | NO | NA | NA | NA | 2022 | Liu, Y.; Huang, X.; Zhou, J.; Yiu, C. K.; Song, Z.; Huang, W.; Nejad, S. K.; Li, H.; Wong, T. H.; Yao, K.; Zhao, L.; Yoo, W.; Park, W.; Li, J.; Huang, Y.; Lam, H. R.; Song, E.; Guo, X.; Wang, Y.; Dai, Z.; Chang, L.; Li, W. J.; Xie, Z.; Yu, X. | Stretchable Sweat-Activated Battery in Skin-Integrated Electronics for Continuous Wireless Sweat Monitoring | Adv Sci (Weinh) |  | 2198-3844 (Electronic) 2198-3844 (Linking) | Wearable electronics have attracted extensive attentions over the past few years for their potential applications in health monitoring based on continuous data collection and real-time wireless transmission, which highlights the importance of portable powering technologies. Batteries are the most used power source for wearable electronics, but unfortunately, they consist of hazardous materials and are bulky, which limit their incorporation into the state-of-art skin-integrated electronics. Sweat-activated biocompatible batteries offer a new powering strategy for skin-like electronics. However, the capacity of the reported sweat-activated batteries (SABs) cannot support real-time data collection and wireless transmission. Focused on this issue, soft, biocompatible, SABs are developed that can be directly integrated on skin with a record high capacity of 42.5 mAh and power density of 7.46 mW cm(-2) among the wearable sweat and body fluids activated batteries. The high performance SABs enable powering electronic devices for a long-term duration, for instance, continuously lighting 120 lighting emitting diodes (LEDs) for over 5 h, and also offers the capability of powering Bluetooth wireless operation for real-time recording of physiological signals for over 6 h. Demonstrations of the SABs for powering microfluidic system based sweat sensors are realized in this work, allowing real-time monitoring of pH, glucose, and Na(+) in sweat. | 44621 | 2022
XG22VQQK | 10.1039/d3mh00989k | https://www.ncbi.nlm.nih.gov/pubmed/37725333 | YES | smart sensing garment/electrochemical sensing fibers |  | NO | Sweat Capture and Biomarker Monitoring | electrochemical | Electrochemical | The sensitivity for glucose and lactate is calculated as 6 nA mM1 and 69 nA mM1, respectively. The limit of detection (LOD) for those two biomarkers is calculated as 4.55 and 11.49 mM, respectively. Fig. 4c–e displays the potential response variation of the Na+, Ca2+, and K+ sensing fibres in electrolyte solutions with concentrations of 10–160 mM, 4–8 mM, and 2–32 mM, respectively. Based on ion-selective membranes, the sensing fibers demonstrate concentration changes through surface potential variations and exhibit near-Nernst behaviour (according to the Nernst equation, the change in potential is proportional to the logarithmic value of ion concentration when the ion concentration in the electrode and solution are in equilibrium) with sensitivities of 67 mV dec 1, 29.4 mV dec 1, and 39 mV dec 1 for Na+, Ca2+, and K+ sensing fibers, respectively. The LODs for the three biomarkers were 0.549 mM, 0.211 mM and 0.134 mM, respectively. Fig. 4f demonstrates the signal response of the pH-sensing fiber based on the protonation variation of polyaniline on the surface at different pH values. The pH-sensing fiber exhibits a sensitivity of 28 mV dec 1 in the pH range of 4 to 7. The selectivity of the sweat sensor is another crucial property, as sweat contains various electrolytes and metabolites that could impact sensor performance. To assess selectivity, we added interfering substances to the relevant solutions and conducted tests. The results show that the response to the target substance is more evident than that to the interfering substances, indicating good selectivity of electrochemical sensing fibers, which is beneficial for integrating multiplexed sensing systems. The sensing fiber samples from different batches exhibited similar signal responses to the same concentration of solution, confirming good reproducibility. Meanwhile, the electrochemical performance of the sensing fibers remained stable for more than 10 h, indicating their good stability. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes |  |  |  |  | yes |  | glucose, lactate, Na+, K+, Ca2+, and pH |  |  |  |  | 1 | 6 | Prototype/Validation | Human | In-situ (human) + ex-situ | NA | NA | Electrochemical sensing fiber with a special core-sheath structure, whose outer layer is wound by nanofibers with a hierarchical Fermat helix structure which has excellent moisture conductivity, and the inner layer is based on CNT fibers covered by three-dimensional reduced graphene oxide folds which have good sensing properties after modification of active materials such as enzymes and selective membranes. | NO | NA | NA | NA | 2023 | Tian, H.; Wang, L.; Yang, W.; Li, K.; Zhang, Q.; Li, Y.; Wang, H.; Hou, C. | Hierarchical Fermat helix-structured electrochemical sensing fibers enable sweat capture and multi-biomarker monitoring | Mater Horiz |  | 2051-6355 (Electronic) 2051-6347 (Linking) | Wearable electrochemical sensors have shown potential for personal health monitoring due to their ability to detect biofluids non-invasively at the molecular level. Smart fibers with high flexibility and comfort are currently ideal for fabricating electrochemical sensors, but little research has focused on fluid transport at the human-machine interface, which is of great significance for continuous and stable monitoring and skin comfort. Here, we report an electrochemical sensing fiber with a special core-sheath structure, whose outer layer is wound by nanofibers with a hierarchical Fermat helix structure which has excellent moisture conductivity, and the inner layer is based on CNT fibers covered by three-dimensional reduced graphene oxide folds which have good sensing properties after modification of active materials such as enzymes and selective membranes. This kind of fiber enables efficient sweat capture, and thus only 0.1 muL of sweat is required to activate the device, and it responds very quickly (1.5 s). The fibers were further integrated into a garment to build a wireless sweat detection system, enabling stable monitoring of six physiological markers in sweat (glucose, lactate, Na(+), K(+), Ca(2+), and pH). This work provides a feasible proposal for future personalized medicine and the construction of "smart sensing garments". | 45229 | 2023
PDVI6XMA | 10.3390/s22218223 | https://www.ncbi.nlm.nih.gov/pubmed/36365929 | YES | skin-attached |  | NO | chloride detection in sweat | electrochemical | Electrochemical | The devices with PLA membrane exhibit a peak current density of 165 µA µM−1 cm−2 ± 4.8%. This is a very good result that shows both the reproducibility of the device (both in terms of fabrication and operation) and, again, the negligible interference of the PLA membrane on the operation of the device. The standard deviation of the proposed system is about 4.5% and possesses a very high selectivity. To validate the proposed sensor platform, sweat chloride was quantified prior to dilution of the sweat collected from a healthy volunteer with a standard solution of 0.1 M sodium nitrate (diluting factor = 30). The test was performed in triplicate, and the results showed a concentration of Cl− of 32.29 ± 3.8 mM. This value is in agreement with the standard concentration of chloride ions in human sweat. The Cl− concentration in the real sweat was also measured by the standard titration method. The measured value is 32.94 ± 5 mM confirming the good performance of the sensor. |  | NA | yes | - | - | sweat | [12 empty] | chloride |  |  |  |  | 1 | 1 | Prototype | Human | Human studies | NA | NA | A green wearable sensor for sweat absorption and chloride sensing is presented. In order to produce a sustainable device, polylactic acid (PLA) was used for both the substrate and the sweat absorption pad fabrication. The sensor material for chloride detection consisted of silver-based reference, working, and counter electrodes obtained from upcycled compact discs. | NO | NA | NA | NA | 2022 | Lopresti, F.; Patella, B.; Divita, V.; Zanca, C.; Botta, L.; Radacsi, N.; O'Riordan, A.; Aiello, G.; Kersaudy-Kerhoas, M.; Inguanta, R.; La Carrubba, V. | Green and Integrated Wearable Electrochemical Sensor for Chloride Detection in Sweat | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | Wearable sensors for sweat biomarkers can provide facile analyte capability and monitoring for several diseases. In this work, a green wearable sensor for sweat absorption and chloride sensing is presented. In order to produce a sustainable device, polylactic acid (PLA) was used for both the substrate and the sweat absorption pad fabrication. The sensor material for chloride detection consisted of silver-based reference, working, and counter electrodes obtained from upcycled compact discs. The PLA substrates were prepared by thermal bonding of PLA sheets obtained via a flat die extruder, prototyped in single functional layers via CO(2) laser cutting, and bonded via hot-press. The effect of cold plasma treatment on the transparency and bonding strength of PLA sheets was investigated. The PLA membrane, to act as a sweat absorption pad, was directly deposited onto the membrane holder layer by means of an electrolyte-assisted electrospinning technique. The membrane adhesion capacity was investigated by indentation tests in both dry and wet modes. The integrated device made of PLA and silver-based electrodes was used to quantify chloride ions. The calibration tests revealed that the proposed sensor platform could quantify chloride ions in a sensitive and reproducible way. The chloride ions were also quantified in a real sweat sample collected from a healthy volunteer. Therefore, we demonstrated the feasibility of a green and integrated sweat sensor that can be applied directly on human skin to quantify chloride ions. | 44861 | 2022
IHL6CF6F | 10.1016/j.talanta.2022.123697 | https://www.ncbi.nlm.nih.gov/pubmed/35752089 | YES | skin-attached |  | NO | detection of biomarkers in undiluted tears in the eye. |  | electronic detection (graphene field-effect transistor (FET) nanosensor) | Limit of detection of 0.34 pM for tumor necrosis factor alpha (TNF-alpha) in undiluted artificial tears |  | NA | - | - | - | undiluted artificial tears | [10 empty] | yes |  | Tumor necrosis factor alpha (TNF-alpha) |  |  |  |  | 1 | 1 | Proof-of-concept |  | In-vitro tests | NA | NA | The nanosensor is a graphene field-effect transistor, in which a nucleic acid aptamer and a biomolecule-permeable polyethylene glycol (PEG) nanolayer are immobilized on the graphene surface. | NO | NA | NA | NA | 2022 | Wang, Z.; Dai, W.; Yu, S.; Hao, Z.; Pei, R.; De Moraes, C. G.; Suh, L. H.; Zhao, X.; Lin, Q. | Towards detection of biomarkers in the eye using an aptamer-based graphene affinity nanobiosensor | Talanta |  | 1873-3573 (Electronic) 0039-9140 (Print) 0039-9140 (Linking) | We present an approach to enable the sensitive and specific detection of biomarkers in undiluted tears in the eye using an aptamer-based graphene affinity nanosensor. The nanosensor is a graphene field-effect transistor, in which a nucleic acid aptamer and a biomolecule-permeable polyethylene glycol (PEG) nanolayer are immobilized on the graphene surface. The aptamer is capable of specifically recognize the target biomarker and induce a change in the carrier concentration of the graphene, which is measured to determine the biomarker concentration. The PEG nanolayer minimizes nonspecific adsorption of background molecules in the sample that would otherwise interfere with the biomarker detection. Experimental results show that tumor necrosis factor alpha (TNF-alpha), an inflammatory cytokine, can be sensitively and specifically detected in undiluted artificial tears with a limit of detection of 0.34 pM. This ability to detect and measure biomarkers in undiluted physiological fluids allows the nanosensor to be potentially used in applications where sample dilutions are not practical, such as wearable measurements of tear-borne biomarkers in the eye. | 44896 | 2022
RUEKFCSI | 10.1126/sciadv.aaw9906 | https://www.ncbi.nlm.nih.gov/pubmed/31453333 | YES | patch | skin-attached | NO | sweat monitoring | electrochemical | Electrochemical | the general performance of a Na+ sensor in 15 to 120 mM NaCl solutions. The sensor remains stable throughout the measurement period with a sensitivity of 56.2 mV/decade, similar to sensitivities achieved by lithography-based Na+ sensors and close to the Nernstian ideal. Similarly, the K+ sensor shows a sensitivity of 51.3 mV/decade in standard solution of 5 to 40 mM KCl. The enzymatic glucose sensor has a sensitivity of 1.0 nA/μM with a linear response over the tested concentration range of 50 to 200 μM. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | Na+, K+, Glucose, (+ Sweat Rate) |  |  |  |  | 1 | 3 | prototype | Human | Human studies | NA | NA | The microfluidic sensing patches are mass-fabricated using R2R processes. | NO | NA | YES | YES | 2019 | Nyein, H. Y. Y.; Bariya, M.; Kivimaki, L.; Uusitalo, S.; Liaw, T. S.; Jansson, E.; Ahn, C. H.; Hangasky, J. A.; Zhao, J.; Lin, Y.; Happonen, T.; Chao, M.; Liedert, C.; Zhao, Y.; Tai, L. C.; Hiltunen, J.; Javey, A. | Regional and correlative sweat analysis using high-throughput microfluidic sensing patches toward decoding sweat | Sci Adv |  | 2375-2548 (Electronic) 2375-2548 (Linking) | Recent technological advancements in wearable sensors have made it easier to detect sweat components, but our limited understanding of sweat restricts its application. A critical bottleneck for temporal and regional sweat analysis is achieving uniform, high-throughput fabrication of sweat sensor components, including microfluidic chip and sensing electrodes. To overcome this challenge, we introduce microfluidic sensing patches mass fabricated via roll-to-roll (R2R) processes. The patch allows sweat capture within a spiral microfluidic for real-time measurement of sweat parameters including [Na(+)], [K(+)], [glucose], and sweat rate in exercise and chemically induced sweat. The patch is demonstrated for investigating regional sweat composition, predicting whole-body fluid/electrolyte loss during exercise, uncovering relationships between sweat metrics, and tracking glucose dynamics to explore sweat-to-blood correlations in healthy and diabetic individuals. By enabling a comprehensive sweat analysis, the presented device is a crucial tool for advancing sweat testing beyond the research stage for point-of-care medical and athletic applications. | 43678 | 2019
5YGDVSZT | 10.3390/s22228971 | https://www.ncbi.nlm.nih.gov/pubmed/36433566 | YES | skin-attached |  | NO |  | electrochemical | Electrochemical | The device is selective for glucose, and it detected glucose accurately in the clinically relevant range (0~2 mM) in an off-body setup. Acceptable sensitivity of 16.8 µA/mM/cm2. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Proof-of-concept | Human | Human studies | NA | NA | Three-dimensional paper-based microfluidic electrochemical integrated device (3D PMED) for measuring glucose concentration in sweat in real-time via simple, non-invasive, capillary-action-based sample collection. | NO | NA | NA | NA | 2022 | Noura, Z.; Shah, I.; Aziz, S.; Ahmed, A.; Jung, D. W.; Brahim, L.; ElMostafa, R. | Wearable Healthcare Monitoring Based on a Microfluidic Electrochemical Integrated Device for Sensing Glucose in Natural Sweat | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | Wearable sweat sensors offer the possibility of continuous real-time health monitoring of an individual at a low cost without invasion. A variety of sweat glucose sensors have been developed thus far to help diabetes patients frequently monitor blood glucose levels through sweat glucose as a surrogate marker. The present study demonstrates the development and characterization of a three-dimensional paper-based microfluidic electrochemical integrated device (3D PMED) for measuring glucose concentration in sweat in real-time via simple, non-invasive, capillary-action-based sample collection. The device was selective for glucose, and it detected glucose accurately in the clinically relevant range (0~2 mM) in an off-body setup. To the best of our knowledge, this is the first time NEXAR has been used for biosensing applications. Further, the developed glucose sensor has acceptable sensitivity of 16.8 microA/mM/cm(2). Importantly, in an on-body setup, the device achieved a significant amperometric response to sweat glucose in a very short amount of time (a few seconds). With detailed investigations, this proof-of-concept study could help further the development of sensitive and selective sweat-based glucose sensing devices for real-time glucose monitoring in diabetes patients. | 44884 | 2022
INWRMTEV | 10.1039/d4ra00110a | https://www.ncbi.nlm.nih.gov/pubmed/38414985 | YES | Electrochemical sensor | microneedle | Minimally | real-time, continuous monitoring of dopamine levels | electrochemical | Electrochemical | The limit of detection (LOD) of the developed sensor was observed to be low, 90 nM in square wave voltammetry and 0.6 μM in chronoamperometric analysis. |  | NA | - | - | yes | interstitial fluid | [12 empty] | Dopamine |  |  |  |  | 1 | 1 | Prototype/Validation |  | In-vitro | NA | NA | A wearable microneedle-based electrochemical sensor, to continuously monitor dopamine in interstitial fluid (ISF). A chitosan-protected hybrid nanomaterial Fe3O4-GO composite has been used as a chemical recognition element protected by Nafion antifouling coating layer. | NO | NA | NA | NA | 2024 | M, R. K.; Panicker, L. R.; Narayan, R.; Kotagiri, Y. G. | Biopolymer-protected graphene-Fe(3)O(4) nanocomposite based wearable microneedle sensor: toward real-time continuous monitoring of dopamine | RSC Adv |  | 2046-2069 (Electronic) 2046-2069 (Linking) | Neurological disorders can occur in the human body as a result of nano-level variations in the neurotransmitter levels. Patients affected by neuropsychiatric disorders, that are chronic require continuous monitoring of these neurotransmitter levels for effective disease management. The current work focus on developing a highly sensitive and personalized sensor for continuous monitoring of dopamine. Here we propose a wearable microneedle-based electrochemical sensor, to continuously monitor dopamine in interstitial fluid (ISF). A chitosan-protected hybrid nanomaterial Fe(3)O(4)-GO composite has been used as a chemical recognition element protected by Nafion antifouling coating layer. The morphological and physiochemical characterizations of the nanocomposite were carried out with XRD, XPS, FESEM, EDAX and FT-IR. The principle of the developed sensor relies on orthogonal detection of dopamine with square wave voltammetry and chronoamperometric techniques. The microneedle sensor array exhibited an attractive analytical performance toward detecting dopamine in phosphate buffer and artificial ISF. The limit of detection (LOD) of the developed sensor was observed to be low, 90 nM in square wave voltammetry and 0.6 muM in chronoamperometric analysis. The practical applicability of the microneedle sensor array has been demonstrated on a skin-mimicking phantom gel model. The microneedle sensor also exhibited good long-term storage stability, reproducibility, and sensitivity. All of these promising results suggest that the proposed microneedle sensor array could be reliable for the continuous monitoring of dopamine. | 45343 | 2024
8YAPSY3W | 10.3390/s23239473 | https://www.ncbi.nlm.nih.gov/pubmed/38067846 | YES | chest band |  | NO | real-time monitoring of athletes' physiological status and performance | electrochemical | Electrochemical | The lactate sensor used in this device is an enzyme-based amperometric sensor that combines the specificity and selectivity of enzymes with the high sensitivity of electrochemical sensing. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate, electrolytes (e.g. Na+) for characterization. NOT WELL SPECIFIED |  |  |  |  | 1 | multiple? NOT WELL SPECIFIED) | Validation | Human | In-vitro + in-vivo (human) | NA | NA | a wearable device designed for continuous athlete monitoring in real time through sweat analysis. This innovative system utilizes a conventional sports band that incorporates a single-use cartridge capable of collecting, processing, and analyzing sweat from the user’s skin, all thanks to an array of sensors integrated into this disposable. | Onalabs Inno-Hub SL | Spain | YES | NA | 2023 | Aguilar-Toran, J.; Rabost-Garcia, G.; Toinga-Villafuerte, S.; Alvarez-Carulla, A.; Colmena-Rubil, V.; Fajardo-Garcia, A.; Cardona-Bonet, A.; Casals-Terre, J.; Munoz-Pascual, X.; Miribel-Catala, P.; Punter-Villagrasa, J. | Novel Sweat-Based Wearable Device for Advanced Monitoring of Athletic Physiological Biometrics | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | Blood testing has traditionally been the gold standard for the physiological analysis and monitoring of professional athletes. In recent years, blood testing has moved out of the laboratory thanks to portable handheld devices, such as lactate meters. However, despite its usefulness and widespread use, blood testing has several drawbacks and limitations, such as the need for the athlete to stop exercising for blood extraction and the inability to have data continuously collected. In this scenario, sweat has become an alternative to blood testing because of its rich content of electrolytes and metabolites, as well as small quantities of sugars, proteins, and ions. Nevertheless, there are few devices capable of analyzing this biofluid and providing useful information to users. In this paper, an electronic system designed for the autonomous analysis of sweat electrolytes and metabolites along with heart rate dynamics is presented. This system is part of a novel wearable device tailored for athletes that offers to the user a real-time assessment of their physiological status and performance. | 45258 | 2023
NNP742K2 | 10.1039/d1an00283j | https://www.ncbi.nlm.nih.gov/pubmed/33999074 | YES | not specified |  | NO | detection of glucose levels in sweat | Colorimetric | Colorimetric | LOD of 0.01 mM and a dynamic range up to 0.15 mM. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Validation | Human | Human studies | NA | NA | Colorimetric wearable biosensors that detect glucose in sweat with a low LOD. The devices are made of filter paper and incorporate a sweat volume sensor and a color chart for signal correction. The biosensors do not suffer from interferences originated by delayed sample readings, or differences in bending angle and sample pH. | NO | NA | YES | NA | 2021 | Vaquer, A.; Baron, E.; de la Rica, R. | Detection of low glucose levels in sweat with colorimetric wearable biosensors | Analyst |  | 1364-5528 (Electronic) 0003-2654 (Linking) | Low glucose levels during exercise may lead to hypoglycemia, which can have grave consequences in diabetic athletes. Mobile colorimetric wearable biosensors that measure glucose levels in sweat are ideal for self-monitoring as they can utilize the camera in smartphones for signal reading. However, colorimetric biosensors proposed thus far have higher limit of detection (LOD) than electrochemical devices, which makes them unsuitable for detecting hypoglycemia. In this manuscript we describe colorimetric wearable biosensors that detect glucose in sweat with an LOD of 0.01 mM and a dynamic range up to 0.15 mM. The devices are made of filter paper and incorporate a sweat volume sensor and a color chart for signal correction. The biosensors do not suffer from interferences originated by delayed sample readings, or differences in bending angle and sample pH. When applied to volunteers performing an exercise routine, sweat glucose levels corrected with sweat volume measurements correlated well with blood glucose measurements performed with a commercial device. The devices are lightweight and easily disposable. These features, along with the smartphone-based colorimetric readout, makes them promising as "over-the-counter" tests for measuring glucose levels non-invasively during exercise. | 44337 | 2021
6EDNGGXD | 10.1021/acs.analchem.2c01052 | https://www.ncbi.nlm.nih.gov/pubmed/35486709 | YES | skin-attached |  | NO | continuous monitoring of sweat conductivity and collection of sweat samples for offline analysis | electrochemical | Electrochemical | A strong linear relationship (R² = 0.97) between the patch's conductivity readings and the laboratory measurements of Na+ and Cl- concentrations |  | NA | yes | - | - | sweat | [12 empty] | Na+, Cl- |  |  |  |  | 1 | 2 | Validation | Human | Human studies | NA | NA | An easy-to-fabricate sweat collection system that allows for continuous electrochemical monitoring, as well as chronological sampling of sweat for offline analysis.The patch consists of ananalysis chamber hosting aconductivity sensor and a sequence of 5 to 10 reservoirs that contain level indicators that monitor the filling speed. | NO | NA | NA | NA | 2022 | Steijlen, A. S. M.; Jansen, K. M. B.; Bastemeijer, J.; French, P. J.; Bossche, A. | Low-Cost Wearable Fluidic Sweat Collection Patch for Continuous Analyte Monitoring and Offline Analysis | Anal Chem |  | 1520-6882 (Electronic) 0003-2700 (Print) 0003-2700 (Linking) | Sweat sensors allow for new unobtrusive ways to continuously monitor an athlete's performance and health status. Significant advances have been made in the optimization of sensitivity, selectivity, and durability of electrochemical sweat sensors. However, comparing the in situ performance of these sensors in detail remains challenging because standardized sweat measurement methods to validate sweat sensors in a physiological setting do not yet exist. Current collection methods, such as the absorbent patch technique, are prone to contamination and are labor-intensive, which limits the number of samples that can be collected over time for offline reference measurements. We present an easy-to-fabricate sweat collection system that allows for continuous electrochemical monitoring, as well as chronological sampling of sweat for offline analysis. The patch consists of an analysis chamber hosting a conductivity sensor and a sequence of 5 to 10 reservoirs that contain level indicators that monitor the filling speed. After testing the performance of the patch in the laboratory, elaborate physiological validation experiments (3 patch locations, 6 participants) were executed. The continuous sweat conductivity measurements were compared with laboratory [Na(+)] and [Cl(-)] measurements of the samples, and a strong linear relationship (R(2) = 0.97) was found. Furthermore, sweat rate derived from ventilated capsule measurement at the three locations was compared with patch filling speed and continuous conductivity readings. As expected from the literature, sweat conductivity was linearly related to sweat rate as well. In short, a successfully validated sweat collection patch is presented that enables sensor developers to systematically validate novel sweat sensors in a physiological setting. | 44691 | 2022
34H6UIRI | 10.1186/s12951-019-0480-4 | https://www.ncbi.nlm.nih.gov/pubmed/30914060 | YES | not specified |  | NO | sweat monitoring | Conductivity | Electrical conductivity | The achieved sensors demonstrate good performance of moisture monitoring and sensitive salt responses with stability and reproducibility. The output current signal increases in a monotonous manner when rising the relative humidity, and the largest responsivity is 244% which is acquired at 80% RH in dark. T he goodness of fit R2 by linear line for the resistancehumidity curve of the sensor is 0.9779 in the dark. Upon illumination, a better linear R2 fitting is equal to 0.9948. |  | NA | yes | - | - | sweat | [12 empty] | water content (moisture) | NaCl |  |  |  | 1 | 2 | validation | Human | Human studies | NA | NA | The sensor utilizes a unique composition-graded CdS1-xSex (x = 0-1) nanowire chip. This means that the composition of the CdSSe nanowires varies along their length, creating a gradient of material properties. | NO | NA | NA | NA | 2019 | Zhang, M.; Guo, S.; Weller, D.; Hao, Y.; Wang, X.; Ding, C.; Chai, K.; Zou, B.; Liu, R. | CdSSe nanowire-chip based wearable sweat sensor | J Nanobiotechnology |  | 1477-3155 (Electronic) 1477-3155 (Linking) | BACKGROUND: Sweat, as an easily accessible bodily fluid, is enriched with a lot of physiological and health information. A portable and wearable sweat sensor is an important device for an on-body health monitoring. However, there are only few such devices to monitor sweat. Based on the fact that sweat is mainly composed of moisture and salt which is much more abundant than other trace ions in sweat, a new route is proposed to realize wearable sweat sensors using CdSSe nanowire-chips coated with a polyimide (PI) membrane. RESULTS: Firstly, the composition-graded CdS(1-x)Se(x) (x = 0-1) nanowire-chip based sensor shows good photo-sensitivity and stress sensitivity which induces linear humidity dependent conductivity. This indicates good moisture response with a maximum responsivity (dI/I) 244% at 80% relative humidity (RH) even in the dark. Furthermore, the linear current decrease with salt increase illustrates the chip sensor has a good salt-sensing ability with the best salt dependent responsivity of 80%, which guarantees the high prediction accuracy in sweat sensing. The sensor current is further proven to nonlinearly correlate to the amount of sweat with excellent stability, reproducibility and recoverability. The wearable sweat sensor is finally applied on-body real-time sweat analysis, showing good consistence with the body status during indoor exercise. CONCLUSIONS: These results suggest that this CdSSe nanowire-chip based PI-coated integrated sensor, combined with inorganic and organic functional layers, provides a simple and reliable method to build up diverse portable and wearable devices for the applications on healthcare and athletic status. | 43550 | 2019
RTVB98DV | 10.1039/d1an02214h | https://www.ncbi.nlm.nih.gov/pubmed/35262099 | YES | skin-attached |  | NO | monitoring of glucose levels in human sweat | electrochemical | Electrochemical | The sensor exhibits a wide linear range of 20 μM to 790 μM for glucose detection. The sensor boasts a high sensitivity of 205.1 μA mM-1 cm-2. | Diabetes | Diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation | Human | Human studies | NA | NA | A wearable electrochemical sweat sensor based on a Ni-Co MOF nanosheet coated Au/polydimethylsiloxane (PDMS) film was prepared for continuous monitoring of the glucose level in sweat with high sensitivity. | NO | NA | NA | NA | 2022 | Shu, Y.; Shang, Z.; Su, T.; Zhang, S.; Lu, Q.; Xu, Q.; Hu, X. | A highly flexible Ni-Co MOF nanosheet coated Au/PDMS film based wearable electrochemical sensor for continuous human sweat glucose monitoring | Analyst |  | 1364-5528 (Electronic) 0003-2654 (Linking) | The development of flexible substrate materials and nanomaterials with high electrochemical performance is of great significance for constructing efficient wearable electrochemical sensors for real-time health monitoring. Herein, a wearable electrochemical sweat sensor based on a Ni-Co MOF nanosheet coated Au/polydimethylsiloxane (PDMS) film was prepared for continuous monitoring of the glucose level in sweat with high sensitivity. First, a stretchable Au/PDMS film based three-electrode system was prepared by chemical deposition of a gold layer on the hydrophilic treated PDMS. Then, Ni-Co MOF nanosheets with high electrocatalytic activity were synthesized by a facile solvothermal method and modified on the Au/PDMS electrode. The electrocatalytic activity of the Ni-Co MOF nanosheets synthesized under different Ni : Co ratios was investigated. The Ni-Co MOF/Au/PDMS (NCAP) film electrode showed excellent electrochemical performance for glucose detection with a wide linear range of 20 muM to 790 muM and a high sensitivity of 205.1 muA mM(-1) cm(-2). In addition, the flexible sensor shows high stability and a good electrochemical response to glucose when stretched and bent to different levels. Moreover, it maintained long-term stability and high selectivity for glucose monitoring. Lastly, a sweat-absorbent cloth was used to cover the working area of the sensor and was fixed with a needle and thread to form a wearable sweat glucose sensor. The sensor can be attached to the skin for stable, accurate and continuous monitoring of glucose levels in human sweat for one day. This work validates the potential of our high-performance wearable sensor for out-of-clinic health monitoring. | 44648 | 2022
ALE8BTQJ | 10.1002/smll.202311802 | https://onlinelibrary.wiley.com/doi/10.1002/smll.202311802 | YES | skin-attached |  | NO | noninvasive, real-time monitoring of potassium ion (K+) concentration in human sweat | electrochemical | Electrochemical | The source states the wearable K+ sensor has a high sensitivity of 62 mV decade-1. This means the sensor can detect subtle changes in K+ concentration and is therefore very precise. The sensor also has a low detection limit of 0.05 mm. These values, combined with the sensor's good selectivity and stability, indicate the device is capable of collecting accurate and reliable K+ concentration data. |  | NA | yes | - | - | sweat | [12 empty] | K+ |  |  |  |  | 1 | 1 | Validation | Human | Laboratory testing + human studies (preliminary) | NA | NA | The main cutting-edge innovation in the wearable K+ sensor is the tailored design and fabrication of polypyrrole (PPy) nanofibers to serve as the ion-to-electron transduction layer. This nanofiber structure offers several advantages over traditional conductive polymer membranes, enhancing the sensor's performance. | NO | NA | NA | NA | 2024 | Yang, Yaqiong; Lv, Tian‐Run; Zhang, Wen‐Hai; Zhang, Jia‐Yue; Yin, Ming‐Jie; An, Quan‐Fu | Tailored Polypyrrole Nanofibers as Ion‐to‐Electron Transduction Membranes for Wearable K<sup>+</sup> Sensors | Small |  | 1613-6810, 1613-6829 | Abstract Conductive polymers are recognized as ideal candidates for the development of noninvasive and wearable sensors for real‐time monitoring of potassium ions (K + ) in sweat to ensure the health of life. However, the low ion‐to‐electron transduction efficiency and limited active surface area hamper the development of high‐performance sensors for low‐concentration K + detection in the sweat. Herein, a wearable K + sensor is developed by tailoring the nanostructure of polypyrrole (PPy), serving as an ion‐to‐electron transduction layer, for accurately and stably tracing the K + fluctuation in human sweat. The PPy nanostructures can be tailored from nanospheres to nanofibers by controlling the supramolecular assembly process during PPy polymerization. Resultantly, the ion‐to‐electron transduction efficiency (17‐fold increase in conductivity) and active surface area (1.3‐fold enhancement) are significantly enhanced, accompanied by minimized water layer formation. The optimal PPy nanofibers‐based K + sensor achieved a high sensitivity of 62 mV decade −1 , good selectivity, and solid stability. After being integrated with a temperature sensor, the manufactured wearable sensor realized accurate monitoring of K + fluctuation in the human sweat. | 45444 | 2024
WK6BCNCY | 10.3390/bios13010105 | https://www.ncbi.nlm.nih.gov/pubmed/36671938 | YES | common rubber glove |  | NO | simultaneous detection of uric acid and glucose levels in sweat | electrochemical | Electrochemical | The sensors exhibited the linear range from 0 to 1.6 mM and 0 to 3.7 mM towards uric acid and glucose electrochemical sensing in phosphate-buffered solution, with the corresponding limit of detection of 3.58 μM and 9.10 μM obtained, respectively. Moreover, the sensors had shown their feasibility of real sample sensing in sweat. The linear detection range for uric acid (0 to 40 μM) and glucose (0 to 1.6 mM) in the sweat can well cover their concentration range in physiological conditions. |  | NA | yes | - | - | sweat |  |  |  |  | yes |  |  | yes |  |  |  |  | uric acid, glucose |  |  |  |  | 1 | 2 | Proof-of-concept/prototype |  | In-vitro tests | NA | NA | A dual-function wearable electrochemical sensor for uric acid and glucose detection in sweat. The sensor with a four-electrode system was prepared by printing the ink on a common rubber glove. CV and chronoamperometry were used to characterize the prepared sensor's electrochemical sensing performance. | NO | NA | NA | NA | 2023 | Li, Z.; Wang, Y.; Fan, Z.; Sun, Y.; Sun, Y.; Yang, Y.; Zhang, Y.; Ma, J.; Wang, Z.; Zhu, Z. | A Dual-Function Wearable Electrochemical Sensor for Uric Acid and Glucose Sensing in Sweat | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | Simultaneous detection of uric acid and glucose using a non-invasive approach can be a promising strategy for related diseases, e.g., diabetes, gout, kidney disease, and cardiovascular disease. In this study, we have proposed a dual-function wearable electrochemical sensor for uric acid and glucose detection in sweat. The sensor with a four-electrode system was prepared by printing the ink on a common rubber glove. CV and chronoamperometry were used to characterize the prepared sensor's electrochemical sensing performance. The sensors exhibited the linear range from 0 to 1.6 mM and 0 to 3.7 mM towards uric acid and glucose electrochemical sensing in phosphate-buffered solution, with the corresponding limit of detection of 3.58 muM and 9.10 muM obtained, respectively. Moreover, the sensors had shown their feasibility of real sample sensing in sweat. The linear detection range for uric acid (0 to 40 muM) and glucose (0 to 1.6 mM) in the sweat can well cover their concentration range in physiological conditions. The prepared dual-function wearable electrochemical sensor features easy preparation, fast detection, high sensitivity, high selectivity, and the practical application potential in uric acid and glucose sensing. | 44932 | 2023
4XGQGTWE | 10.1021/acs.analchem.9b03110 | https://www.ncbi.nlm.nih.gov/pubmed/31553565 | YES | skin-attached |  | NO | monitoring of glucose levels in sweat | Optical | Optical | The device demonstrates a linear response to glucose concentrations ranging from 0.1 to 0.5 mM, with a detection limit of 0.03 mM. |  |  | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | Colorimetric detection method using the GOD-peroxidase-o-dianisidine system is presented as a more sensitive approach compared to the conventional GOD-peroxidase-KI system. | NO | NA | NA | NA | 2019 | Xiao, J.; Liu, Y.; Su, L.; Zhao, D.; Zhao, L.; Zhang, X. | Microfluidic Chip-Based Wearable Colorimetric Sensor for Simple and Facile Detection of Sweat Glucose | Anal Chem |  | 1520-6882 (Electronic) 0003-2700 (Linking) | This study reports a microfluidic chip-based wearable colorimetric sensor for detecting sweat glucose. The device consisted of five microfluidic channels branching out from the center and connected to the detection microchambers. The microchannels could route the sweat excreted from the epidermis to the microchambers, and each of them was integrated with a check valve to avoid the risk of the backflow of the chemical reagents from the microchamber. The microchambers contained the pre-embedded glucose oxidase (GOD)-peroxidase-o-dianisidine reagents for sensing the glucose in sweat. It was found that the color change caused by the enzymatic oxidation of o-dianisidine could show a more sensitive response to the glucose than that of the conventional GOD-peroxidase-KI system. This sensor could perform five parallel detections at one time. The obtained linear range for sweat glucose was 0.1-0.5 mM with a limit of detection of 0.03 mM. The sensor was also used to detect the glucose in sweat samples from a group of subjects engaged in both fasting and postprandial trials. The results showed that our wearable colorimetric sensor can reveal the subtle differences existing in the sweat glucose concentration after the fasting and the oral glucose uptake. | 43802 | 2019
VDCRBS5S | 10.1021/acsami.1c11446 | https://www.ncbi.nlm.nih.gov/pubmed/34340308 | YES | skin-attached |  | NO | monitoring sweat | Optical | Optical | The responses of both biosensors showed good linearity, between 0.1 and 1.0 mM, for lactate and between 0.1 and 0.8 mM for glucose. The dependency between the sensor response and lactate or glucose concentration can be approximated by a linear function with a correlation factor R2 of 0.981 for lactate and 0.989 for glucose. The statistical limit of detection (LOD) was 0.069 mM with a limit of quantification (LOQ) = 0.23 mM for lactate. In the case of glucose, LOD = 0.044 mM and LOQ = 0.15 mM (LOD = 3S/K, LOQ = 10S/K, where S is the standard deviation of the blank sample and K is the slope of the calibration curve). |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | glucose, lactate |  |  |  |  | 1 | 2 | Proof-of-concept |  | In-vitro tests | NA | NA | A novel three-dimensional titanium dioxide nanotubes/alginate hydrogel scaffold is proposed for the detection of sweat biomarkers, lactate, and glucose in artificial sweat. | NO | NA | NA | NA | 2021 | Gunatilake, U. B.; Garcia-Rey, S.; Ojeda, E.; Basabe-Desmonts, L.; Benito-Lopez, F. | TiO(2) Nanotubes Alginate Hydrogel Scaffold for Rapid Sensing of Sweat Biomarkers: Lactate and Glucose | ACS Appl Mater Interfaces |  | 1944-8252 (Electronic) 1944-8244 (Print) 1944-8244 (Linking) | Versatile sensing matrixes are essential for the development of enzyme-immobilized optical biosensors. A novel three-dimensional titanium dioxide nanotubes/alginate hydrogel scaffold is proposed for the detection of sweat biomarkers, lactate, and glucose in artificial sweat. Hydrothermally synthesized titanium dioxide nanotubes were introduced to the alginate polymeric matrix, followed by cross-linking nanocomposite with dicationic calcium ions to fabricate the scaffold platform. Rapid colorimetric detection (blue color optical signal) was carried out for both lactate and glucose biomarkers in artificial sweat at 4 and 6 min, respectively. The superhydrophilicity and the capillarity of the synthesized titanium dioxide nanotubes, when incorporated into the alginate matrix, facilitate the rapid transfer of the artificial sweat components throughout the sensor scaffold, decreasing the detection times. Moreover, the scaffold was integrated on a cellulose paper to demonstrate the adaptability of the material to other matrixes, obtaining fast and homogeneous colorimetric detection of lactate and glucose in the paper substrate when image analysis was performed. The properties of this new composite provide new avenues in the development of paper-based sensor devices. The biocompatibility, the efficient immobilization of biological enzymes/colorimetric assays, and the quick optical signal readout behavior of the titanium dioxide nanotubes/alginate hydrogel scaffolds provide a prospective opportunity for integration into wearable devices. | 44419 | 2021
5IXMVHDJ | 10.1038/s41598-017-02133-0 | https://www.ncbi.nlm.nih.gov/pubmed/28512341 | YES | not specified | nanoporous, flexible polymer membranes | NO | detection of analytes | Immuneassay | Immunoassay | IL-6 over 0.2–200 pg/ mL in human sweat 10 hours post-antibody sensor functionalization. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | IL-6, cortisol |  |  |  |  | 1 | 2 | prototype |  | Laboratory testing + in-vitro studies | NA | NA | Sensors on nanoporous polymer membranes antibody-functionalized with room temperature ionic liquids (RTILs) | NO | NA | NA | NA | 2017 | Munje, R. D.; Muthukumar, S.; Jagannath, B.; Prasad, S. | A new paradigm in sweat based wearable diagnostics biosensors using Room Temperature Ionic Liquids (RTILs) | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | Successful commercialization of wearable diagnostic sensors necessitates stability in detection of analytes over prolonged and continuous exposure to sweat. Challenges are primarily in ensuring target disease specific small analytes (i.e. metabolites, proteins, etc.) stability in complex sweat buffer with varying pH levels and composition over time. We present a facile approach to address these challenges using RTILs with antibody functionalized sensors on nanoporous, flexible polymer membranes. Temporal studies were performed using both infrared spectroscopic, dynamic light scattering, and impedimetric spectroscopy to demonstrate stability in detection of analytes, Interleukin-6 (IL-6) and Cortisol, from human sweat in RTILs. Temporal stability in sensor performance was performed as follows: (a) detection of target analytes after 0, 24, 48, 96, and 168 hours post-antibody sensor functionalization; and (b) continuous detection of target analytes post-antibody sensor functionalization. Limit of detection of IL-6 in human sweat was 0.2 pg/mL for 0-24 hours and 2 pg/mL for 24-48 hours post-antibody sensor functionalization. Continuous detection of IL-6 over 0.2-200 pg/mL in human sweat was demonstrated for a period of 10 hours post-antibody sensor functionalization. Furthermore, combinatorial detection of IL-6 and Cortisol in human sweat was established with minimal cross-talk for 0-48 hours post-antibody sensor functionalization. | 42871 | 2017
GKNF76ZE | 10.1016/j.bios.2016.09.038 | https://www.ncbi.nlm.nih.gov/pubmed/27743863 | YES | patch |  | NO | continuous, non-invasive monitoring of biomarkers in sweat | electrochemical | Electrochemical | Na+ and pH sensors had relative standard deviations (R.S.D.) ~0.6% and the lactate sensor has RSD of ~4% in sensitivity. Sensitivity of the pH sensor is 71.90±0.8 mV/unit and for [Na+]is 56±1mV/unit. Long-term studies showed stability and repeatability of over 8 months for the pH sensor, over 5 weeks for the sodium sensor and over 4 weeks for the enzymatic sensor if kept properly. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes |  |  |  |  | yes |  | lactate, K+, Glucose, pH |  |  |  |  | 1 | 4 | prototype | Human | Human studies | NA | NA | Microfluidic platform, fully integrated. Wireless for data transmission. | NO | NA | NA | NA | 2017 | Anastasova, S.; Crewther, B.; Bembnowicz, P.; Curto, V.; Ip, H. M.; Rosa, B.; Yang, G. Z. | A wearable multisensing patch for continuous sweat monitoring | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Linking) | In sport, exercise and healthcare settings, there is a need for continuous, non-invasive monitoring of biomarkers to assess human performance, health and wellbeing. Here we report the development of a flexible microfluidic platform with fully integrated sensing for on-body testing of human sweat. The system can simultaneously and selectively measure metabolite (e.g. lactate) and electrolytes (e.g. pH, sodium) together with temperature sensing for internal calibration. The construction of the platform is designed such that continuous flow of sweat can pass through an array of flexible microneedle type of sensors (50microm diameter) incorporated in a microfluidic channel. Potentiometric sodium ion sensors were developed using a polyvinyl chloride (PVC) functional membrane deposited on an electrochemically deposited internal layer of Poly(3,4-ethylenedioxythiophene) (PEDOT) polymer. The pH sensing layer is based on a highly sensitive membrane of iridium oxide (IrOx). The amperometric-based lactate sensor consists of doped enzymes deposited on top of a semipermeable copolymer membrane and outer polyurethane layers. Real-time data were collected from human subjects during cycle ergometry and treadmill running. A detailed comparison of sodium, lactate and cortisol from saliva is reported, demonstrating the potential of the multi-sensing platform for tracking these outcomes. In summary, a fully integrated sensor for continuous, simultaneous and selective measurement of sweat metabolites, electrolytes and temperature was achieved using a flexible microfluidic platform. This system can also transmit information wirelessly for ease of collection and storage, with the potential for real-time data analytics. | 42931 | 2017
YCWTJYK2 | 10.1002/cphc.201701312 | https://chemistry-europe.onlinelibrary.wiley.com/doi/10.1002/cphc.201701312 | YES | wrist-worn | smartwatch-like | NO | real-time monitoring of sodium levels in sweat. | electrochemical | Electrochemical | Reproducibility: Consistent responses across calibrations, with similar results for both single and batch ISE-REs. Linearity: Near-Nernstian slope of ~57 mV/decade ensures accurate sodium quantification over a wide range. Rapid Response: Stabilization within seconds allows real-time monitoring of dynamic sodium level changes. Calibration Stability: Maintains stable response characteristics during on-body trials, validated through 4-point calibrations. The platform was characterised in the lab and in exercise trials over a period of about 60 minutes continuous monitoring. Sweat sodium concentration was found to rise initially to approximately 17 mM and decline gradually over the period of the trial to about 11–12 mM. |  | NA | yes | - | - | sweat | [12 empty] | Na+ |  |  |  |  | 1 | 1 | Validation | Human | laboratory testing + human studies | NA | NA | 3D Printing: Simplifies and accelerates fabrication while enabling precise component alignment and a larger reservoir for extended monitoring. Thread Microfluidics: Employs passive sweat wicking through hydrophilic threads, allowing tunable flow rates without external pumps for energy efficiency. Compact ISE-RE Design: Combines solid-state ion-selective and reference electrodes into a miniaturized unit with a PEDOT layer, enhancing sensitivity, stability, and biocompatibility. Wireless Connectivity: Utilizes Bluetooth for real-time data transmission and analysis, improving user convenience. | NO | NA | NA | NA | 2018 | McCaul, Margaret; Porter, Adam; Barrett, Ruairi; White, Paddy; Stroiescu, Florien; Wallace, Gordon; Diamond, Dermot | Wearable Platform for Real‐time Monitoring of Sodium in Sweat | ChemPhysChem |  | 1439-4235, 1439-7641 | Abstract A fully integrated and wearable platform for harvesting and analysing sweat sodium concentration in real time during exercise has been developed and tested. The platform was largely produced using 3D printing, which greatly simplifies fabrication and operation compared to previous versions generated with traditional production techniques. The 3D printed platform doubles the capacity of the sample storage reservoir to about 1.3 ml, reduces the assembly time and provides simple and precise component alignment and contact of the integrated solid‐state ion‐selective and reference electrodes with the sorbent material. The sampling flowrate in the device can be controlled by introducing threads to enhance wicking of sweat from the skin, across the electrodes to the storage area. The platform was characterised in the lab and in exercise trials over a period of about 60 minutes continuous monitoring. Sweat sodium concentration was found to rise initially to approximately 17 mM and decline gradually over the period of the trial to about 11–12 mM. | 43270 | 2018
ZVVI72SE | 10.1021/acsami.4c10033 | https://www.ncbi.nlm.nih.gov/pubmed/39178237 | YES | skin-attached |  | NO | real-time, noninvasive monitoring of health-associated biochemical parameters in sweat | electrochemical | Electrochemical | The molecularly imprinted biosensor can achieve sensitive and specific detection of cortisol at concentrations as low as 1 pM, 1000-fold lower than previously reported MIP cortisol sensors. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | Cortisol, Na+ |  |  |  |  | 1 | 2 | Proof-of-concept | Human | In-situ (human) | NA | NA | This sensor design utilizes MIPs, which are synthetic polymers with specific recognition sites for target molecules. In this case, the MIP is designed to bind selectively to cortisol, enabling highly sensitive and specific detection. This approach eliminates the need for labels or redox probes commonly used in other electrochemical sensors, enhancing the sensitivity and simplifying the detection process | NO | NA | NA | NA | 2024 | Garg, M.; Guo, H.; Maclam, E.; Zhanov, E.; Samudrala, S.; Pavlov, A.; Rahman, M. S.; Namkoong, M.; Moreno, J. P.; Tian, L. | Molecularly Imprinted Wearable Sensor with Paper Microfluidics for Real-Time Sweat Biomarker Analysis | ACS Appl Mater Interfaces |  | 1944-8252 (Electronic) 1944-8244 (Print) 1944-8244 (Linking) | The urgent need for real-time and noninvasive monitoring of health-associated biochemical parameters has motivated the development of wearable sweat sensors. Existing electrochemical sensors show promise in real-time analysis of various chemical biomarkers. These sensors often rely on labels and redox probes to generate and amplify the signals for the detection and quantification of analytes with limited sensitivity. In this study, we introduce a molecularly imprinted polymer (MIP)-based biochemical sensor to quantify a molecular biomarker in sweat using electrochemical impedance spectroscopy, which eliminates the need for labels or redox probes. The molecularly imprinted biosensor can achieve sensitive and specific detection of cortisol at concentrations as low as 1 pM, 1000-fold lower than previously reported MIP cortisol sensors. We integrated multimodal electrochemical sensors with an iontophoresis sweat extraction module and paper microfluidics for real-time sweat analysis. Several parameters can be simultaneously quantified, including sweat volume, secretion rate, sodium ion, and cortisol concentration. Paper microfluidic modules not only quantify sweat volume and secretion rate but also facilitate continuous sweat analysis without user intervention. While we focus on cortisol sensing as a proof-of-concept, the molecularly imprinted wearable sensors can be extended to real-time detection of other biochemicals, such as protein biomarkers and therapeutic drugs. | 45539 | 2024
SWC7JXLD | 10.1002/adma.201707442 | https://www.ncbi.nlm.nih.gov/pubmed/29663538 | YES | patch |  | NO | drug monitoring | electrochemical | Electrochemical | There is a linear relationship of the sensor’s responses to the caffeine concentrations with a highsensitivity of 110 nA mm−1 |  | NA | yes | - | - | sweat | [12 empty] | Methylxanthine (Caffeine) |  |  |  |  | 1 | 1 | prototype | Human | Human studies | NA | NA | Electrochemical Differential Pulse Voltammetry Sensing Module | NO | NA | NA | NA | 2018 | Tai, L. C.; Gao, W.; Chao, M.; Bariya, M.; Ngo, Q. P.; Shahpar, Z.; Nyein, H. Y. Y.; Park, H.; Sun, J.; Jung, Y.; Wu, E.; Fahad, H. M.; Lien, D. H.; Ota, H.; Cho, G.; Javey, A. | Methylxanthine Drug Monitoring with Wearable Sweat Sensors | Adv Mater |  | 1521-4095 (Electronic) 0935-9648 (Linking) | Drug monitoring plays crucial roles in doping control and precision medicine. It helps physicians tailor drug dosage for optimal benefits, track patients' compliance to prescriptions, and understand the complex pharmacokinetics of drugs. Conventional drug tests rely on invasive blood draws. While urine and sweat are attractive alternative biofluids, the state-of-the-art methods require separate sample collection and processing steps and fail to provide real-time information. Here, a wearable platform equipped with an electrochemical differential pulse voltammetry sensing module for drug monitoring is presented. A methylxanthine drug, caffeine, is selected to demonstrate the platform's functionalities. Sweat caffeine levels are monitored under various conditions, such as drug doses and measurement time after drug intake. Elevated sweat caffeine levels upon increasing dosage and confirmable caffeine physiological trends are observed. This work leverages a wearable sweat sensing platform toward noninvasive and continuous point-of-care drug monitoring and management. | 43252 | 2018
RKP65NT7 | 10.1007/s00216-022-04510-9 | https://www.ncbi.nlm.nih.gov/pubmed/36622394 | YES | textile |  | NO | real-time sweat analysis and the prediction of physiological parameters. | electrochemical | Electrochemical | The all-in-one sensor integrated all the functions exhibite UA sensitivity of 0.62 ± 0.1 nA/µM and Na+ ion sensitivity of 43.1 ± 3.5 mV/decade in the artificial sweat, the performance of which can meet requirements for on-body applications. The fluid delivery capability of its internal microfluidic channel for cleaning the sensing surface has also been confirmed. In addition, all the functional components on the fiber-based sensor show excellent longterm stability in continuous testing for up to 1 or 2 months. |  | NA | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | Na+, Uric Acid (UA) |  |  |  |  | 1 | 2 | Prototype/Validation |  | Laboratory testing + in-vitro tests | NA | NA | A flexible microelectronic fiber with bioanalytical functions that could be woven into textiles as a new form of wearable bioelectronics. Within a single strand of fiber, we successfully integrated all-in-one multiplexed electrochemical sensing capabilities, with the sweat as the primary object. Adopting the laser micromachining technique, we developed biosensing functions on the longitudinal surface of the fiber with two sensing electrodes for Na+ and uric acid (UA), respectively, together with a pseudo reference electrode (p-RE). | NO | NA | NA | NA | 2023 | Wu, J.; Sato, Y.; Guo, Y. | Microelectronic fibers for multiplexed sweat sensing | Anal Bioanal Chem |  | 1618-2650 (Electronic) 1618-2642 (Print) 1618-2642 (Linking) | Wearable bioelectronics are gaining extraordinary attention due to their capabilities to achieve continuous monitoring of human health status. However, mainstream manufacturing technologies, including photolithography and printing technology, limit current wearable bioelectronics on 2D planar structures with little surface area in contact with the body. It thus limits the amount of physiological information that current wearable bioelectronics could obtain. Furthermore, they need to be firmly attached to the body, affecting the wearing comfort. In this study, we leveraged the versatile thermal drawing process and developed a flexible microelectronic fiber with bioanalytical functions that could be woven into textiles as a new form of wearable bioelectronics. Within a single strand of fiber, we successfully integrated all-in-one multiplexed electrochemical sensing capabilities, with the sweat as the primary object. Adopting the laser micromachining technique, we developed biosensing functions on the longitudinal surface of the fiber with two sensing electrodes for Na(+) and uric acid (UA), respectively, together with a pseudo reference electrode (p-RE). We carefully characterized the all-in-one multiplexed sensing performance of the fiber and demonstrated its successful application in sweat sensing based on its textile forms. The results show significant potential for application in wearable textiles for monitoring key health signals of humans. | 45108 | 2023
74YJ6EGN | 10.3390/bios13020184 | https://www.ncbi.nlm.nih.gov/pubmed/36831950 | YES | skin-attached |  | NO | detection of cortisol levels in human sweat- monitoring of stress levels | Optical | localized surface plasmon resonance (LSPR) | The biosensor exhibits a detection limit of 0.1 nM for cortisol. The device was tested in cortisol solutions with concentrations ranging from 0.1 to 1000 nM, demonstrating its ability to quantify cortisol across a wide range of concentrations. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | Prototype/Validation | Human | In-vitro + human studies | NA | NA | A flexible LSPR biosensor was manufactured to detect cortisol levels in the human body by depositing gold nanoparticle (AuNP) layers on a 3-aminopropyltriethoxysilane (APTES)-functionalized poly (dimethylsiloxane) (PDMS) substrate. Subsequently, an aptamer was immobilized on the surface of the LSPR substrate, enabling highly sensitive and selective cortisol capture owing to its specific cortisol recognition. | NO | NA | NA | NA | 2023 | Nan, M.; Darmawan, B. A.; Go, G.; Zheng, S.; Lee, J.; Kim, S.; Lee, T.; Choi, E.; Park, J. O.; Bang, D. | Wearable Localized Surface Plasmon Resonance-Based Biosensor with Highly Sensitive and Direct Detection of Cortisol in Human Sweat | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | Wearable biosensors have the potential for developing individualized health evaluation and detection systems owing to their ability to provide continuous real-time physiological data. Among various wearable biosensors, localized surface plasmon resonance (LSPR)-based wearable sensors can be versatile in various practical applications owing to their sensitive interactions with specific analytes. Understanding and analyzing endocrine responses to stress is particularly crucial for evaluating human performance, diagnosing stress-related diseases, and monitoring mental health, as stress takes a serious toll on physiological health and psychological well-being. Cortisol is an essential biomarker of stress because of the close relationship between cortisol concentration in the human body and stress level. In this study, a flexible LSPR biosensor was manufactured to detect cortisol levels in the human body by depositing gold nanoparticle (AuNP) layers on a 3-aminopropyltriethoxysilane (APTES)-functionalized poly (dimethylsiloxane) (PDMS) substrate. Subsequently, an aptamer was immobilized on the surface of the LSPR substrate, enabling highly sensitive and selective cortisol capture owing to its specific cortisol recognition. The biosensor exhibited excellent detection ability in cortisol solutions of various concentrations ranging from 0.1 to 1000 nM with a detection limit of 0.1 nM. The flexible LSPR biosensor also demonstrated good stability under various mechanical deformations. Furthermore, the cortisol levels of the flexible LSPR biosensor were also measured in the human epidermis before and after exercise as well as in the morning and afternoon. Our biosensors, which combine easily manufactured flexible sensors with sensitive cortisol-detecting molecules to measure human stress levels, could be versatile candidates for human-friendly products. | 44950 | 2023
XGB4JHY2 | 10.1016/j.bios.2020.112782 | https://linkinghub.elsevier.com/retrieve/pii/S0956566320307697 | YES | skin-attached |  | NO | real-time on-body sweat cortisol detection | electrochemical | Electrochemical | Cortisol sensor: the sensor exhibits a linear response to cortisol concentrations ranging from 0 to 500 nM. The limit of detection (LOD) for cortisol is estimated to be ~7.47 nM. This suggests that the cortisol sensor can reliably detect cortisol levels within the typical range found in sweat (22 nM to 386 nM). |  | NA | yes | - | - | sweat |  |  |  |  | yes |  | yes | [5 empty] | Cortisol, Glucose |  |  |  |  | 1 | 2 | Validation | Human | Human studies | NA | NA | The patch is integrated with all-printed flexible electrochemical immunosensor, which was used to detect cortisol through differential pulse voltammetry (DPV). The near field communication (NFC) module on the patch enables wireless power harvesting and data interaction with an NFC-enabled smartphone, which makes the patch get rid of rigid batteries and realize epidermal on-body testing. | NO | NA | YES | NA | 2021 | Cheng, Chen; Li, Xin; Xu, Gang; Lu, Yanli; Low, Sze Shin; Liu, Guang; Zhu, Long; Li, Caidong; Liu, Qingjun | Battery-free, wireless, and flexible electrochemical patch for in situ analysis of sweat cortisol via near field communication | Biosensors and Bioelectronics |  | 9565663 | Wearable and flexible biosensing devices have been widely developed for in situ detections. Cortisol is a vital biomarker which plays crucial regulatory role in numerous physiological processes of the human body. Here, a wireless, battery-free, and flexible integrated patch is developed for real-time on-body sweat cortisol detection. The patch integrated with all-printed flexible electrochemical immunosensor, which was used to detect cortisol through differential pulse voltammetry (DPV). The near field communication (NFC) module on the patch enables wireless power harvesting and data interaction with an NFC-enabled smartphone, which makes the patch get rid of rigid batteries and realize epidermal on-body testing. Multiple in situ detections on volunteers’ sweat on the surface of skin showed that the flexible integrated patch could reflect the circadian rhythm of the body’s sweat cortisol level changes in relaxed mood or under stress, which could be confirmed with the enzyme linked immunosorbent assay (ELISA) kit. In this way, the patch provides a rapid-detecting, convenient, and noninvasive sensing solution for in situ analysis of sweat cortisol, which can be applied for the personalized mental health management. | 44197 | 2021
AXIHAVDI | 10.1038/s41587-019-0321-x | https://www.ncbi.nlm.nih.gov/pubmed/31768044 | YES | patch |  | NO | simultaneous sweat sampling, chemical sensing and vital-sign monitoring | electrochemical | Electrochemical. | We determined the accuracy of the wearable sensor for analyzing raw sweat samples using high-performance liquid chromatography (HPLC) and commercial colorimetric assay kits and obtained high Pearson correlation coefficients of 0.963 and 0.965 for UA and Tyr sensing, respectively, suggesting the reliability of the sensor measurements. |  |  | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | uric acid, tyrosin |  |  |  |  | 1 | 2 | Validation | Human | Human studies | NA | NA | The entire sensor, including the microfluidic channels for sweat sampling and the sensing components, is fabricated using a laser-engraving process | NO | NA | NA | NA | 2020 | Yang, Y.; Song, Y.; Bo, X.; Min, J.; Pak, O. S.; Zhu, L.; Wang, M.; Tu, J.; Kogan, A.; Zhang, H.; Hsiai, T. K.; Li, Z.; Gao, W. | A laser-engraved wearable sensor for sensitive detection of uric acid and tyrosine in sweat | Nat Biotechnol |  | 1546-1696 (Electronic) 1087-0156 (Linking) | Wearable sweat sensors have the potential to provide continuous measurements of useful biomarkers. However, current sensors cannot accurately detect low analyte concentrations, lack multimodal sensing or are difficult to fabricate at large scale. We report an entirely laser-engraved sensor for simultaneous sweat sampling, chemical sensing and vital-sign monitoring. We demonstrate continuous detection of temperature, respiration rate and low concentrations of uric acid and tyrosine, analytes associated with diseases such as gout and metabolic disorders. We test the performance of the device in both physically trained and untrained subjects under exercise and after a protein-rich diet. We also evaluate its utility for gout monitoring in patients and healthy controls through a purine-rich meal challenge. Levels of uric acid in sweat were higher in patients with gout than in healthy individuals, and a similar trend was observed in serum. | 43862 | 2020
AHVPZ34C | 10.1016/j.talanta.2022.123481 | https://linkinghub.elsevier.com/retrieve/pii/S0039914022002776 | YES | skin-attached |  | NO | monitoring of pH and K+ | electrochemical | Electrochemical | pH Sensitivity: the pH sensor exhibits a sensitivity of 54 mV/dec when tested using a traditional electrode setup (β-CD/RGO/GCE). This value is close to the theoretical Nernst response for a pH electrode, indicating a high level of sensitivity to changes in pH. The sensitivity of the flexible pH electrode (β-CD/RGO/CC ink/PET) integrated into the wearable device is slightly lower, ranging from 48 mV/dec to 51 mV/dec under different bending conditions. The variation in sensitivity is attributed to the influence of the flexible substrate and the use of a flexible reference electrode. K+ Sensitivity: The K+-selective electrode (K+-ISM/β-CD/RGO/GCE) demonstrates a sensitivity of 56.0 mV/dec for K+ detection in a traditional setup. This value is also close to the Nernst response, indicating a high degree of sensitivity to changes in K+ concentration. The flexible K+ electrode (K+-ISM/β-CD/RGO/CC ink/PET) in the wearable device maintains a consistent sensitivity of 55 mV/dec under different bending states. This finding suggests that the flexibility of the device does not compromise the sensitivity of the K+ sensor. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | pH, K+ |  |  |  |  | 1 | 2 | validation | Human | Human studies | NA | NA | The sensor was composed of flexible reference electrode, pH response electrode and K+ selective electrode, which were prepared through printing β-CD functionalized graphene (β-CD/RGO) water suspension on conductive PET substrate with microelectronic printer. | NO | NA | NA | NA | 2022 | Cui, Xiaoqing; Bao, Yu; Han, Tingting; Liu, Zhenbang; Ma, Yingming; Sun, Zhonghui | A wearable electrochemical sensor based on β-CD functionalized graphene for pH and potassium ion analysis in sweat | Talanta |  | 399140 | Here, we developed a wearable electrochemical sensor for pH and K+ monitoring in sweat. The sensor was composed of flexible reference electrode, pH response electrode and K+ selective electrode, which were prepared through printing β-CD functionalized graphene (β-CD/RGO) water suspension on conductive PET substrate with microelectronic printer. β-CD/RGO not only served as the pH sensitive material for pH response electrode with good sensitivity, selectivity and reproducibility due to its abundant oxygen-containing functional groups, but also worked as the ion-to-electron transducer for K+ selective electrode with good sensitivity. The wearable sensor exhibited good potential stability at different bending states. On-body sweat pH and K+ measurements showed high accuracy compared with ex-situ analysis. | 44774 | 2022
7GJ6KKK8 | 10.1038/s41598-024-62236-3 | https://www.ncbi.nlm.nih.gov/pubmed/38773136 | YES | skin-attached |  | NO | monitoring the levels of certain ions and pH in human sweat. | electrochemical | Potentiometry | Na+ sensor sensitivity: 59.7 ± 0.8 mV/decade; K+ sensor sensitivity: 57.8 ± 0.9 mV/decade; pH sensor sensitivity: 54.7 ± 0.6 mV/pH |  | NA | yes | - | - | sweat | [10 empty] | yes |  | Na+, K+, pH |  |  |  |  | 1 | 3 | Prototype/Validation | Human | Human studies | NA | NA | Flexible substrates in the sensor's construction, allowing it to conform to the contours of the body for comfortable and unobtrusive wearing. This design is essential for a device intended for continuous or real-time monitoring of sweat composition during daily activities or exercise.The device is equipped with wireless communication capabilities, | NO | NA | YES | NA | 2024 | Jalal, N. R.; Madrakian, T.; Ahmadi, M.; Afkhami, A.; Khalili, S.; Bahrami, M.; Roshanaei, M. | Wireless wearable potentiometric sensor for simultaneous determination of pH, sodium and potassium in human sweat | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | This paper reports on the development of a flexible-wearable potentiometric sensor for real-time monitoring of sodium ion (Na(+)), potassium ion (K(+)), and pH in human sweat. Na(0.44)MnO(2), polyaniline, and K(2)Co[Fe(CN)(6)] were used as sensing materials for Na(+), H(+) and K(+) monitoring, respectively. The simultaneous potentiometric Na(+), K(+), and pH sensing were carried out by the developed sensor, which enables signal collection and transmission in real-time to the smartphone via a Wi-Fi access point. Then, the potentiometric responses were evaluated by a designed android application. Na(+), K(+), and pH sensors illustrated high sensitivity (59.7 +/- 0.8 mV/decade for Na(+), 57.8 +/- 0.9 mV/decade for K(+), and 54.7 +/- 0.6 mV/pH for pH), excellent stability, and good batch-to-batch reproducibility. The results of on-body experiments demonstrated that the proposed platform is capable of real-time monitoring of the investigated ions. | 45433 | 2024
FTD47T4B | 10.1016/j.bios.2024.116662 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324006687 | YES | skin-attached |  | NO | monitoring of specific biomarkers in sweat | Raman | Surface-Enhanced Raman Spectroscopy (SERS) | The SERS biosensor can detect creatinine and UA at concentrations as low as 10-8 M in both artificial and real sweat samples. The sensor exhibits a strong linear correlation between SERS intensity and the logarithm of analyte concentration over a specific range. This range is not explicitly stated in the sources, but linear regression curves are provided for both creatinine and UA. The device can simultaneously detect both creatinine and UA in artificial sweat at concentrations as low as 10-7 M. |  | NA | yes | - | - | sweat | [7 empty] | yes |  |  |  | yes | creatinine, uric acid (UA) |  |  |  |  | 1 | 2 | Validation | Human | in-vitro tests + laboratory testing + human studies | NA | NA | The wearable SERS biosensor for sweat creatinine and uric acid monitoring integrates key innovations: Bimetallic Core-Shell SERS Substrate: A gold-silver (Au@Ag) structure enhances sensitivity, stability, and signal reproducibility, with a silk fibroin-based anti-opal array amplifying electromagnetic fields and ensuring mechanical durability. Integrated Microfluidics: Seamless channels guide efficient sweat sampling, enabling dynamic, real-time biomarker monitoring with optimized design via COMSOL simulations. Silk Fibroin Substrate: A biocompatible, flexible, and adhesive natural material ensures wearability and eco-friendliness. Label-Free Detection: Simplifies analysis by eliminating the need for additional reagents, enhancing convenience and reducing costs. | NO | NA | NA | NA | 2024 | Hu, Mengsu; Zhu, Kai; Wei, Jinxiu; Yang, Kuo; Wu, Lei; Zong, Shenfei; Wang, Zhuyuan | Silk fibroin-based wearable SERS biosensor for simultaneous sweat monitoring of creatinine and uric acid | Biosensors and Bioelectronics |  | 9565663 | Sweat biomarkers have the potential to offer valuable clinical insights into an individual’s health and disease condition. Current sensors predominantly utilize enzymes and antibodies as biometric components to measure biomarkers present in sweat quantitatively. However, enzymes and antibodies are susceptible to interference by environmental factors, which may affect the performance of the sensor. Herein, we present a wearable microfluidic surface-enhanced Raman scattering (SERS) biosensor that enables the non-invasive and label-free detection of biomarkers in sweat. Concretely, we developed a bimetallic self-assembled anti-opal array structure with uniform hot spots, enhanced the Raman scattering effect, and integrated it into a silk fibroin-based sensing patch. Utilizing a silk fibroin substrate in the wearable SERS sensor imparts desirable properties such as softness, breathability, and biocompatibility, which enables the sensor to establish close contact with the skin without causing chemical or physical irritation. In addition, introducing microfluidic channels enables the controlled and high temporal resolution management of sweat, facilitating more efficient sweat collection. The proposed label-free SERS sensor can offer chemical ‘fingerprint’ information, enabling the identification of sweat analytes. As an illustration of the feasibility, we have effectively monitored the creatinine and uric acid levels in sweat. This study presents a versatile and highly sensitive approach for the simultaneous detection of biomarkers in human sweat, showcasing significant potential for application in point-of-care monitoring. | 45627 | 2024
GVRTPDR7 | 10.1007/s00421-023-05364-4 | https://link.springer.com/10.1007/s00421-023-05364-4 | YES | skin-attached |  | NO | monitor sodium concentration in sweat during exercise | electrochemical | Electrochemical | The Wearable Fabric Sensor (WFS) demonstrates strong accuracy in measuring sweat sodium concentration compared to the gold standard Sweat Sodium Analyzer (SSA). The WFS readings (54 mM ± 22 mM) closely matched the SSA (59 mM ± 22 mM), with no significant difference (p = -0.45) and a mean bias of 5.35 mM. Statistical analysis confirmed high agreement, with Lin’s concordance correlation coefficient at 0.95 and a Pearson r-squared value of 0.9. Error metrics, including a MAPE of 12% and MAE of 7.2 mM, indicate robust performance. Although sensitivity is not explicitly reported, the WFS effectively detects sodium variations during exercise. |  | NA | yes | - | - | sweat | [12 empty] | Na+ |  |  |  |  | 1 | 1 | validation | Human | Human studies | NA | NA | The Wearable Fabric Sensor (WFS) features a novel material design using electrospun nylon-6/MWCNT nanocomposites for sodium sensing. This structure provides high surface area, selective sodium ion binding, and resistance-based sensing, enabling accurate quantification of sweat sodium concentration. Integrated into a flexible fabric platform, the sensor ensures comfort, conformal skin contact, and moisture-wicking for continuous monitoring. Custom-built electronics with a microcontroller and ADC process and store resistance data, while a thermistor compensates for temperature variations. The WFS directly measures sodium, offering advantages over chloride-based sensors, with simpler, cost-effective fabrication and future potential for wireless data transmission. | RooSense LLC | USA | YES | NA | 2024 | Dyshko, Kristina; Nicodemus, Mary Pat; Otterstetter, Ronald; Ghadimi, Hanieh; Daniels, Shelby; Fulmer, Michael S.; Cheney, Zachary; Ellis, Rebecca; Stege, Victoria; Monty, Chelsea N. | Evaluation of a wearable fabric-based sensor for accurate sodium determination in sweat during exercise | European Journal of Applied Physiology |  | 1439-6319, 1439-6327 | Introduction Newly developed wearable fabric sensors (WFS) can increase the ease and accuracy of sweat sodium measurements by performing simultaneous sampling and analysis on the body during exercise. Purpose Determine the accuracy of a WFS for measurement of sodium concentration in sweat. METHODS: Subjects wore a WFS prototype and sweat collectors on their forearm during cycle ergometry. Subjects exercised at a moderate intensity (~ 65% heart rate reserve) for 30–60 min. Sweat samples were collected and analyzed using a commercial sweat sodium analyzer (SSA) every 10–15 min. WFS were adhered with an armband and connected to custom built electronics. Accuracy was determined by comparing predicted WFS concentration to the actual concentration from the commercial SSA and analyzed statistically using ANOVA and Bland−Altman plots. Results A total of 19 subjects completed the study. The average sweat sodium concentration was 59 mM ± 22 mM from a SSA compared with 54 mM ± 22 mM from the WFS. Overall, the average accuracy of the WFS was 88% in comparison to the SSA with p = 0.45. A line of best fit comparing predicted versus actual sweat sodium concentration had a slope of 0.99, intercept of – 4.46, and an r2 of 0.90. Bland−Altman analysis showed the average concentration difference between the WFS and the SSA was 5.35 mM, with 99% of data points between ± 1.96 times the standard deviation. Conclusion The WFS accurately predicted sweat sodium concentration during moderate intensity cycle ergometry. With the need for precise assessment of sodium loss, especially during long duration exercise, this novel analysis method can benefit athletes and coaches. Further research involving longer duration and more intense exercise is warranted. | 45413 | 2024
PTQLBWUT | 10.1016/j.bios.2024.116136 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324001416 | YES | patch | skin-attached | NO | non-invasive monitoring of riboflavin levels in sweat. | electrochemical | Electrochemical | Sensitivity: the rGO/PtNPs modified electrode achieved an average sensitivity of 0.0067 μA/nM for riboflavin detection, with a relative standard deviation (RSD) of 6.18%. This high sensitivity is attributed to the synergistic effect of the rGO and PtNPs, which enhance the electroactive surface area and charge transfer efficiency. Detection Limit: the device boasts a low detection limit of 1.2 nM for riboflavin, calculated using the 3δ/slope method (where δ represents sensor noise). This ultra-low detection limit enables the system to accurately measure riboflavin within the physiologically relevant range found in sweat. Linearity: the sensor demonstrated excellent linearity (R2 = 0.9935) within the riboflavin concentration range of 10 nM to 1000 nM, covering the typical physiological concentrations found in sweat. Accuracy: on-body validation studies comparing the wearable sensor's measurements with gold-standard UHPLC analysis showed strong agreement. A high correlation (R2 = 0.9783) and a slope rate of nearly 1 (1.049) confirm the high accuracy of the wearable biosensing system for riboflavin detection in sweat. | Metabolic Syndrome | nutritional health management (focusing on riboflavin status). | yes | - | - | sweat | [10 empty] | yes |  | riboflavin, pH |  |  |  |  | 1 | 2 | Prototype/Validation | Human | in-vitro tests + human studies | NA | NA | Flexible, Battery-Free Design: a flexible printed circuit board (PCB) enables comfortable skin attachment, powered wirelessly via NFC technology for lightweight, unobtrusive wear. Nanomaterial-Enhanced Sensitivity: electrodes with reduced graphene oxide (rGO) and platinum nanoparticles (PtNPs) provide high sensitivity and selectivity for riboflavin detection. Efficient Sweat Sampling: a microfluidic patch with hydrophilic microchannels ensures continuous sweat flow and real-time monitoring with optimized temporal resolution. Dual-Channel Sensing: a multi-analyte sensor array simultaneously measures riboflavin and pH, improving accuracy and expanding applications for skin health insights. | NO | NA | YES | NA | 2024 | Shi, Zhenghan; Deng, Peixue; Zhou, Li-ang; Jin, Meng; Fang, Feiyue; Chen, Tao; Liu, Guang; Wen, Hao; An, Zijian; Liang, Hao; Lu, Yanli; Liu, Jun; Liu, Qingjun | Wireless and battery-free wearable biosensing of riboflavin in sweat for precision nutrition | Biosensors and Bioelectronics |  | 9565663 | Nutrition assessment is crucial for dietary guidance and prevention of malnutrition. Recent endeavors in wearable biochemical sensors have enabled real-time, in situ analysis of nutrients in sweat. However, the monitoring of riboflavin, an indispensable vitamin B involved in energy metabolism, remains challenging due to its trace level and variations in the sweat matrix. Herein, we report a wireless, battery-free, and flexible wearable biosensing system for the in situ monitoring of sweat riboflavin. Highly sensitive and selective electrochemical voltammetric detection is realized based on the synergistic effect of electrodeposited reduced graphene oxide (rGO) and platinum nanoparticles (PtNPs) with a low detection limit of 1.2 nM. The fully integrated system is capable of sweat sampling with the microfluidic patch, real-time riboflavin analysis and pH calibration with the flexible electrode array, as well as wirelessly simultaneous near field communication (NFC) energy harvesting and data transmission with the flexible circuit and a smartphone. On-body human sweat analysis demonstrates high accuracy cross-validated with gold-standard measurements, and reveals a strong correlation between sweat and urine riboflavin levels. The proposed wearable platform opens up attractive possibilities for noninvasive nutrient tracking, providing strong potential for personalized dietary guidance towards precision nutrition. | 45413 | 2024
5X9VZ3MQ | 10.1016/j.talanta.2020.121815 | https://www.ncbi.nlm.nih.gov/pubmed/33379040 | YES | patch |  | NO | alcohol monitoring through sweat analysis | electrochemical | Electrochemical. | The current response increased linearly (R 2 = 0.993) over alcohol concentration range making the presented methodology suitable for the detection of small ingestions of alcohol. The limit of detection (LOD) was 0.045 mM/L which is 10 times lower than the amount of alcohol after the ingestion of 1 drink. |  | NA | yes | - | - | sweat | [12 empty] | Ethanol |  |  |  |  | 1 | 1 | proof-of-concept |  | In-vitro tests | NA | NA | electrochemical sensor within a continuous monitoring application. Utilising disposable screen-printed electrodes modified with a PANI containing AOX film, a low-cost alternative to the complex instrumentation currently utilised for analysis of such samples is offered. | NO | NA | NA | NA | 2021 | Biscay, J.; Findlay, E.; Dennany, L. | Electrochemical monitoring of alcohol in sweat | Talanta |  | 1873-3573 (Electronic) 0039-9140 (Linking) | Rapid, periodic monitoring and detection of ethanol (EtOH) after consumption via a non-invasive measurement has been an area of increased research in recent years. Current point-of-care or on-site detection strategies rely on single use sensors which are inadequate for monitoring during a longer period. A low cost, portable and novel approach is developed here for real-time monitoring over several days utilising electrochemical techniques. The sensor shows oxidation of the ethanol in phosphate buffer and artificial sweat using the amperometric response from the application of +0.9 V to the polyaniline modified screen printed electrode using 1 mM EtOH as the averaged amount of EtOH eliminated in sweat after the consumption of one alcoholic beverage. Our enzyme based electrochemical sensor exhibits a qualitative assessment of the presence of EtOH in small volumes (</=40 muL) of 0.1 M sodium bicarbonate and subsequently artificial sweat, with 50 measurements taken daily over 11 days. While quantitative information is not obtained, the sensor system exhibits excellent stability after 3 months' dried storage in this complex biological matrix in an oxygen free cabinet. This addresses one of the key challenges for enzyme based electrochemical sensors, namely, the ability for real-time monitoring in complex biological matrices. The qualitative response illustrates the potential for this sensor to be exploited by non-experts which suggests the promise for their wider application in next-generation wearable electronics necessary for alcohol monitoring. | 44256 | 2021
8M49IZSY | 10.1038/s41467-021-22109-z | https://www.ncbi.nlm.nih.gov/pubmed/33758197 | YES | patch |  | NO | continuous sweat monitoring at rest | electrochemical | Electrochemical | We observed that levodopa concentration in sweat generally increases with increasing doses. When other foods with minimal levodopa is consumed, no significant signal is observed.This indicates that monitoring sweat levodopa may be a promising way to keep track of blood levodopa to assist medication management of Parkinson’s disease patients. However, the exact relations between sweat levodopa concentration, plasma levels, and intake dose can depend on diet, hydration, other physiological conditions that impact absorption and metabolism rates, and on sweat rate and secretion mechanisms. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | pH, Cl-, levodopa |  |  |  |  | 1 | 3 | Prototype/Validation | Human | Human studies | NA | NA | the wearable sweat patch incorporates several cutting-edge technologies, including microfluidics for controlled sweat management, hydrophilic fillers for rapid sample uptake, integrated electrochemical sensors for biomarker detection, and a flexible substrate for comfortable wearability. | NO | NA | NA | NA | 2021 | Nyein, H. Y. Y.; Bariya, M.; Tran, B.; Ahn, C. H.; Brown, B. J.; Ji, W.; Davis, N.; Javey, A. | A wearable patch for continuous analysis of thermoregulatory sweat at rest | Nat Commun |  | 2041-1723 (Electronic) 2041-1723 (Linking) | The body naturally and continuously secretes sweat for thermoregulation during sedentary and routine activities at rates that can reflect underlying health conditions, including nerve damage, autonomic and metabolic disorders, and chronic stress. However, low secretion rates and evaporation pose challenges for collecting resting thermoregulatory sweat for non-invasive analysis of body physiology. Here we present wearable patches for continuous sweat monitoring at rest, using microfluidics to combat evaporation and enable selective monitoring of secretion rate. We integrate hydrophilic fillers for rapid sweat uptake into the sensing channel, reducing required sweat accumulation time towards real-time measurement. Along with sweat rate sensors, we integrate electrochemical sensors for pH, Cl(-), and levodopa monitoring. We demonstrate patch functionality for dynamic sweat analysis related to routine activities, stress events, hypoglycemia-induced sweating, and Parkinson's disease. By enabling sweat analysis compatible with sedentary, routine, and daily activities, these patches enable continuous, autonomous monitoring of body physiology at rest. | 44278 | 2021
LTY9X8HQ | 10.1016/j.bios.2024.116358 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324003634 | YES | patch |  | NO | continuous, non-invasive glucose monitoring using human sweat | electrochemical | Electrochemical | The GOx/PB/MWCNT-Fe3O4 printed sensor exhibits a sensitivity of 4.495 μAμM−1cm−2 in a phosphate-buffered saline (PBS) solution. Further testing reveals that the sensitivity slightly improves in an artificial sweat solution, reaching 4.59 μAμM−1cm−2. The sensor also demonstrates a remarkable limit of detection (LOD) of 0.38 μM in PBS and 0.31 μM in artificial sweat. These sensitivity and LOD values, combined with a linear detection range of 1 μM to 400 μM, highlight the sensor's accuracy and capability to effectively monitor glucose levels within the physiological range relevant for diabetes management. | Diabetes | diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation | Human | human studies | NA | NA | Nanocomposite Ink: a custom ink with MWCNTs, Fe₃O₄ nanoparticles, and PEDOT:PSS enhances conductivity, catalytic activity, and biocompatibility. All-Printed Flexible Design: screen-printing on a Mylar® substrate enables low-cost, scalable production, with excellent wearability and comfort. Prussian Blue Nanoparticles: serve as efficient redox mediators, boosting sensitivity and stability during glucose detection. Partial Enzyme Encapsulation: MWCNTs and glutaraldehyde ensure long-term enzyme stability, protecting glucose oxidase from degradation. | NO | NA | NA | NA | 2024 | Shamili, Chandradas; Pillai, Adarsh Sivan; Saisree, S; Chandran, Achu; Varma, Manoj Raama; Kuzhichalil Peethambharan, Surendran | All-printed wearable biosensor based on MWCNT-iron oxide nanocomposite ink for physiological level detection of glucose in human sweat | Biosensors and Bioelectronics |  | 9565663 | Wearable sensors for sweat glucose monitoring are gaining massive interest as a patient-friendly and noninvasive way to manage diabetes. The present work offers an alternative on-body method employing an allprinted flexible electrochemical sensor to quantify the amount of glucose in human sweat. The working electrode of the glucose sensor was printed using a custom-formulated ink containing multi-walled carbon nanotube (MWCNT), poly (3,4-ethylenedioxythiophene) polystyrene sulfonate (PEDOPT: PSS), and iron (II, III) oxide (Fe3O4) nanoparticles. This novel ink composition has good conductivity, enhanced catalytic activity, and excellent selectivity. The working electrode was modified using Prussian blue (PB) nanoparticles and glucose oxidase enzyme (GOx). The sensor displayed a linear chronoamperometric response to glucose from 1 μM to 400 μM, with a precise detection limit of ~0.38 μM and an impressive sensitivity of ~4.495 μAμM 1cm 2. The sensor stored at 4 ◦C exhibited excellent stability over 60 days, high selectivity, and greater reproducibility. The glucose detection via the standard addition method in human sweat samples acquired a high recovery rate of 96.0–98.6%. Examining human sweat during physical activity also attested to the biosensor’s real-time viability. The results also show an impressive correlation between glucose levels obtained from a commercial blood glucose meter and sweat glucose concentrations. Remarkably, the present results outperform previously published printed glucose sensors in terms of detection range, low cost, ease of manufacturing, stability, selectivity, and wearability. | 45505 | 2024
PQAUHNH2 | 10.1038/s41551-022-00916-z | https://www.ncbi.nlm.nih.gov/pubmed/35970928 | YES | patch |  | NO | monitoring of multiple metabolites and nutrients in sweat. | electrochemical | Electrochemical | The sensitivity for the in vitro indirect detection of target molecules using the LEG-AQCA-MIP sensors was reported as 1×10−5AV−1 for negative differential pulse voltammetry (DPV) analysis. |  | NA | yes | - | - | sweat |  |  |  |  | yes |  |  | yes |  | yes | yes | yes | All nine essential AAs (Leu, Ile, Val, Trp, Phe, His, Lys, Met, Thr) + vitamin B6, C, D3, E, glucose, uric acid, creatine, creatinine, and cholesterol. |  |  |  |  | 1 | 18 | Prototype/Validation | Human | Human studies | NA | NA | A wearable electrochemical biosensor for the continuous analysis, in sweat during physical exercise and at rest, of trace levels of multiple metabolites and nutrients, including all essential amino acids and vitamins. The biosensor consists of graphene electrodes that can be repeatedly regenerated in situ, functionalized with metabolite-specific antibody-like molecularly imprinted polymers and redox-active reporter nanoparticles, and integrated with modules for iontophoresis-based sweat induction, microfluidic sweat sampling, signal processing and calibration, and wireless communication. | NO | NA | YES | NA | 2022 | Wang, M.; Yang, Y.; Min, J.; Song, Y.; Tu, J.; Mukasa, D.; Ye, C.; Xu, C.; Heflin, N.; McCune, J. S.; Hsiai, T. K.; Li, Z.; Gao, W. | A wearable electrochemical biosensor for the monitoring of metabolites and nutrients | Nat Biomed Eng |  | 2157-846X (Electronic) 2157-846X (Linking) | Wearable non-invasive biosensors for the continuous monitoring of metabolites in sweat can detect a few analytes at sufficiently high concentrations, typically during vigorous exercise so as to generate sufficient quantity of the biofluid. Here we report the design and performance of a wearable electrochemical biosensor for the continuous analysis, in sweat during physical exercise and at rest, of trace levels of multiple metabolites and nutrients, including all essential amino acids and vitamins. The biosensor consists of graphene electrodes that can be repeatedly regenerated in situ, functionalized with metabolite-specific antibody-like molecularly imprinted polymers and redox-active reporter nanoparticles, and integrated with modules for iontophoresis-based sweat induction, microfluidic sweat sampling, signal processing and calibration, and wireless communication. In volunteers, the biosensor enabled the real-time monitoring of the intake of amino acids and their levels during physical exercise, as well as the assessment of the risk of metabolic syndrome (by correlating amino acid levels in serum and sweat). The monitoring of metabolites for the early identification of abnormal health conditions could facilitate applications in precision nutrition. | 44866 | 2022
BHGBZAVU | 10.1002/smtd.202200653 | https://onlinelibrary.wiley.com/doi/10.1002/smtd.202200653 | YES | skin-attached |  | NO | continuously and wirelessly monitor biomarkers in sweat at the molecular level | electrochemical | Electrochemical | The Na+ sensor exhibits a near-Nernstian sensitivity of 67.43 mV per decade of ion concentration. This means that for every tenfold increase in the concentration of Na+ ions in the sweat, the sensor's output voltage increases by approximately 67.43 mV. The K+ sensor demonstrates a near-Nernstian sensitivity of 30.42 mV per decade of ion concentration. Similarly, for a tenfold increase in K+ ion concentration, the K+ sensor's output voltage increases by about 30.42 mV. |  | NA | yes | - | - | sweat | [12 empty] | Na+, K+ |  |  |  |  | 1 | 2 | Validation | Human | in-vitro tests + laboratory testing + human studies | NA | NA | The Sweat Wearable Sensor Analysis System (SWSAS) integrates Hybrid Nanogenerator Modules (HNGMs), which combine Triboelectric Nanogenerators (TENGs) and Electromagnetic Generators (EMGs) to harvest mechanical energy from motion. This hybrid system efficiently powers the wearable device. A flexible sweat sensor patch, made from laser-induced graphene (LIG), enables comfortable ion detection (e.g., Na+, K+) using ion-selective electrodes (ISEs) and microfluidic channels for precise sweat sampling. The low-power circuit board includes energy management, signal amplification, and a Bluetooth Low Energy (BLE) module for wireless data transmission, making SWSAS a reliable, real-time sweat monitoring system. | NO | NA | YES | NA | 2022 | Gai, Yansong; Wang, Engui; Liu, Minghao; Xie, Lirong; Bai, Yuan; Yang, Yuan; Xue, Jiangtao; Qu, Xuecheng; Xi, Yuan; Li, Linlin; Luo, Dan; Li, Zhou | A Self‐Powered Wearable Sensor for Continuous Wireless Sweat Monitoring | Small Methods |  | 2366-9608, 2366-9608 | Abstract Wireless wearable sweat analysis devices can monitor biomarkers at the molecular level continuously and in situ, which is highly desired for personalized health care. The miniaturization, integration, and wireless operation of sweat sensors improve the comfort and convenience while also bringing forward new challenges for power supply technology. Herein, a wireless self‐powered wearable sweat analysis system (SWSAS) is designed that effectively converts the mechanical energy of human motion into electricity through hybrid nanogenerator modules (HNGMs). The HNGM shows stable output characteristics at low frequency with a current of 15 mA and a voltage of 60 V. Through real‐time on‐body sweat analysis powered by HNGM, the SWSAS is demonstrated to selectively monitor biomarkers (Na + and K + ) in sweat and wirelessly transmit the sensing data to the user interface via Bluetooth. | 44835 | 2022
LIKNRV8Y | 10.3390/bios12100854 | https://www.ncbi.nlm.nih.gov/pubmed/36290991 | YES | patch | textile | NO | cortisol detection in sweat | electrochemical | Electrochemical | The cortisol sensor rapidly (<2 min) responded to 9.8−49.5 ng/mL of cortisol, with an average relative standard deviation (%RSD) of 6.4% across the dynamic range, indicating excellent precision. The cortisol sensor yielded an excellent limit of detection (LOD) of 8.00 ng/mL, which is within the typical physiological levels in human sweat. A single cortisol sensor patch could be reused 15 times over a 30-day period with no loss in performance, attesting to excellent reusability. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | Prototype | Human | Human studies | NA | NA | The wearable cortisol sensor was fabricated via layer-by-layer assembly (LbL) on a flexible cotton textile substrate coated with a conductive nanoporous carbon nanotube/cellulose nanocrystal (CNT/CNC) composite suspension, conductive polyaniline (PANI), and a selective cortisol-imprinted poly(glycidylmethacrylate-co-ethylene glycol dimethacrylate) (poly(GMA-co-EGDMA)) decorated with gold nanoparticles (AuNPs), or plated with gold. | NO | NA | NA | NA | 2022 | Mugo, S. M.; Lu, W.; Robertson, S. | A Wearable, Textile-Based Polyacrylate Imprinted Electrochemical Sensor for Cortisol Detection in Sweat | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | A wearable, textile-based molecularly imprinted polymer (MIP) electrochemical sensor for cortisol detection in human sweat has been demonstrated. The wearable cortisol sensor was fabricated via layer-by-layer assembly (LbL) on a flexible cotton textile substrate coated with a conductive nanoporous carbon nanotube/cellulose nanocrystal (CNT/CNC) composite suspension, conductive polyaniline (PANI), and a selective cortisol-imprinted poly(glycidylmethacrylate-co-ethylene glycol dimethacrylate) (poly(GMA-co-EGDMA)) decorated with gold nanoparticles (AuNPs), or plated with gold. The cortisol sensor rapidly (<2 min) responded to 9.8-49.5 ng/mL of cortisol, with an average relative standard deviation (%RSD) of 6.4% across the dynamic range, indicating excellent precision. The cortisol sensor yielded an excellent limit of detection (LOD) of 8.00 ng/mL, which is within the typical physiological levels in human sweat. A single cortisol sensor patch could be reused 15 times over a 30-day period with no loss in performance, attesting to excellent reusability. The cortisol sensor patch was successfully verified for use in quantification of cortisol levels in human sweat. | 44844 | 2022
GNHCYBKN | 10.1016/j.bios.2022.114994 | https://linkinghub.elsevier.com/retrieve/pii/S095656632201034X | YES | skin-attached |  | NO | monitoring urea levels in sweat | electrochemical | Electrochemical | The biosensor's sensitivity varies with the pH of the sweat sample. At lower pH values (more acidic conditions), the sensitivity is higher but with a narrower linear range. The sources provide specific sensitivity values (slopes of the calibration curves) for different pH levels: pH 4.9: -0.17 V/decade of urea concentration. pH 5.7: -0.062 V/decade of urea concentration; pH 7.0: -0.053 V/decade of urea concentration; pH 7.4: -0.045 V/decade of urea concentration. Incorporating a PVC membrane on the sensor surface improves sensitivity by extending the linear range and delaying the saturation point at higher urea concentrations. The biosensor demonstrated good accuracy in measuring urea levels in artificial sweat samples, achieving recovery values of 105% and 98%. Recovery refers to the percentage of the known amount of analyte that is measured by the sensor. The device was also tested on real sweat samples collected from a human volunteer. The average urea concentration obtained (10.9 mM) was within the expected physiological range for healthy individuals, suggesting accurate measurements. The integrated pH sensor plays a critical role in ensuring accurate urea measurements by correcting for variations in sweat pH. The sources note that without pH correction, significant errors in urea quantification could occur, particularly in more acidic samples. The reproducibility of sensor fabrication can also affect accuracy. The sources report a relative standard deviation (RSD) of 13.9% when comparing the sensitivity of five printed biosensors, indicating some variability in sensor performance. |  | NA | yes | - | - | sweat | [8 empty] | yes |  | yes |  | urea, pH |  |  |  |  | 1 | 2 | Validation | Human | in-vitro + human studies | NA | NA | Integrated pH Sensor for Accurate Urea Monitoring: Incorporates real-time pH correction, enabling precise urea quantification despite individual sweat pH variations, while simultaneously monitoring pH as an additional biomarker for hydration status. Flexible Screen-Printed Electrodes (SPCEs): Fabricated on flexible substrates, ensuring skin conformability, mechanical durability, and scalability for mass production. Exploration of Roll-to-Roll (R2R) Manufacturing: High-throughput techniques, such as screen-printing and slot-die coating, reduce costs, enable large-scale production, and preserve biomolecule integrity. Specialized Sensor Inks: Polyaniline (PANI) nanoparticles, urease bioink, and PVC-based membrane inks enhance conductivity, enzyme activity, and dynamic range, supporting robust sensor performance and manufacturability. | NO | NA | NA | NA | 2023 | Ibáñez-Redín, Gisela; Rosso Cagnani, Giovana; O. Gomes, Nathalia; Raymundo-Pereira, Paulo A.; S. Machado, Sergio A.; Gutierrez, Marco Antonio; Krieger, Jose Eduardo; Oliveira, Osvaldo N. | Wearable potentiometric biosensor for analysis of urea in sweat | Biosensors and Bioelectronics |  | 9565663 | Herein, we introduce wearable potentiometric biosensors on screen-printed carbon electrodes (SPCEs) for onbody and on-site monitoring of urea in sweat. The biosensor architecture was judiciously designed to detect urea at different pHs and incorporate a pH sensor, thus containing polyaniline ink, urease bioink and a polyvinylchloride membrane. Urea detection could be performed in the wide range from 5 to 200 mM at pH 7.0, encompassing urea levels in human sweat. The biosensor response was fast (incubation time 5 min), with no interference from other substances in sweat. Reliable urea detection could be done in undiluted human sweat with a skin-worn flexible device using the pH correction strategy afforded by the pH sensor. The performance of the epidermal biosensor was not affected by severe bending strains. The feasibility of mass production was demonstrated by fabricating epidermal flexible biosensors using slot-die coating with a roll-to-roll technique. | 44986 | 2023
9ATFLSW3 | 10.1016/j.bios.2024.116560 | https://www.ncbi.nlm.nih.gov/pubmed/39018979 | YES | patch |  | NO | real-time, continuous monitoring of specific biomarkers in sweat. | electrochemical | Electrochemical | ISFET sensor devices can be readily applied for measuring ion concentrations in sweat. pH, Na +, and K + ISFETs were characterized in standard and artificial sweat solutions, demonstrating an overall good performance in the range of analyte concentrations present in sweat. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | pH, Na+, K+ |  |  |  |  | 1 | 3 | Validation | Human | Human studies | NA | NA | The platform comprises Ion-Sensitive Field-Effect Transistors (ISFETs) manufactured on silicon, a tailor-made solid-state reference electrode, and a temperature sensor integrated into a patch-like polymeric substrate, together with the component that easily collects and drives samples under continuous capillary flow to the sensor areas. | NO | NA | NA | NA | 2024 | Rovira, M.; Lafaye, C.; Demuru, S.; Kunnel, B. P.; Aymerich, J.; Cuenca, J.; Serra-Graells, F.; Margarit-Taule, J. M.; Haque, R.; Saubade, M.; Fernandez-Sanchez, C.; Jimenez-Jorquera, C. | Assessing the performance of a robust multiparametric wearable patch integrating silicon-based sensors for real-time continuous monitoring of sweat biomarkers | Biosens Bioelectron |  | 1873-4235 (Electronic) 0956-5663 (Linking) | The development of wearable devices for sweat analysis has experienced significant growth in the last two decades, being the main focus the monitoring of athletes health during workouts. One of the main challenges of these approaches has been to attain the continuous monitoring of sweat for time periods over 1 h. This is the main challenge addressed in this work by designing an analytical platform that combines the high performance of potentiometric sensors and a fluidic structure made of a plastic fabric into a multiplexed wearable device. The platform comprises Ion-Sensitive Field-Effect Transistors (ISFETs) manufactured on silicon, a tailor-made solid-state reference electrode, and a temperature sensor integrated into a patch-like polymeric substrate, together with the component that easily collects and drives samples under continuous capillary flow to the sensor areas. ISFET sensors for measuring pH, sodium, and potassium ions were fully characterized in artificial sweat solutions, providing reproducible and stable responses. Then, the real-time and continuous monitoring of the biomarkers in sweat with the wearable platform was assessed by comparing the ISFETs responses recorded during an 85-min continuous exercise session with the concentration values measured using commercial Ion-Selective Electrodes (ISEs) in samples collected at certain times during the session. The developed sensing platform enables the continuous monitoring of biomarkers and facilitates the study of the effects of various real working conditions, such as cycling power and skin temperature, on the target biomarker concentration levels. | 45580 | 2024
W72UAR3C | 10.1002/adhm.201901321 | https://onlinelibrary.wiley.com/doi/10.1002/adhm.201901321 | YES | skin-attached |  | NO | multiplexed sensing of ammonium (NH4+) and calcium (Ca2+) ions from human perspiration | electrochemical | Electrochemical | Ion concentrations detected: 0.01 × 10⁻³ to 100 × 10⁻³ M for both NH4⁺ and Ca²⁺. Covers typical physiological levels in human sweat. Achieves steady-state current in <1 second, enabling real-time monitoring. Relative Standard Deviation (RSD): 5% across three measurements per subject. Stability test over 100 cycles: Signal loss of only 1.1–2.9% in artificial sweat solution. Minimal impact from interfering ions (e.g., K⁺, Na⁺). Maintains accurate NH4⁺ and Ca²⁺ detection above background noise. Strong correlation with standard ion-selective electrodes. |  | NA | yes | - | - | sweat | [12 empty] | NH4+, Ca2+ |  |  |  |  | 1 | 2 | Validation | Human | In-vitro, Human-studies | NA | NA | Substrate: Styrene-Ethylene-Butylene-Styrene (SEBS) elastomer, highly stretchable and durable. Sweat Acquisition Layer: Laser-patterned microcapillary channels on polyimide tape for efficient sample collection. Organic Electrochemical Transistors (OECTs): Transconductance: ≈2 mS; Operating Voltage: <0.5 V. Ion-Selective Membranes (ISMs): Selective for NH4⁺ and Ca²⁺ via ionophore shuttles. Two PEDOT:PSS OECT channels functionalized for NH4⁺ and Ca²⁺ detection. Scalable design for additional ions with tailored ISMs. Fully integrated system with real-time sweat analysis. Lightweight, wearable design compatible with portable electronics. | NO | NA | NA | NA | 2019 | Keene, Scott T.; Fogarty, Daragh; Cooke, Ross; Casadevall, Carlos D.; Salleo, Alberto; Parlak, Onur | Wearable Organic Electrochemical Transistor Patch for Multiplexed Sensing of Calcium and Ammonium Ions from Human Perspiration | Advanced Healthcare Materials |  | 2192-2640, 2192-2659 | Abstract Wearable health monitoring has garnered considerable interest from the healthcare industry as an evolutionary alternative to standard practices with the ability to provide rapid, off‐site diagnosis and patient‐monitoring. In particular, sweat‐based wearable biosensors offer a noninvasive route to continuously monitor a variety of biomarkers for a range of physiological conditions. Both the accessibility and wealth of information of sweat make it an ideal target for noninvasive devices that can aid in early diagnosis of disease or to monitor athletic performance. Here, the integration of ammonium (NH 4 + ) and calcium (Ca 2+ ) ion‐selective membranes with a poly(3,4‐ethylenedioxythiophene):poly(styrenesulfonate)‐based (PEDOT:PSS) organic electrochemical transistor (OECT) for multiplexed sensing of NH 4 + and Ca 2+ in sweat with high sensitivity and selectivity is reported for the first time. The presented wearable sweat sensor is designed by combining a flexible and stretchable styrene‐ethylene‐butene‐styrene substrate with a laser‐patterned microcapillary channel array for direct sweat acquisition and delivery to the ion‐selective OECT. The resulting dermal sensor exhibits a wide working range between 0.01 × 10 −3 and 100 × 10 −3 m , well within the physiological levels of NH 4 + and Ca 2+ in sweat. The integrated devices are successfully implemented with both ex situ measurements and on human subjects with real‐time analysis using a wearable sensor assembly. | 43800 | 2019
W57E676X | 10.1016/j.bios.2022.114450 | https://linkinghub.elsevier.com/retrieve/pii/S0956566322004900 | YES | skin-attached |  | NO | real-time, non-invasive glucose monitoring in sweat. | electrochemical | Electrochemical | Limit of Detection (LOD): The LOD is the lowest concentration of an analyte that can be reliably detected by the sensor. The glucose tag has an LOD of 24 μM. Limit of Quantification (LOQ): The LOQ is the lowest concentration of an analyte that can be quantitatively measured with acceptable accuracy and precision. The glucose tag boasts an LOQ of 74 μM. Sensitivity: The sensitivity of the sensor refers to the change in output signal per unit change in analyte concentration. The glucose tag has a sensitivity of 1.27 μA cm−2 mM−1 | Diabetes | diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Validation | Human | in-vitro tests + Ex-situ (Off-Body) Testing + Human Studies (preliminary human trials) + laboratory Testing | NA | NA | Femtosecond Laser Ablation: Enables precise, damage-free fabrication of a flexible NFC antenna on a PCB, eliminating cleanroom needs while achieving high performance in a miniaturized design. Battery-Free NFC Technology: Utilizes NFC for wireless powering and data transfer, harvesting energy from smartphones, which reduces size and weight. Flexible, Layered Structure: Ensures comfort and reliable function during movement, integrating antenna, sensor, and microfluidic channels into a thin, flexible assembly. Ultra-Compact Design: With a footprint of 1.2 cm² and weighing just 0.13 g, it is among the smallest and lightest wearable glucose sensors available. Microfluidic Sweat Sampling: Efficiently collects and delivers sweat to the sensor via laser-cut channels, ensuring accurate glucose measurement. | NO | NA | YES | NA | 2022 | Mirzajani, Hadi; Abbasiasl, Taher; Mirlou, Fariborz; Istif, Emin; Bathaei, Mohammad Javad; Dağ, Çağdaş; Deyneli, Oğuzhan; Yazıcı, Dilek; Beker, Levent | An ultra-compact and wireless tag for battery-free sweat glucose monitoring | Biosensors and Bioelectronics |  | 9565663 | Glucose monitoring before, during, and after exercise is essential for people with diabetes as exercise increases the risk of activity-induced hyper- and hypo-glycemic events. The situation is even more challenging for athletes with diabetes as they have impaired metabolic control compared to sedentary individuals. In this regard, a compact and noninvasive wearable glucose monitoring device that can be easily worn is critical to enabling glucose monitoring. This report presents an ultra-compact glucose tag with a footprint and weight of 1.2 cm2 and 0.13 g, respectively, for sweat analysis. The device comprises a near field communication (NFC) chip, antenna, electrochemical sensor, and microfluidic channels implemented in different material layers. The device has a flexible and conformal structure and can be easily attached to different body parts. The battery-less operation of the device was enabled by NFC-based wireless power transmission and the compact antenna. Femtosecond laser ablation was employed to fabricate a highly compact and flexible NFC antenna. The proposed device demonstrated excellent operating characteristics with a limit of detection (LOD), limit of quantification (LOQ), and sensitivity of 24 μM, 74 μM, and 1.27 μA cm 2 mM 1, respectively. The response of the proposed sensor in sweat glucose detection and quantification was validated by nuclear magnetic resonance spectroscopy (NMR). Also, the device’s capability in attachment to the body, sweat collection, and glucose measurement was demonstrated through in vitro and in vivo experiments, and satisfactory results were obtained. | 44835 | 2022
IGNEK2NR | 10.1016/j.bios.2022.114565 | https://linkinghub.elsevier.com/retrieve/pii/S0956566322006054 | YES | patch |  | NO | monitoring lactate levels in sweat | electrochemical | Electrochemical + optical | A LOD of 3.70 mM was calculated based on the RGB analysis of images captured during in vitro testing. This LOD was determined using a linear regression of data points collected after 120 seconds of exposure to varying lactate concentrations. The authors note that this LOD is sufficient for discrimination within the physiological sweat lactate range of 4-80 mM. The researchers employed RGB analysis of images to quantify the color changes of the electrochromic display. This approach provides a more reliable quantification method than naked-eye observation, especially for subtle color variations at lower lactate concentrations. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate |  |  |  |  | 1 | 1 | Prototype/Validation | Human | in-vitro tests + laboratory testing + human studies | NA | NA | The ESPS device integrates advanced features for wearable biosensing: Flexible Electronics: Fabricated on a flexible PET substrate using cost-effective screen printing, it employs customized conductive inks for transparent and efficient electrochromic display electrodes. Self-Powered Operation: A dual biofuel cell (BFC) system generates energy through lactate oxidation and oxygen reduction, eliminating external power needs and enhancing portability. Resettable Electrochromic Display: BFC-driven reversible color changes allow continuous lactate measurements, with modifications to bioanodes and biocathodes for efficient reactions. Integrated Sensing and Display: Combines biofuel cells, an electrochromic display, and a switch mechanism for direct, visual lactate readout without external devices. | NO | NA | NA | NA | 2022 | Hartel, Martin C.; Lee, Dongwook; Weiss, Paul S.; Wang, Joseph; Kim, Jayoung | Resettable sweat-powered wearable electrochromic biosensor | Biosensors and Bioelectronics |  | 9565663 | Wearable biofuel cells (BFCs) are being widely studied as self-powered sensors to monitor one’s physiological status non-invasively. However, the energy generated by such devices is often insufficient to power accompanying measurement read-out systems and communication protocols. Here, we report a reversible fully printed electrochromic self-powered sensor (ESPS) toward on-body sweat lactate monitoring. The device was realized by integrating a lactate-oxidizable BFC anode, oxygen-reducible BFC cathode, and a reversible Prussian blue (PB) electrochromic display on a single wearable patch. Through modification of the BFC electrodes, the formal potential of the PB display (0.14 V) could be positioned in between the potentials of the lactate oxidation ( 0.05 V) and oxygen reduction (0.35 V). Based on this system, the ESPS enables reversible and robust color changes with distinct time- and concentration-dependent profiles. The results demonstrate that the self-powered device has practical on-body applications for continual lactate monitoring with several advantages, including facile reversibility, ease of fabrication, and naked-eye colorimetric read-out. | 44866 | 2022
9SYQF7VI | 10.1016/j.bios.2022.113970 | https://linkinghub.elsevier.com/retrieve/pii/S0956566322000094 | YES | skin-attached |  | NO | real-time, continuous monitoring of lactate levels in sweat. | electrochemical | Electrochemical | Pearson Correlation Coefficient: 0.97 (comparison with reference analytical method, 20+ sweat samples). Deviation between wearable readings and reference: within 10%. Reference method: Integrated lactate concentration over 4–6 minutes. Wearable device: Real-time, continuous monitoring. Sensitivity (S): Enhanced by Prussian Blue nanozymes due to increased surface area and reduced diffusion limitations. Michaelis Constant (Kₘ): Modified by adding PFSI to siloxane membranes: Lower sensitivity but increased Kₘ. Adjusted for higher lactate concentrations in undiluted sweat. High correlation maintained across varied sweating rates. Accurate lactate level detection unaffected by sweat flow rate. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate |  |  |  |  | 1 | 1 | Validation | Human | In-vitro, Human-studies | NA | NA | Flow-Independent Biosensor: Co-immobilized Lactate Oxidase (LOD) and Prussian Blue Nanozymes ensure accurate lactate detection independent of sweat rate by creating a diffusion-controlled environment with a PFSI membrane. High-Sensitivity Nanozymes: Prussian Blue nanozymes offer enhanced sensitivity, reduced diffusion limitations, and extended operational life due to their large surface area and stability. Low-Power Controller (UMKA): Features include short-circuit current measurement, programmable gain amplifier, Bluetooth connectivity, and battery life up to 100 hours. Capillary Sweat System: Optimized capillary dimensions ensure self-driven sweat flow and close contact with the sensor for continuous and efficient analysis. | NO | NA | NA | NA | 2022 | Komkova, Maria A.; Eliseev, Artem A.; Poyarkov, Andrei A.; Daboss, Elena V.; Evdokimov, Pavel V.; Eliseev, Andrei A.; Karyakin, Arkady A. | Simultaneous monitoring of sweat lactate content and sweat secretion rate by wearable remote biosensors | Biosensors and Bioelectronics |  | 9565663 | We report on the simultaneous monitoring of sweat lactate concentration and sweat secretion rate. For this aim lactate oxidase-Prussian Blue enzyme-nanozyme type lactate biosensors were elaborated. The use of siloxaneperfluorosulfonated ionomer composite membrane for enzyme-nanozyme immobilization results in the biosensor displaying flux independence in the whole range of physiological sweat secretion rates (0.025–2 μl cm2 min-1). On the contrary, current response of the biosensor based on solely siloxane membranes becomes saturated at physiological sweat lactate concentration, depending mostly on the flow rate. Accordingly, for simultaneous monitoring of sweat lactate concentration and its secretion rate both flow-through biosensors were integrated with high-accuracy wearable electronic devices allowing real-time remote monitoring. As found, during exhaustive physical exercise sweat secretion rate and lactate content are independent of each other, thus, confirming that this excretory liquid is suitable for non-invasive diagnostics. | 44652 | 2022
MV4SJ9RX | 10.1016/j.bios.2020.112133 | https://linkinghub.elsevier.com/retrieve/pii/S0956566320301317 | YES | skin-attached |  | NO | non-invasive cortisol detection in sweat | electrochemical | Electrochemical | Sensor-Level Sensitivity: LOD of 1 pg/mL, sensitivity of 0.25 Ω/ng·mL⁻¹ (cortisol in artificial sweat). Integrated LOP Sensitivity: Sensitivity of 0.273 Ω/ng·mL⁻¹ after integrating with the microfluidic device. 3D Electrode Enhancement: Nanostructured gold electrodes improve sensitivity via increased antibody binding sites. Comparison with ELISA: Relative difference of 14.7% (±9.38%) based on sweat samples from three subjects. Sample Size: Limited evaluation with 6 samples from 3 subjects. Accuracy may be impacted by sensor drift, sweat composition variations, and environmental conditions. Selectivity against neuropeptide Y (NPY) was confirmed, but interference from other components needs further validation. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | Validation | Human | in-vitro tests + laboratory testing + human studies | NA | NA | Enhanced sensitivity via 3D gold electrodes with increased antibody binding sites. Mechanical stability on a flexible mogul-patterned substrate ensures durability during skin movement. Compact design for discreet wearability. Hydrophilic PVP-treated channels ensure efficient sweat flow. Zig-zag capillaries and advanced valves enable autonomous sweat metering, precise incubation, and one-touch reagent delivery. Simplifies protocols by eliminating labeling steps. High sensitivity and wide dynamic range suitable for physiological cortisol levels. | NO | NA | NA | NA | 2020 | Lee, Han-Byeol; Meeseepong, Montri; Trung, Tran Quang; Kim, Bo-Yeong; Lee, Nae-Eung | A wearable lab-on-a-patch platform with stretchable nanostructured biosensor for non-invasive immunodetection of biomarker in sweat | Biosensors and Bioelectronics |  | 9565663 | Conformable, wearable biosensor-integrated systems are a promising approach to non-invasive and quantitative on-body detection of biomarkers in body fluids. However, realizing such a system has been slowed by the difficulty of fabricating a soft affinity-based biosensor patch capable of precise on-body fluid handling with minimal wearer intervention and a simple measurement protocol. Herein, we demonstrate a conformable, wearable labon-a-patch (LOP) platform composed of a stretchable, label-free, impedimetric biosensor and a stretchable microfluidic device for on-body detection of the hormone biomarker, cortisol. The all-in-one, stretchable microfluidic device can precisely collect and deliver sweat for cortisol quantitation and offers one-touch operation of reagent delivery for simultaneous electrochemical signal generation and washing. Three-dimensional nanostructuring of the Au working electrode enables the high sensitivity required to detect the pM-levels of cortisol in sweat. Our integrated LOP detected sweat cortisol quantitatively and accurately during exercise. This LOP will open a new horizon for non-invasive, highly sensitive, and quantitative on-body immunodetection for wearable personal diagnostics. | 43952 | 2020
GM54SNAE | 10.1016/j.bios.2024.116280 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324002859 | YES | patch |  | minimally | continuous glucose monitoring in interstitial fluid (ISF). | electrochemical | Electrochemical | In-vitro Sensitivity: Achieves 77.721 nA mM⁻¹ mm⁻² over a glucose concentration range of 0–15 mM, with a linear correlation coefficient (R²) of 0.9987. Limit of Detection (LOD): Detects glucose concentrations as low as 36.437 μM (signal-to-noise ratio of 3). Differential Sensing Accuracy: Employs a differential sensor to reduce interference, improving the Pearson correlation coefficient from 0.5617 to 0.9882. Clarke Error Grid Analysis: In human trials, 89.47% of readings fell within Zone A (clinically acceptable), and 10.53% in Zone B, with no data in Zones C–E, ensuring high accuracy in real-world use. These metrics highlight the sensor’s sensitivity, reliability, and suitability for continuous glucose monitoring. | Diabetes | diabetes management | - | - | yes | interstitial fluid |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Validation | Human | in-vitro, ex-vivo, in-silico (FEM), animal models, and human studies, | NA | NA | Differential Microneedle Array: Combines acupuncture needles with PCB boards to integrate multiple sensing electrodes on a single microneedle system. Differential Sensing Electrodes: Uses two electrochemical three-electrode systems to eliminate interference, enhancing glucose detection accuracy. Wireless Data Transmission: Sends glucose data via Bluetooth to a smartphone app for real-time monitoring and analysis. Finite Element Modeling (FEM): Optimizes microneedle design for efficient skin penetration, reducing discomfort. Cylindrical Microneedles: Offers superior penetration efficiency with less force and improved biosafety. Prussian Blue-Based Glucose Detection: Utilizes Prussian blue for highly sensitive glucose detection, ensuring accurate measurements. | NO | NA | YES | NA | 2024 | Yang, Yong; Sheng, Can; Dong, Fang; Liu, Sheng | An integrated wearable differential microneedle array for continuous glucose monitoring in interstitial fluids | Biosensors and Bioelectronics |  | 9565663 | Monitoring biomarkers in human interstitial fluids (ISF) using microneedle sensors has been extensively studied. However, most of the previous studies were limited to simple in vitro demonstrations and lacked system integration and analytical performance. Here we report a miniaturized, high-precision, fully integrated wearable electrochemical microneedle sensing device that works with a customized smartphone application to wirelessly and in real-time monitor glucose in human ISF. A microneedle array fabrication method is proposed which enables multiple individually addressable, regionally separated sensing electrodes on a single microneedle system. As a demonstration, a glucose sensor and a differential sensor are integrated in a single sensing patch. The differential sensing electrodes can eliminate common-mode interference signals, thus significantly improving the detection accuracy. The basic mechanism of microneedle penetration into the skin was analyzed using the finite element method (FEM). By optimizing the structure of the microneedle, the puncture efficiency was improved while the puncture force was reduced. The electrochemical properties, biocompatibility, and system stability of the microneedle sensing device were characterized before human application. The test results were closely correlated with the gold standard (blood). The platform can be used not only for glucose detection, but also for various ISF biomarkers, and it expands the potential of microneedle technology in wearable sensing. | 45474 | 2024
G2KZXD8L | 10.1016/j.bios.2020.112828 | https://linkinghub.elsevier.com/retrieve/pii/S0956566320308149 | YES | patch |  | NO | simultaneously measure glucose and lactate levels in sweat | electrochemical | Electrochemical | Glucose Detection: Sensitivity: 2.4 nA μM⁻¹ for concentrations 0.08–1.25 mM. 1.0 nA μM⁻¹ for higher concentrations. Limit of Detection (LOD): 17.05 μM. Lactate Detection: Sensitivity: 0.49 μA mM⁻¹ within the physiological range (0.3–20.3 mM). LOD: 3.73 μM. Additional Performance Metrics: Reproducibility: Consistent responses across different sensor batches. Stability: Maintains electrochemical performance over 14 days. Selectivity: Minimal interference from other sweat components like ascorbic acid and uric acid. Comparison with Colorimetric Assay: Reported "acceptable differences" suggest potential accuracy, though no specific statistical values were provided. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | glucose, lactate |  |  |  |  | 1 | 2 | Proof-of-concept | Human | in-vitro tests + laboratory testing + human studies | NA | NA | Flexible All-Paper Substrate: Lightweight, biocompatible, and hydrophobically patterned for guided sweat flow and improved wearability. Three-Dimensional Folded Structure: Enhances sweat absorption, vertical diffusion, and electrode accessibility while avoiding direct skin contact. Advanced Sensing Materials: Ti₃C₂Tx MXene electrodes modified with methylene blue boost conductivity, adsorption, and electrochemical sensitivity. Enzyme Immobilization: Glucose oxidase and lactate oxidase provide selective biomarker detection with measurable signals. 3D Electrode Design: Optimized electrode placement ensures efficient enzyme immobilization and sweat contact for accurate sensing. Integrated System: Combines the sensor patch, a flexible circuit board, and smartphone connectivity for real-time biomarker monitoring. | NO | NA | YES | NA | 2021 | Li, Meng; Wang, Lichao; Liu, Rui; Li, Jianmin; Zhang, Qinghong; Shi, Guoying; Li, Yaogang; Hou, Chengyi; Wang, Hongzhi | A highly integrated sensing paper for wearable electrochemical sweat analysis | Biosensors and Bioelectronics |  | 9565663 | Wearable electrochemical sensors have attracted tremendous attention in recent years. Significant progress has been achieved, particularly in device integration. Most wearable devices are integrated on thin-film polymer, however, less attention is paid to the sweat flow at human-device interfaces, which is of great significance for continuous real-time analysis and long-term skin comfort. Here, we reported a low-cost, freestanding and disposable highly integrated sensing paper (HIS paper) for real-time analysis of sweat. By using a simple printing process, the HIS paper combining hydrophobic protecting wax, conducting electrodes, and the incorporated MXene/methylene blue (Ti3C2Tx/MB) active materials was assembled. In particular, the printed paper was folded into a multi-layer structure, in which a reasonable designed three-dimensional (3D) sweat diffusion path is established by connecting the hydrophilic regions of each layer, providing efficient pathways for the collection and diffusion of sweat along the vertical direction of the folded HIS paper. More importantly, the independent 3D position of three-electrode facilitates the decoration and fixation of enzymes, as well as the accessibility of electrolytes. In addition, a dual-channel electrochemical sensor that can simultaneously detect glucose and lactate with sensitivity of 2.4 nA μM 1 and 0.49 μA mM 1 respectively was produced based on the HIS paper. This HIS paper provides a miniaturized, low-cost and flexible solution for a range of biochemical platforms, including wearable bioelectronics. | 44228 | 2021
MCDSNJZK | 10.1002/smtd.202000926 | https://onlinelibrary.wiley.com/doi/10.1002/smtd.202000926 | YES | skin-attached |  | NO | monitoring glucose and lactate levels in sweat | electrochemical | Electrochemical | Glucose Sensor: Detection limit: 2 µM. Sensitivity: 31.02 µA mM⁻¹ cm⁻² within a linear range of 2 µM to 1 mM. Sensitivity with silk fibroin nanofibrils (SFNFs) is 1.58x higher than silk fibroin (SF) and 1.95x higher than chitosan. Lactate Sensor: Linear range: 0.4 to 6 mM. Sensitivity: 6.68 µA cm⁻² [log₁₀ (mM)]⁻¹. Stability and Selectivity: Operational stability: 25 hours (glucose) and 23.6 hours (lactate). Strong anti-interference against substances like uric acid, urea, and other analytes. Enhanced sensitivity arises from the PEN membrane and Pt-G nanocomposite, ensuring efficient analyte penetration, electrocatalytic activity, and reliable measurements. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | glucose, lactate |  |  |  |  | 1 | 1 | prototype | Human | human studies | NA | NA | Silk fibroin nanofibrils (SFNFs) to create a porous membrane that enhances enzyme stability and lifespan. An ultra-thin PtNPs/graphene (Pt-G) film to facilitate efficient electron transport, boosting sensitivity. A flexible and biocompatible polyurethane/silk fibroin (PSF) substrate ensuring comfort for the wearer. Integration into a wearable platform with wireless data transmission capabilities for convenient, real-time data analysis. | NO | NA | NA | NA | 2021 | Liu, Xiaotian; Zhang, Wenli; Lin, Zaifu; Meng, Zhaohui; Shi, Chenyang; Xu, Zhijun; Yang, Likun; Liu, Xiang Yang | Coupling of Silk Fibroin Nanofibrils Enzymatic Membrane with Ultra‐Thin PtNPs/Graphene Film to Acquire Long and Stable On‐Skin Sweat Glucose and Lactate Sensing | Small Methods |  | 2366-9608, 2366-9608 | Abstract The applications of enzymatic biosensors are largely limited by their relatively poor stability and short lifespan. Herein, a bio‐active porous enzymatic nanofiber (PEN) membrane composed of silk fibroin nanofibrils (SFNFs) and enzymes is developed to effectively retain the enzymes in the 3D space. The 3D functional scaffolds formed by SFNFs can immobilize enzymes and provide a large surface area for molecular/ion diffusion and biochemical reactions. The PEN membrane is subsequently attached to an ultra‐thin PtNPs/graphene (Pt‐G) nanocomposite film to facilitate the electron transport between the enzymes and electrodes, permitting highly effective glucose and lactate sensing with long and stable performance. The as‐assembled glucose and lactate sensors demonstrate high sensitivity, good cyclic reproducibility, and in particular long‐term stability of up to 25 and 23.6 h, respectively. These glucose sensors have a working life that is ≈1.25‐times longer than that of the best available sensors reported so far. Moreover, a wearable platform based on the sensors is developed for real‐time analysis of sweat during outdoor exercising to transmit signals to a mobile handset. The high sensitivity, comfort and long‐term stability of the device can benefit for long‐term in‐line surveillance of physiological conditions. | 44256 | 2021
X83KNB2P | 10.1016/j.bios.2024.116299 | https://linkinghub.elsevier.com/retrieve/pii/S095656632400304X | YES | patch |  | NO | monitoring of sweat rate and electrolytes detection | electrochemical | Electrochemical | Linearity with Syringe Pump: The Art-Sweat patch demonstrated excellent linearity when tested with a syringe pump delivering a 50 mM NaCl solution at known flow rates. Dynamic Flow Rate Detection: The patch reliably detected flow rates ranging from 1.0 to 4.0 μL/min with a rapid response time. Comparison with Absorbent Pads: The patch showed consistent trends with sweat rates determined using absorbent pads during on-body tests. Linear Relationship with Conductance: The device accurately measured ionic concentrations, showing a linear relationship for NaCl solutions ranging from 10 mM to 200 mM. |  | NA | yes | - | - | sweat | [12 empty] | electrolytes* (not specified) |  |  |  |  | 1 | NA* | Validation | Human | in-vitro + human studies | NA | NA | Resettable Microfluidic Channel with Droplet Separation: Combines horizontal microchannels with sweat absorption for continuous, long-term monitoring (capacity up to 333 μL), surpassing standard devices. Side-Wall Vent with Stop Valve: Ensures stable droplet formation, air pressure equilibrium, and prevents leakage for accurate data in motion-intensive scenarios. Dual-Mode Signal Decoupling: Simultaneously measures sweat rate (via interdigital electrodes) and ionic concentration (via conductance) from a single channel. Wireless Real-Time Feedback: A custom wireless module enables mobile app integration, offering real-time hydration monitoring, data visualization, and personalized management. Flexible and Precise Design: Employs laser-cut polyimide (PI) and PET layers for a conformable patch with potential high sensitivity. | NO | NA | YES | NA | 2024 | Liu, Mengyuan; Liu, Siyuan; Zhang, Tong; Zhou, Dengfeng; Li, Lianhui; Gao, Qiang; Liu, Yujie; Ge, Changlei; Wang, Yongfeng; Wang, Mingxu; Wen, Feng; Xiong, Zuoping; Zhou, Zhen; Wang, Shuqi; Zhang, Ting | Adaptively resettable microfluidic patch for sweat rate and electrolytes detection | Biosensors and Bioelectronics |  | 9565663 | Skin-interfaced microfluidic patch has become a reliable device for sweat collection and analysis. However, the intractable problems of emptying the microchannel for reuse, and the channel’s volumetric capacity limited by the size of the patch, directly hinder the practical application of sweat sensors. Herein, we report an adaptively resettable microfluidic sweat patch (Art-Sweat patch) capable of continuously monitoring both sweat rate (0.2–4.0 μL min 1) and total ionic charge concentration (10–200 mmol L 1). We develop a platform with a vertical and horizontal microchannel combined strategy, enabling repeatedly filling sweat and emptying the microchannel for autonomously resetting and detecting. The variation in the emptied volume is designed to be adaptively identified by the sensor, resulting in enhanced stability and an enlarged volumetric capacity of over 300 μL. By integrating with self-designed wireless transmission modules, the proposed Art-Sweat patch shows product-level wearability and high performance in monitoring variations in regional sweat rate and concentration for hydration status assessment. | 45505 | 2024
3HA757EG | 10.1016/j.aca.2023.342154 | https://linkinghub.elsevier.com/retrieve/pii/S0003267023013752 | YES | skin-attached |  | NO | detection of uric acid in sweat | electrochemical | Electrochemical | Sensitivity: The sensor demonstrates a sensitivity of 0.87 μA μM− 1 cm− 21. This value indicates a significant change in the measured current for each unit change in uric acid concentration. Limit of Detection (LOD): The LOD, determined based on a signal-to-noise ratio of 3, is 0.87 μM2. This value represents the lowest concentration of uric acid that the sensor can reliably distinguish from background noise. Linear Range: The sensor exhibits a linear response to uric acid concentrations ranging from 0.5 μM to 1.46 mM. This wide linear range covers the typical uric acid levels found in human sweat (1–43 μM), enabling accurate measurements across a broad spectrum of physiological conditions. Response Time: The sensor achieves 95% of the steady-state current in just 6 seconds, indicating a rapid response time. |  | NA | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | uric acid (UA) |  |  |  |  | 1 | 1 | Prototype/Validation |  | in-vitro + laboratory testing | NA | NA | Flexible and Stretchable Substrate: Uses a 3D graphene foam (GF) embedded in polydimethylsiloxane (PDMS), allowing for high stretchability (strain up to 110%), flexibility, and durability. High Sensitivity: Leveraging graphene's excellent conductivity and surface area for precise uric acid detection. Pt/Pd Nanoparticle Enhancement: Modifies the substrate with nanoparticles to boost electrocatalytic activity, increasing sensitivity and lowering detection limits. Dual Working Electrode Design: Eliminates interference through a novel configuration with detection and control electrodes for high selectivity. | NO | NA | NA | NA | 2024 | Lv, Mingrui; Qiao, Xiujuan; Li, Yanxin; Zeng, Xianghua; Luo, Xiliang | A stretchable wearable sensor with dual working electrodes for reliable detection of uric acid in sweat | Analytica Chimica Acta |  | 32670 |  | 45292 | 2024
5MB5SR5A | 10.1007/s10544-016-0142-3 | http://link.springer.com/10.1007/s10544-016-0142-3 | YES | skin-attached |  | NO | biomarkers detection | electrochemical | Electrochemical | Limit of Detection (LOD): The sensor detects BNP concentrations as low as 100 pg/mL, enabling early heart failure detection. Specificity: A 112-fold higher response to BNP compared to non-specific targets (e.g., BSA, IgG), ensuring accuracy. Sensitivity and Linearity: Exhibits a linear response to BNP concentrations from 100 pg/mL to 1 ng/mL, supporting quantification of BNP levels. Surface Morphology Contribution: Enhanced sensitivity is attributed to the nanogranular 2-D PANI structure, which increases the binding sites and charge gating efficiency. |  | NA | - | - | - | PBS | [12 empty] | BNP |  |  |  |  | 1 | 1 | Proof-of-concept |  | in-vitro + laboratory testing | NA | NA | Novel Material: Utilizes 2-D polyaniline nanostructures with high surface-area-to-volume (SA/V) ratios, enhancing sensitivity through increased antibody binding sites. Simplified Fabrication: Combines CMOS-compatible methods and chemical synthesis for cost-effective, uniform production on flexible substrates. Flexible Substrate: Built on polyimide (PI) for wearable applications, offering stable performance and predictable conductivity under bending conditions. Microfluidic Integration: Includes a PDMS microchannel for precise sample delivery, reducing noise and ensuring safety. High-Specificity Detection: Focuses on BNP biomarkers but can be adapted for other analytes via functionalization, broadening its application scope. | NO | NA | NA | NA | 2016 | Liu, Pei; Zhu, Yisi; Lee, Seung Hee; Yun, Minhee | Two-dimensional polyaniline nanostructure to the development of microfluidic integrated flexible biosensors for biomarker detection | Biomedical Microdevices |  | 1387-2176, 1572-8781 | In this work, we report a flexible field-effecttransistor (FET) biosensor design based on two-dimensional (2-D) polyaniline (PANI) nanostructure. The flexible biosensor devices were fabricated through a facile and inexpensive method that combines top-down and bottom-up processes. The chemically synthesized PANI nanostructure showed excellent p-type semiconductor properties as well as good compatibility with flexible design. With the 2-D PANI nanostructure being as thin as 80 nm and its extremely large surface-area-to-volume (SA/V) ratio due to the intrinsic properties of PANI chemical synthesis, the developed flexible biosensor exhibited outstanding sensing performance in detecting B-type natriuretic peptide (BNP) biomarkers, and was able to achieve high specificity (averagely 112 folds) with the limit of detection as low as 100 pg/mL. PANI nanostructure under bending condition was also investigated and showed controllable conductance changes being less than 20% with good restorability which may open up the possibility for wearable applications. | 42705 | 2016
HBCN9CMW | 10.1016/j.talanta.2020.121502 | https://linkinghub.elsevier.com/retrieve/pii/S0039914020307931 | YES | skin-attached |  | NO | real-time pH monitoring in sweat | electrochemical | Electrochemical | Sensitivity: The sensor exhibits a super-Nernstian response with a slope of -0.079 ± 0.002 V per unit of pH (RSD% = 2.5%). Comparison with Portable Potentiostat: Measurements using the integrated device were compared to those obtained with a laboratory-grade portable potentiostat. The results showed a strong correlation (R2= 0.997) between the two methods, confirming the accuracy of the wearable sensor. Agreement with pH Meter in Collected Sweat: The sensor, with a cotton pad simulating on-body conditions, measured a pH of 5.3 ± 0.3 in collected sweat, which was consistent with the value of 5.2 ± 0.2 obtained using a commercial pH meter. On-Body Measurement Validation: During a real-time on-body test, the device recorded a pH of approximately 5.5, which aligned with the pH meter measurement of 5.4 ± 0.2, further validating its accuracy in a real-world scenario7. Stability and Repeatability: The sensor exhibited good stability over a 2-hour continuous monitoring period, with only a slight decrease in potential observed. The device showed no significant hysteresis effect when tested across a range of pH values, indicating its reliable and repeatable performance. Researchers investigated various storage conditions to optimize the sensor's long-term stability. Storing the modified electrodes at -26°C for 20 days resulted in an unvaried potential response (RSD% = 1.2%), indicating good stability under low-temperature storage. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | pH |  |  |  |  | 1 | 1 | Prototype | Human | laboratory testing + human studies | NA | NA | The flexible pH sensor integrates several innovative features:Kapton Substrate: Combines flexibility and durability for a comfortable, long-term wearable design with integrated electrodes, circuitry, and RFID antenna. Screen-Printed Sensor: Cost-effective and scalable fabrication with precise deposition of conductive inks and a sensitive iridium oxide layer. High-Sensitivity Iridium Oxide: Ensures accurate detection of subtle pH changes in sweat, offering stability and reliability. Wireless RFID Data Transmission: Enables real-time monitoring or data logging with advantages in range and power autonomy over NFC and BLE. Miniaturized Electronics: Incorporates a sensor-oriented IC with a 10-bit ADC and temperature sensor, enhancing compactness and portability | NO | NA | NA | NA | 2021 | Mazzaracchio, Vincenzo; Fiore, Luca; Nappi, Simone; Marrocco, Gaetano; Arduini, Fabiana | Medium-distance affordable, flexible and wireless epidermal sensor for pH monitoring in sweat | Talanta |  | 399140 | In the last decade, wearable sensors have gained a key role on biomedical research field for reliable health state monitoring. A wide plethora of physics marker sensors is already commercially available, including activity tracker, heart rate devices, and fitness smartwatch. On the contrary, wearable and epidermal sensors for chemical biomarker monitoring in several biofluids are not ready yet. Herein, we report a wireless and flexible epidermal device for pH monitoring in sweat, fabricated by encompassing a screen-printed potentiometric sensor, an integrated circuit, and antenna embedded onto the same Kapton substrate. An iridium oxide film was electrodeposited onto the graphite working electrode providing the pH sensitive layer, while the integrated circuit board allows for data acquisition and storing. Furthermore, a radio frequency identification antenna surrounding the entire system enables data transmission to an external reader up to nearly 2 m in the most favourable case. The potentiometric sensor was firstly characterised by cyclic voltammetry experiments, then the iridium oxide electrodeposition procedure was optimised. Next, the sensor was tested toward pH detection in buffer solutions with a near-Nernstian response equal to 0.079 ± 0.002 V for unit of pH. Interference studies of common sweat ions, including Na+, K+ and Cl , showed any influence on the pH sensor response. Finally, the integrated epidermal device was tested for real-time on-body pH sweat monitoring during a running activity. Data recorded for a running subject were wireless transmitted to an external receiver, showing a pH value close to 5.5, in agreement with value obtained by pH-meter reference measurement. | 44197 | 2021
Q3RX89I4 | 10.1016/j.bios.2023.115198 | https://linkinghub.elsevier.com/retrieve/pii/S0956566323001409 | YES | skin-attached |  | NO | monitor levels of H2O2 and glucose | electrochemical | Electrochemical | For H2O2, the detection sensitivity is 11.74 nA mM− 1 mm− 2 in the range of 0.002–2 mM and 0.68 nA mM− 1 mm− 2 in the range of 2–100 mM. The limit of detection (LOD) for H2O2 is 1.88 μM.1 This LOD is reportedly superior to some previously reported H2O2 electrochemical sensors. For glucose, the sensitivity is reported as 0.19 nA μM− 1 with a linear range of 40–200 x 10−6 M. The LOD for glucose is 23 μM. This linear range encompasses typical sweat glucose concentrations in healthy individuals. The reported LOD is also noted to be superior to other carbon-based glucose sensors. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose, H2O2 |  |  |  |  | 1 | 2 | Prototype | Human | laboratory testing + in-vitro tests + human studies (limited, only 1 volunteer) | NA | NA | Biomimetic Enzyme Activity: Gold nanoparticles (AuNPs) within N-doped porous carbonized silk fibroin (CSF) mimic natural enzyme activity, enhancing sensitivity and stability for H₂O₂ and glucose detection. Flexible Silk Fibroin Substrate: A modified silk fibroin substrate ensures biocompatibility, flexibility, and a stable skin interface. Polyurethane Reinforcement: PU integration boosts flexibility, durability, and resistance to degradation for wearability. Porous N-Doped Carbon Scaffold: The high surface area and N-doping amplify sensitivity and catalytic performance. Synergy of Components: The combination of AuNPs and CSF delivers superior enzyme-like activity via optimized interactions. Miniaturized Wearable Design: The sensor supports sweat biomarker monitoring with potential Bluetooth connectivity for mobile integration. | NO | NA | NA | NA | 2023 | Meng, Guoqing; Long, Fen; Zeng, Zhicheng; Kong, Lingqing; Zhao, Bicheng; Yan, Jiaqi; Yang, Likun; Yang, Yun; Liu, Xiang-Yang; Yan, Zhengquan; Lin, Naibo | Silk fibroin based wearable electrochemical sensors with biomimetic enzyme-like activity constructed for durable and on-site health monitoring | Biosensors and Bioelectronics |  | 9565663 | Flexible biomimetic sensors have encountered a bottleneck of sensitivity and durability, as the sensors must directly work within complex body fluid with ultra-trace biomarkers. In this work, a wearable electrochemical sensor on a modified silk fibroin substrate is developed using gold nanoparticles hosted into N-doped porous carbonizated silk fibroin (AuNPs@CSF) as active materials. Taking advantage of the inherent biocompatibility and flexibility of CSF, and the high stability and enzyme-like catalytic activity of AuNPs, AuNPs@CSF-based sensor exhibits durable stability and superior sensitivity to monitor H2O2 released from cancer cell (4T1) and glucose in sweat. The detection limits for H2O2 and glucose are low to be 1.88 μM and 23 μM respectively, and the sensor can be applied in succession within 30 days at room temperature. Further, physical cross-linking of polyurethane (PU) with SF well matches with the skin tissue mechanically and provides a flexible, robust and stable electrode-tissue interface. AuNPs@CSF is applied successfully for wearable electrochemical monitoring of glucose in human sweat.The present AuNPs@CSF will possess a potential application in clinical diagnosing of H2O2- or glucose-related diseases in future. | 45047 | 2023
2MG8QZR7 | 10.1016/j.bios.2021.113765 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321008022 | YES | skin-attached |  | NO | non-invasive, real-time monitoring of sweat metabolites and electrolytes | electrochemical | Electrochemical | Sodium (Na+) Sensor: The Na+ sensor exhibits a linear relationship between the open-circuit potential and the logarithm of the Na+ concentration. The standard curve equation is: Y = 0.0196 × log2(X) + 0.1695, with an R² value of 0.9984. This indicates a high degree of correlation and sensitivity to changes in Na+ concentration. Potassium (K+) Sensor: The K+ sensor also demonstrates a linear response to the logarithm of the K+ concentration. Its standard curve is represented by the equation: Y = 0.0146 × log2(X) + 0.1773, with an R² value of 0.9957. Similar to the Na+ sensor, this signifies a strong correlation and sensitivity to K+ concentration variations. pH Sensor: The pH sensor exhibits a sub-Nernstian response with a sensitivity of 55.50 ± 6.31 mV/pH within the pH range of 5.91 to 8.04 at room temperature. This means that the sensor's output voltage changes by approximately 55.5 mV for every unit change in pH. Glucose Sensor: The glucose sensor's performance is described by a standard curve equation: Y = 0.0151 × log2(X) + 0.2071, with an R² value of 0.9862. This curve covers the normal range of glucose in sweat (0.33-0.65 mM), indicating the sensor's sensitivity and ability to detect glucose fluctuations within this clinically relevant range. |  | NA | yes | - | - | sweat |  |  |  |  | yes |  |  |  | yes |  | yes |  | Glucose, Na+, K+, pH, Urea |  |  |  |  | 1 | 5 | Prototype/Validation | Human | Laboratory testing + human studies | NA | NA | The Liquid Metal-Polymer Conductor (LMPC) patch introduces groundbreaking innovations for flexible, battery-free health monitoring: LMPC Material: Embeds liquid metal particles in a polymer matrix, enabling exceptional flexibility, stretchability, and biocompatibility while ensuring stable electrical performance during movement. LMPC-Based Antenna: Features a wireless, highly conductive antenna powered by LMPC, enabling communication and power transfer without bulky batteries or wires. Screen-Printed Ion-Selective Electrodes (ISEs): Enables affordable, miniaturized detection of glucose, sodium, and potassium in sweat through scalable mass production techniques. Multimodal Sensing Integration: Combines electrochemical ISEs with colorimetric sensors to monitor pH and urea, broadening biomarker analysis for a holistic health overview. Battery-Free Operation: Wireless power transfer enhances convenience, durability, and usability by eliminating battery reliance. | NO | NA | YES | NA | 2022 | Mou, Lei; Xia, Yong; Jiang, Xingyu | Liquid metal-polymer conductor-based wireless, battery-free epidermal patch | Biosensors and Bioelectronics |  | 9565663 | Wearable epidermal patch can seamlessly monitor biological signals in real-time. Here, we report a liquid metalpolymer conductor-based wireless epidermal patch. The epidermal patch is made of a new conductive material called liquid metal-polymer conductors (LMPC). LMPC is made by casting and peeling off polymers from patterned liquid metal particles. Our printable conductors present good stretchability, repeatability, and biocompatibility. We fabricate LMPC-based antenna and wire, which achieves wireless signal communication and power supply. To demonstrate the capability of our LMPC-based antenna and wire, we fabricate an epidermal patch to analyze metabolites, electrolytes, and urea in sweat. When a portable device is close to the epidermal sensor, the device can power and read the sensor through LMPC-based antenna. The epidermal patch exhibited good analytical performance for sweat analysis with a low limit of detection, fast response time, and multiplex detection capabilities. This epidermal patch opens the possibility for a broad range of non-invasive diagnostic tools that can be used for health monitoring in the general population. | 44593 | 2022
HDFXAHHP | 10.1093/ibd/izaa191 | https://www.ncbi.nlm.nih.gov/pubmed/32720974 | YES | wrist-worn | smartwatch-like | NO | monitoring specific biomarkers in eccrine sweat | Impedance spectroscopy | Impedance Spectroscopy | The device achieved a Pearson correlation of r = 0.99 for IL-1β and r = 0.95 for CRP compared to ELISA results. Bland-Altman Analysis showed a mean bias of -0.25 pg/mL for IL-1β and -3.9 pg/mL for CRP. | Inflammation | Inflammatory bowel disease (IBD). | yes | - | - | eccrine sweat | [12 empty] | Interleukin-1β (IL-1β), C-reactive protein (CRP) |  |  |  |  | 1 | 2 | Validation | Human | Human studies | NA | NA | A novel sensing platform that facilitates simultaneous multiplexed detection of biomarkers through the use of monoclonal antibodies and impedimetric detection modality. | EnLiSense LLC | USA | NA | NA | 2020 | Jagannath, B.; Lin, K. C.; Pali, M.; Sankhala, D.; Muthukumar, S.; Prasad, S. | A Sweat-based Wearable Enabling Technology for Real-time Monitoring of IL-1beta and CRP as Potential Markers for Inflammatory Bowel Disease | Inflamm Bowel Dis |  | 1536-4844 (Electronic) 1078-0998 (Linking) | BACKGROUND: More than 1.2 million people in the United States are affected by inflammatory bowel disease (IBD). Inflammatory bowel disease has a natural course characterized by alternating periods of remission and relapse. Currently, disease flares are unpredictable as they occur in a random way. Further, current testing methods and practices lack the ability for real-time tracking of flares. There exists no technology that can be utilized for continuous monitoring of biomarkers, as most of these rely on samples such as blood, feces, and testing methods by which continuous monitoring is not feasible. Cytokines play a key role in IBD; the development, recurrence, and exacerbation of the inflammatory process are orchestrated by their levels in time and space. Cytokines are also present in sweat. We hypothesize that demonstrating real-time continuous monitoring of interleukin-1beta (IL-1beta) and C-reactive protein (CRP) may help create an enabling technology to track inflammation in IBD patients and identify flare-ups and assess efficacy of therapy. METHODS: A multiplexed SWEATSENSER was used for noninvasive continuous monitoring of interleukin-1beta and C-reactive protein in human eccrine sweat. Impedance spectroscopy was used to measure the sensor response. Sweat was collected using an FDA-approved PharmChek patch from 26 healthy human subjects to determine the levels of the 2 study inflammatory markers. Correlation analysis was performed for preclinical validation of the SWEATSENSER with ELISA as the reference method. On-body continuous monitoring measurements were performed on 20 human subjects using EnLiSense's SWEATSENSER wearable device for real-time monitoring studies. RESULTS: The sensor device can detect interleukin-1beta and C-reactive protein in sweat over a dynamic range of 3 log orders. Pearson correlation of r = 0.99 and r = 0.95 was achieved for IL-1beta and CRP, respectively, for the SWEATSENSER with ELISA. Bland-Altman results further confirmed a good agreement (mean bias of -0.25 and -3.9 pg/mL for IL-1beta and CRP, respectively) of the device with the reference method, demonstrating applicability of the device for real-time monitoring. Continuous on-body measurements were performed in 20 healthy human subjects for the detection of IL-1beta to establish the preclinical utility of the sensor device. The continuous on-body measurements in healthy cohort reported a mean IL-1beta concentration of ~28 pg/mL. Stable measurements for over continuous 30 hours was reported by the device. CONCLUSION: This work demonstrates the first proof-of-feasibility of multiplexed cytokine and inflammatory marker detection in passively expressed eccrine sweat in a wearable form-factor that can be utilized toward better management of inflammatory bowel disease. This is a first step toward demonstrating a noninvasive enabling technology that can enable baseline tracking of an inflammatory response. Furthermore, this is the first study to report and quantify the presence of CRP in human eccrine sweat. | 44092 | 2020
72CRFVBG | 10.1016/j.talanta.2022.123695 | https://www.ncbi.nlm.nih.gov/pubmed/35728453 | YES | microneedle | hollow; array | Minimally | monitoring of glucose levels in interstitial fluid (ISF). | electrochemical | Electrochemical | The MN biosensor exhibited excellent stability (i.e., RSD = 1.9%) using interval monitoring for 7.5 h in a high concentration of glucose (i.e., 10 mM) and employing continuous monitoring mode (RSD = 6.8%). | Diabetes | Diabetes management | - | - | yes | interstitial fluid |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation |  | In-vitro + ex-vivo | NA | NA | Highly stable electrochemical biosensor based on a redox mediator bilayer consisting of Prussian blue and iron-nickel hexacyanoferrate to increase the long-term stability of the readout coupled with a hollow MN array as a sampling unit for ISF uptake. | NO | NA | NA | NA | 2022 | Parrilla, M.; Detamornrat, U.; Dominguez-Robles, J.; Donnelly, R. F.; De Wael, K. | Wearable hollow microneedle sensing patches for the transdermal electrochemical monitoring of glucose | Talanta |  | 1873-3573 (Electronic) 0039-9140 (Linking) | According to the World Health Organization, about 422 million people worldwide have diabetes, with 1.5 million deaths directly attributed each year. Therefore, there is still a need to effectively monitor glucose in diabetic patients for proper management. Recently, wearable patches based on microneedle (MN) sensors provide minimally invasive analysis of glucose through the interstitial fluid (ISF) while exhibiting excellent correlation with blood glucose. Despite many advances in wearable electrochemical sensors, long-term stability and continuous monitoring remain unsolved challenges. Herein, we present a highly stable electrochemical biosensor based on a redox mediator bilayer consisting of Prussian blue and iron-nickel hexacyanoferrate to increase the long-term stability of the readout coupled with a hollow MN array as a sampling unit for ISF uptake. First, the enzymatic biosensor is developed by using affordable screen-printed electrodes (SPE) and optimized for long-term stability fitting the physiological range of glucose in ISF (i.e., 2.5-22.5 mM). In parallel, the MN array is assessed for minimally invasive piercing of the skin. Subsequently, the biosensor is integrated with the MN array leaving a microfluidic spacer that works as the electrochemical cell. Interestingly, a microfluidic channel connects the cell with an external syringe to actively and rapidly withdraw ISF toward the cell. Finally, the robust MN sensing patch is characterized during in vitro and ex vivo tests. Overall, affordable wearable MN-based patches for the continuous monitoring of glucose in ISF are providing an advent in wearable devices for rapid and life-threatening decision-making processes. | 44866 | 2022
9H9B8TQX | 10.1002/smll.202306769 | https://onlinelibrary.wiley.com/doi/10.1002/smll.202306769 | YES | skin-attached |  | NO | real-time, continuous monitoring of multiple biomarkers in sweat. | electrochemical | Electrochemical | Potassium (K+): The potassium ion-selective sensor showed a linearity of 0.989 within its detection range, with a sensitivity of 48 mV per decade of concentration. Sodium (Na+): The sodium ion-selective sensor exhibited a linearity of 0.999 and a sensitivity of 87.9 mV per decade of concentration. Glucose: The glucose amperometric sensor had a linearity of 0.956 within its detection range and a sensitivity of 1.3 nA μm−1. Lactate: The lactate amperometric sensor demonstrated a linearity of 0.988 and a sensitivity of 210 nA−1. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | K+, Na+, Lactate, Glucose |  |  |  |  | 1 | 4 | Prototype/Validation | Human | laboratory testing + human studies. | NA | NA | Tree-Bionic Microfluidic Chip: Biomimetic design mimics tree transpiration, with root-like collection (covering ~200 sweat glands), trunk-like flow stabilization, and leaf-like discharge to enhance efficiency and prevent blockage. Ultra-Thin Elastic Sensor: A 20 μm MEMS-fabricated sensor integrates potentiometric and amperometric detection for multiple biomarkers (K⁺, Na⁺, glucose, lactate) with high flexibility, sensitivity, and robustness. Elastic Circuit Board (ECB): A custom elastic PCB enables stable, low-impedance signal processing and wireless transmission, maintaining flexibility during movement. Integrated Micro-Heater: A PI-Pt-PI micro-heater boosts sweat secretion and stabilizes temperature, ensuring consistent sensor performance. Advanced Data Processing: Incorporates FFT filtering for noise reduction and BLE for seamless real-time data transfer to a user interface. | NO | NA | YES | NA | 2024 | Niu, Jiaqi; Lin, Shujing; Chen, Di; Wang, Zhitao; Cao, Cheng; Gao, Ang; Cui, Shengsheng; Liu, Yanlei; Hong, Yuping; Zhi, Xiao; Cui, Daxiang | A Fully Elastic Wearable Electrochemical Sweat Detection System of Tree‐Bionic Microfluidic Structure for Real‐Time Monitoring | Small |  | 1613-6810, 1613-6829 | Abstract Fresh sweat contains a diverse range of physiological indicators that can effectively reflect changes in the body. However, existing wearable sweat detection systems face challenges in efficiently collecting and detecting fresh sweat in real‐time. Additionally, they often lack the necessary deformation capabilities, resulting in discomfort for the wearer. Here, a fully elastic wearable electrochemical sweat detection system is developed that integrates a sweat‐collecting microfluidic chip, a multi‐parameter electrochemical sensor, a micro‐heater, and a sweat detection elastic circuit board system. The unique tree‐bionic structure of the microfluidic chip significantly enhances the efficiency of fresh sweat collection and discharge, enabling real‐time detection by the electrochemical sensors. The sweat multi‐parameter electrochemical sensor offers high‐precision and high‐sensitivity measurements of sodium ions, potassium ions, lactate, and glucose. The electronic system is built on an elastic circuit board that matches perfectly to wrinkled skin, ensuring improved wearing comfort and enabling multi‐channel data sampling, processing, and wireless transmission. This state‐of‐the‐art system represents a significant advancement in the field of elastic wearable sweat detection and holds promising potential for extending its capabilities to the detection of other sweat markers or various wearable applications. | 45352 | 2024
H7K9JYDI | 10.1016/j.ijbiomac.2023.124757 | https://linkinghub.elsevier.com/retrieve/pii/S0141813023016513 | YES | skin-attached |  | NO | non-invasive detection of sweat urea | electrochemical | Electrochemical + Laser Desorption/Ionization Mass Spectrometry (LDI-MS) | The carbonized AuNPs-silk wearable skin patch exhibits promising accuracy and sensitivity for sweat urea detection. The device boasts a linear detection range of 0–100 mM for urea, covering the clinically relevant cut-off value of 60 mM for chronic kidney disease (CKD). The device employs two detection methods: Electrochemical sensor: This method demonstrates a limit of detection (LOD) of 20 mM. LDI-MS: This technique offers a higher sensitivity, achieving an LOD of 8 mM. Further, the study tested the patch's performance in artificial sweat, revealing accurate detection of spiked urea concentrations at 20 mM, 60 mM (cut-off), and 80 mM. With standard deviation (SD) values below 10, the platform showcased high precision in these analyses | Metabolic Syndrome | chronic kidney disease (CKD) management | yes | - | - | sweat | [8 empty] | yes |  |  |  | urea |  |  |  |  | 1 | 1 | proof-of-concept |  | in-vitro test + laboratory testing | NA | NA | Self-Reduction of AuNPs: Tyrosine in silk facilitates the self-reduction of gold nanoparticles, eliminating external reducing agents and lowering production costs. Carbonization for Conductivity: High-temperature carbonization transforms silk into a highly conductive carbon structure, enabling advanced electrochemical sensing. Dual-Detection System: Combines electrochemical sensing with LDI-MS for precise and confirmatory sweat urea analysis. AuNPs as Self-Matrix for LDI-MS: Gold nanoparticles enhance conductivity and serve as a self-matrix, improving ionization efficiency without external matrices. Flexible Wearable Substrate: Silk ensures biocompatibility, flexibility, and seamless skin conformity for effective sweat collection and analysis. | NO | NA | NA | NA | 2023 | Promphet, Nadtinan; Phamonpon, Wisarttra; Karintrithip, Wimala; Rattanawaleedirojn, Pranee; Saengkiettiyut, Kanokwan; Boonyongmaneerat, Yuttanant; Rodthongkum, Nadnudda | Carbonization of self-reduced AuNPs on silk as wearable skin patches for non-invasive sweat urea detection | International Journal of Biological Macromolecules |  | 1418130 | Flexible conductive skin patches were readily fabricated on silk fabric by in situ deposition of gold nanoparticles (AuNPs) followed by carbonization. The carbonized AuNPs-silk with high flexibility was characterized by field emission scanning electron microscopy (FESEM), atomic force microscopy (AFM) and Fourier transform Raman spectroscopy (FT-Raman) to verify the well arrangement surface and desired chemical binding. The conductivity of silk skin patch, measured by a four-point probe, was found to be 109.24 ± 13 S cm 1 × 10 3, verifying the potential application as a working electrode in electrochemical sensor and a sweat collection patch for direct detection by laser desorption/ionization mass spectrometry (LDI-MS). This silk skin patch offered a linear range of 0–100 mM with a detection limit (LOD) of 20 mM for electrochemical sensor and 8 mM for LDI-MS, respectively. Ultimately, this skin patch is successfully applied for the detection of sweat urea at its cut-off value (60 mM) for indicating chronic kidney disease (CKD) in artificial sweat with satisfactory results. By using dual-detection technique on single silk substrate, this platform might be an alternative approach for a noninvasive sweat urea detection with high precision. | 45108 | 2023
F6ADUP3R | 10.1109/TBME.2014.2369991 | http://ieeexplore.ieee.org/document/6953152/ | YES | skin-attached |  | NO | monitor electrolyte levels in sweat | electrochemical | Electrochemical (potentiometric) | Sensitivity: the sensor demonstrated a sensitivity of 0.3 mV/mM when integrated into the patch. This means that for every 1 mM increase in sodium concentration, the sensor output voltage increased by 0.3 mV. While slightly lower than the sensitivity of a commercial sensor (0.5 mV/mM), stand-alone sensors exhibited a near-Nernstian behavior of approximately 57 mV/decade of Na+ concentration, suggesting the limitation might stem from the RFID chip's ADC. Accuracy: at a sodium concentration of 50 mM, the patch sensor showed an accuracy of 96%. This was determined by comparing the sensor's output voltage to the expected value based on a calibration curve. Precision: the article reports a precision of 28% for the patch sensor. Precision represents the variability or spread of measurements. |  | NA | yes | - | - | sweat | [12 empty] | Na+ |  |  |  |  | 1 | 1 | prototype |  | in-vitro tests + laboratory testing | NA | NA | Flexible Printed Circuit Board (FPCB): a Dupont Pyralux-based substrate ensures body-conforming flexibility and comfort, reducing sweat pooling. Solder Reflow Integration: surface-mount packaging creates a slim, unobtrusive design by integrating components directly onto the FPCB. Electroplated Ion-Selective Electrodes (ISEs): palladium and silver layers enhance sensitivity for potentiometric sodium ion detection in sweat. Hermetic Sealing with Parylene C: biocompatible coating protects electronics from sweat, ensuring durability and device longevity. Paper Microfluidics: a porous layer wicks sweat to electrodes efficiently, improving response times for sweat analysis. Battery-Free Operation: RFID chip energy harvesting eliminates batteries, enhancing wearability and extending device lifespan. | NO | NA | YES | NA | 2015 | Rose, Daniel P.; Ratterman, Michael E.; Griffin, Daniel K.; Hou, Linlin; Kelley-Loughnane, Nancy; Naik, Rajesh R.; Hagen, Joshua A.; Papautsky, Ian; Heikenfeld, Jason C. | Adhesive RFID Sensor Patch for Monitoring of Sweat Electrolytes | IEEE Transactions on Biomedical Engineering |  | 0018-9294, 1558-2531 | Wearable digital health devices are dominantly found in rigid form factors such as bracelets and pucks. An adhesive radio-frequency identiﬁcation (RFID) sensor bandage (patch) is reported, which can be made completely intimate with human skin, a distinct advantage for chronological monitoring of biomarkers in sweat. In this demonstration, a commercial RFID chip is adapted with minimum components to allow potentiometric sensing of solutes in sweat, and surface temperature, as read by an Android smartphone app with 96% accuracy at 50 mM Na+ (in vitro tests). All circuitry is solder-reﬂow integrated on a standard Cu/polyimide ﬂexible-electronic layer including an antenna, but while also allowing electroplating for simple integration of exotic metals for sensing electrodes. Optional paper microﬂuidics wick sweat from a sweat porous adhesive allowing ﬂow to the sensor, or the sensor can be directly contacted to the skin. The wearability of the patch has been demonstrated for up to seven days, and includes a protective textile which provides a feel and appearance similar to a standard Band-Aid. Applications include hydration monitoring, but the basic capability is extendable to other mM ionic solutes in sweat (Cl−, K+ , Mg2+ , NH4 + , and Zn2+ ). The design and fabrication of the patch are provided in full detail, as the basic components could be useful in the design of other wearable sensors. | 42156 | 2015
3GEB3FTY | 10.1016/j.aca.2024.342761 | https://linkinghub.elsevier.com/retrieve/pii/S0003267024005622 | YES | body strap worn on the abdomen or lower back |  | NO | monitor glucose levels in sweat | electrochemical | Electrochemical | The wearable sweat glucose sensing device exhibits a linear detection range of 0.1–1.5 mM of glucose, with a detection limit of 0.1 mM. This detection range corresponds with typical sweat glucose levels, making it suitable for monitoring fluctuations in glucose concentrations. The device demonstrated a high correlation coefficient (R²) of 0.997 during calibration using a commercial potentiostat. Furthermore, when tested with artificial sweat spiked with standard glucose at low, medium, and high concentrations, the device achieved an outstanding recovery percentage ranging from 99% to 102%. | Diabetes | diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | proof-of-concept | Human | laboratory testing + human studies (preliminary on-body trials) | NA | NA | Nanomaterial-Modified Electrode: a CNT-CNF nanocomposite boosts conductivity, surface area, and enzyme loading for high glucose sensitivity. Prussian Blue Electrodeposition: enhances redox activity and electrocatalysis, improving the detection of H₂O₂ from glucose oxidation. Chitosan Coating: provides a biocompatible layer for enzyme immobilization, ensuring stability and extended sensor lifespan. Flexible Substrate: a polycarbonate base conforms comfortably to the body, ensuring seamless wearability. Machine Learning Integration: XGBoost and SHAP optimize fabrication parameters, maximizing sensor performance with cutting-edge predictive modeling. Smartphone Integration: a portable potentiostat and Bluetooth-linked app enable real-time glucose monitoring and display. | NO | NA | YES | NA | 2024 | Promphet, Nadtinan; Thanawattano, Chusak; Buekban, Chatchai; Laochai, Thidarut; Lormaneenopparat, Panlop; Sukmas, Wiwittawin; Rattanawaleedirojn, Pranee; Puthongkham, Pumidech; Potiyaraj, Pranut; Leewattanakit, Worapong; Rodthongkum, Nadnudda | Smartphone based wearable sweat glucose sensing device correlated with machine learning for real-time diabetes screening | Analytica Chimica Acta |  | 32670 |  | 45474 | 2024
QDZIKGDT | 10.1016/j.aca.2024.342441 | https://linkinghub.elsevier.com/retrieve/pii/S0003267024002423 | YES | skin-attached |  | NO | simultaneous detection of uric acid (UA) and dopamine (DA) in sweat | electrochemical | Electrochemical | For Dopamine (DA): Linear Response Range: 1–500 μM Detection Limit: 9.41 μM Sensitivity: 0.019 μA μM−1 cm−2 For Uric Acid (UA): Linear Response Range: 10–1000 μM Detection Limit: 9.09 μM Sensitivity: 0.026 μA μM−1 cm−2 The researchers conducted various tests, including: Anti-interference experiments to evaluate the sensor's selectivity in the presence of other potential interferents in sweat. Repeatability and reproducibility tests to assess the sensor's consistency in providing accurate readings. Long-term stability tests to determine the sensor's performance over extended periods. Real sweat sample analysis comparing the sensor's readings with those obtained from a standard uric acid detection kit, demonstrating good agreement between the two methods. While the sensor showed good performance for UA detection in real sweat samples, the researchers note that further development is needed to improve the sensitivity for dopamine detection, as its concentration in sweat is typically below the current detection limit. Overall, the wearable sensor exhibits promising accuracy and sensitivity for UA detection, paving the way for its potential use in real-time sweat analysis and health monitoring applications. |  | NA | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | Dopamine, Uric Acid |  |  |  |  | 1 | 2 | Prototype/Validation | Human | laboratory testing + human studies | NA | NA | High-Performance Nanomaterials: a CuNi-MOF@rGO nanocomposite enhances sensitivity and selectivity for uric acid and dopamine detection, leveraging high surface area, porosity, and excellent conductivity. Flexible Substrate: built on a durable, flexible polyimide (PI) substrate, the sensor ensures comfort, thermal stability, and chemical resistance for wearable use. Biomimetic Microfluidics: a 3D-printed sweat collection system mimics biological designs, featuring circular chambers and hydrophilic channels for efficient, reliable sweat sampling. Enhanced Hydrophilicity: silwet-L77 modification boosts sweat flow within microfluidic channels, optimizing collection efficiency. Real-Time Wearable Sensing: the device enables precise, non-invasive, real-time analysis of key biomarkers in sweat. Health Monitoring Applications: this cutting-edge design supports personalized healthcare and advanced physiological monitoring. | NO | NA | NA | NA | 2024 | Wang, Cuncun; Zhang, Yong; Liu, Yiyi; Zeng, Xin; Jin, Changpeng; Huo, Danqun; Hou, Jingzhou; Hou, Changjun | A wearable flexible electrochemical biosensor with CuNi-MOF@rGO modification for simultaneous detection of uric acid and dopamine in sweat | Analytica Chimica Acta |  | 32670 | Background: In health assessment and personalized medical services, accurate detection of biological markers such as dopamine (DA) and uric acid (UA) in sweat is crucial for providing valuable physiological information. However, there are challenges in detecting sweat biomarkers due to their low concentrations, variations in sweat yield among individuals, and the need for efficient sweat collection. | 45383 | 2024
AFBA9PGG | 10.1016/j.bios.2021.113078 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321001159 | YES | skin-attached |  | NO | non-invasive blood glucose monitoring | electrochemical | Electrochemical | Sensitivity of the G/CNTs/GOx Composite Textile: when tested in a three-electrode system using PBS solution, the G/CNTs/GOx composite textile (the working electrode) demonstrated high sensitivity to glucose: 574.26 μA mM−1 cm−2 for glucose concentrations ranging from 0 to 0.1 mM. 12.11 μA mM−1 cm−2 for glucose concentrations ranging from 1.0 to 30 mM. Sensitivity of the Two-Electrode Glucose Sensor: the fully assembled two-electrode sensor, incorporating both the G/CNTs/GOx working electrode and the G/CNTs/Ag/AgCl counter electrode, exhibited the following sensitivities: 1.39 to 1.81 μA μM−1 cm−2 in PBS solution. 14.45 ± 2.97 μA mM−1 cm−2 when tested on porcine skin using the reverse iontophoresis extraction method. Limit of Detection: the two-electrode glucose sensor achieved a limit of detection (LOD) of 0.06 μM for glucose. Correlation with Invasive Glucose Meter: the research demonstrated a high degree of consistency between the glucose concentrations measured by the non-invasive wearable sensor and those obtained using a standard invasive glucose meter in tests on both nude mice and human subjects. This correlation strongly suggests the accuracy and reliability of the device's glucose readings. | Diabetes | diabetes management | - | - | yes | interstitial fluid |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Validation | Human | In-vitro tests + ex-vivo tests + animal models + human studies | NA | NA | Two-Electrode System: Uses a single pair of electrodes for both ISF extraction via reverse iontophoresis and glucose detection, simplifying the design. Graphene/Carbon Nanotubes Composite Textiles: Provides excellent conductivity and flexibility, enabling efficient electrochemical reactions and wearability. Simple Fabrication: The composite material is synthesized using a two-step chemical vapor deposition process for controlled growth and enzyme immobilization. High Sensitivity and Accuracy: Demonstrates precise glucose measurements with sensitivities comparable to commercial glucose meters. Semi-Continuous Monitoring: Allows intermittent glucose monitoring for extended periods, tested in human trials. | NO | NA | NA | NA | 2021 | Yao, Yao; Chen, Jingyao; Guo, Yuhan; Lv, Tian; Chen, Zilin; Li, Ning; Cao, Shaokui; Chen, Bingdi; Chen, Tao | Integration of interstitial fluid extraction and glucose detection in one device for wearable non-invasive blood glucose sensors | Biosensors and Bioelectronics |  | 9565663 | Wearable non-invasive glucose sensors that can provide human a painless and portable means to monitor their blood glucose and manage their health condition draw great attentions, recently. Non-invasive human glucose sensors by detecting glucose in interstitial fluid (ISF) extracted through a reverse iontophoresis (RI) approach have been widely investigated, but the current challenges are their complex structure and instability for continuous monitor. Herein, we demonstrate a simple two-electrode non-invasive blood glucose sensor, which is fabricated by using graphene/carbon nanotubes/glucose oxidase composite textile and graphene/carbon nanotube/silver/silver chloride composite textile as the working electrode and counter electrode, respectively. By using one single device, extraction of ISF through RI process is firstly conducted by loading a certain electric current between two electrodes, then the glucose concentration in the ISF is detected through an amperometric approach by using the same two electrodes. The feasibility of these non-invasive glucose sensors is validated on porcine skin, nude mice and human. The blood glucose concentration calculated according to the response currents of the two-electrode sensors is highly consistent with that measured by commercial glucose meter. Furthermore, the used textile-like electrodes provide the non-invasive blood glucose sensors with excellent flexible and wearable properties, which make them promising to be integrated with other electronic units for monitor and management of human health. | 44317 | 2021
6NEEUXNI | 10.1016/j.aca.2023.341754 | https://linkinghub.elsevier.com/retrieve/pii/S0003267023009753 | YES | patch |  | NO | real-time detection of glucose levels in human sweat | electrochemical | Electrochemical | The source states that the Ni-Co MOF modified CMP (NCMP) electrode achieved high sensitivity for glucose detection with a low detection limit of 6.78 μM and a wide linear range of 20 μM - 1.1 mM. The NCMP film electrode-based flexible sensor has a sensitivity of 71.62 μA mM− 1 cm− 2 for glucose detection | Diabetes | diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Proof-of-concept | Human | In-vitro tests + human studies | NA | NA | SVFDT Method: a novel, cost-effective fabrication approach combining stamping and vacuum filtration to create flexible three-electrode arrays, enabling scalability and reducing dependency on specialized equipment. Flexible Substrates: incorporates stretchable PDMS for enhanced comfort and adaptability to body movements. Advanced Nanomaterials: uses MWCNTs to boost electrode conductivity and durability under mechanical stress. Enzyme-like Ni-Co MOF: features a highly porous, electrocatalytic material for glucose oxidation, providing high sensitivity and low detection limits. | NO | NA | NA | NA | 2023 | Xia, Youyuan; Su, Tong; Mi, Ziyi; Feng, Ziyou; Hong, Yawen; Hu, Xiaoya; Shu, Yun | Wearable electrochemical sensor based on bimetallic MOF coated CNT/PDMS film electrode via a dual-stamping method for real-time sweat glucose analysis | Analytica Chimica Acta |  | 32670 |  | 45200 | 2023
PLTIUXQL | 10.1177/2472630318759257 | https://www.ncbi.nlm.nih.gov/pubmed/29447045 | YES | skin-attached |  | NO | monitoring stress | electrochemical | Electrochemical impedance spectroscopy (EIS) | Detection of cortisol was shown for concentrations of 1 pg/mL to 200 ng/mL in both synthetic and perspired human sweat, |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | prototype |  | Laboratory testing | NA | NA | A four-channel electrochemical impedance spectroscopy (EIS) analyzer module. Flexible chemi-impedance sensors fabricated with gold electrodes for wearable applications. | NO | NA | NA | NA | 2018 | Sankhala, D.; Muthukumar, S.; Prasad, S. | A Four-Channel Electrical Impedance Spectroscopy Module for Cortisol Biosensing in Sweat-Based Wearable Applications | SLAS Technol |  | 2472-6311 (Electronic) 2472-6303 (Linking) | A four-channel electrochemical impedance spectroscopy (EIS) analyzer module has been demonstrated on flexible chemi-impedance sensors fabricated with gold electrodes for wearable applications. The module can perform time division multiplexed (TDM) impedance measurements on four sensors at 1 kHz. In this work, we characterize the system for the detection of cortisol in an ultralow volume (1-3 microL) of perspired human sweat, sensor performance, and effects during continuous cortisol dosing and with pH and temperature variations expected on the surface of the skin that would be representative of use conditions as seen by a wearable device. Detection of cortisol was shown for concentrations of 1 pg/mL to 200 ng/mL in both synthetic and perspired human sweat, and output response reported as a change in impedance with varying cortisol concentrations. Continuous dose testing was performed to demonstrate the stability of the sensors over prolonged periods of operation for cortisol concentrations within the physiologically relevant range of 10-200 ng/mL reported in human sweat. Temperature and pH effects testing was performed for pH range 4-8 and in a temperature chamber for the clinical range reported on the surface of human skin: 25-40 degrees C. The cortisol sensor demonstrated stability of operation with 7.58% variability under these conditions. | 43435 | 2018
Z322KISI | 10.1126/sciadv.aau6356 | https://www.ncbi.nlm.nih.gov/pubmed/30746456 | YES | skin-attached |  | NO | monitoring of physiological status | electrochemical | Electrochemical + optical | The calibration curve indicates the ability to measure chloride levels with an accuracy of ±10mM |  | NA | yes | - | - | sweat | [12 empty] | Chlorides |  |  |  |  | 1 | 1 | validation | Human | Human studies | NA | NA | The device incorporates thermographic capabilities, enabling simultaneous measurement of skin temperature along with sweat analysis | NO | NA | NA | NA | 2019 | Reeder, J. T.; Choi, J.; Xue, Y.; Gutruf, P.; Hanson, J.; Liu, M.; Ray, T.; Bandodkar, A. J.; Avila, R.; Xia, W.; Krishnan, S.; Xu, S.; Barnes, K.; Pahnke, M.; Ghaffari, R.; Huang, Y.; Rogers, J. A. | Waterproof, electronics-enabled, epidermal microfluidic devices for sweat collection, biomarker analysis, and thermography in aquatic settings | Sci Adv |  | 2375-2548 (Electronic) 2375-2548 (Linking) | Noninvasive, in situ biochemical monitoring of physiological status, via the use of sweat, could enable new forms of health care diagnostics and personalized hydration strategies. Recent advances in sweat collection and sensing technologies offer powerful capabilities, but they are not effective for use in extreme situations such as aquatic or arid environments, because of unique challenges in eliminating interference/contamination from surrounding water, maintaining robust adhesion in the presence of viscous drag forces and/or vigorous motion, and preventing evaporation of collected sweat. This paper introduces materials and designs for waterproof, epidermal, microfluidic and electronic systems that adhere to the skin to enable capture, storage, and analysis of sweat, even while fully underwater. Field trials demonstrate the ability of these devices to collect quantitative in situ measurements of local sweat chloride concentration, local sweat loss (and sweat rate), and skin temperature during vigorous physical activity in controlled, indoor conditions and in open-ocean swimming. | 43466 | 2019
MVR2F84L | 10.1039/d0ay02014a | https://www.ncbi.nlm.nih.gov/pubmed/33290448 | YES | skin-attached |  | NO | monitoring of sweat electrolyte concentrations during exercise | electrochemical | Electrochemical detection (ion chromatography). | Standard deviation experiments were conducted to quantify the accuracy of the IC method. Standard solutions with varying concentrations (0.1-100 ppm) of Na+, NH4+, K+, and Cl− ions were prepared and analyzed to determine the standard deviation at different concentration levels. The relative standard deviations (RSDs) for Cl− and Na+ were relatively small. For K+, the RSDs were more significant, likely because the K+ concentrations in sweat are about ten times lower than Na+ and Cl−. |  | NA | yes | - | - | sweat | [12 empty] | Na+, Cl-, K+. |  |  |  |  | 1 | 3 | Validation | Human | Human studies | NA | NA | A novel design of a sweat collection system is created. The sweat collection patch, that is made from hydrophilized foil and a double-sided acrylate adhesive, consists of a reservoir array that collects samples consecutively in time. | NO | NA | NA | NA | 2020 | Steijlen, A. S. M.; Bastemeijer, J.; Groen, P.; Jansen, K. M. B.; French, P. J.; Bossche, A. | A wearable fluidic collection patch and ion chromatography method for sweat electrolyte monitoring during exercise | Anal Methods |  | 1759-9679 (Electronic) 1759-9660 (Linking) | This paper presents a method to continuously collect and reliably measure sweat analyte concentrations during exercise. The method can be used to validate newly developed sweat sensors and to obtain insight into intraindividual variations of sweat analytes in athletes. First, a novel design of a sweat collection system is created. The sweat collection patch, that is made from hydrophilized foil and a double-sided acrylate adhesive, consists of a reservoir array that collects samples consecutively in time. During a physiological experiment, sweat can be collected from the back of a participant and the filling speed of the collector is monitored by using a camera. After the experiment, Na+, Cl- and K+ levels are measured with ion chromatography. Sweat analyte variations are measured during exercise for an hour at three different locations on the back. The Na+ and Cl- variations show a similar trend and the absolute concentrations vary with the patch location. Na+ and Cl- concentrations increase and K+ concentrations seem to decrease during this exercise. With this new sweat collection system, sweat Na+, Cl- and K+ concentrations can be collected over time during exercise at medium to high intensity, to analyse the trend in electrolyte variations per individual. | 44188 | 2020
9ILIEEN9 | 10.1002/advs.202201507 | http://www.ncbi.nlm.nih.gov/pubmed/35657078 | YES | textile |  | NO | TEXTILE CHARACTERIZATION (WOUND-HEALING APPLICATIONS - NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [18 empty] | 1 | [7 empty] | NO | NA |  |  | 2022 | Chen, Shuo-Wen; Huang, Shih-Min; Wu, Han-Song; Pan, Wei-Pang; Wei, Shih-Min; Peng, Chih-Wei; Ni, I. Chih; Murti, Bayu Tri; Tsai, Meng-Lin; Wu, Chih- I.; Yang, Po-Kang | A Facile, Fabric Compatible, and Flexible Borophene Nanocomposites for Self-Powered Smart Assistive and Wound Healing Applications | Advanced Science (Weinheim, Baden-Wurttemberg, Germany) |  | 2198-3844 | Smart fabrics that can harvest ambient energy and provide diverse sensing functionality via triboelectric effects have evoked great interest for next-generation healthcare electronics. Herein, a novel borophene/ecoflex nanocomposite is developed as a promising triboelectric material with tailorability, durability, mechanical stability, and flexibility. The addition of borophene nanosheets enables the borophene/ecoflex nanocomposite to exhibit tunable surface triboelectricity investigated by Kelvin probe force microscopy. The borophene/ecoflex nanocomposite is further fabricated into a fabric-based triboelectric nanogenerator (B-TENG) for mechanical energy harvesting, medical assistive system, and wound healing applications. The durability of B-TENG provides consistent output performance even after severe deformation treatments, such as folding, stretching, twisting, and washing procedures. Moreover, the B-TENG is integrated into a smart keyboard configuration combined with a robotic system to perform an upper-limb medical assistive interface. Furthermore, the B-TENG is also applied as an active gait phase sensing system for instantaneous lower-limb gait phase visualization. Most importantly, the B-TENG can be regarded as a self-powered in vitro electrical stimulation device to conduct continuous wound monitoring and therapy. The as-designed B-TENG not only demonstrates great potential for multifunctional self-powered healthcare sensors, but also for the promising advancements toward wearable medical assistive and therapeutic systems. | 44774 | 2022
R2PIKEJS | 10.1002/adma.202108607 | https://onlinelibrary.wiley.com/doi/10.1002/adma.202108607 | YES | patch |  | minimally | real-time and continuous monitoring of sodium levels in interstitial fluid (ISF) | electrochemical | Electrochemical | Sensitivity: The sensor's sensitivity is reported as 5.61 mA mm⁻¹. This value represents the change in drain-source current per unit change in sodium concentration. A higher sensitivity value indicates that the sensor can detect smaller variations in sodium levels. Limit of Detection (LOD): The LOD is reported as 2.78 × 10⁻⁶ m. This value represents the lowest concentration of sodium that the sensor can reliably distinguish from the background noise. A lower LOD indicates better sensitivity and the ability to detect trace amounts of the analyte. |  | NA | - | - | yes | interstitial fluid |  |  |  |  | yes |  | yes |  |  |  | yes |  | Sodium, pH, glucose, cortisol |  |  |  |  | 1 | 4 | Prototype/Validation | Human | in-vitro tests + animal models + human studies | NA | NA | Stretchable Microneedle Array: combines rigid microneedles for interstitial fluid (ISF) access with a stretchable substrate for skin conformability, enabling minimally invasive, long-term biomarker sampling. Extended Gate Field-Effect Transistor (EGFET): separates the sensing element from the transistor for enhanced sensitivity, stability, and rapid electrical detection, crucial for real-time monitoring. Silver Nanowire (AgNW) Electrodes: provides flexibility, conductivity, and durability, maintaining performance under stretching and bending for wearable applications. Multiplexed Sensing Capability: allows detection of various biomarkers (e.g., sodium, pH, glucose, cortisol) via modifiable microneedle electrodes. Wireless and IoT Integration: incorporates Bluetooth for remote monitoring and IoT connectivity, enabling real-time health tracking and telehealth applications. | NO | NA | NA | NA | 2022 | Zheng, Youbin; Omar, Rawan; Zhang, Rongjun; Tang, Ning; Khatib, Muhammad; Xu, Qi; Milyutin, Yana; Saliba, Walaa; Broza, Yoav Y.; Wu, Weiwei; Yuan, Miaomiao; Haick, Hossam | A Wearable Microneedle‐Based Extended Gate Transistor for Real‐Time Detection of Sodium in Interstitial Fluids | Advanced Materials |  | 0935-9648, 1521-4095 | Abstract Sodium is a prominent prognostic biomarker for assessing health status, such as dysnatremia. As of now, detection and monitoring of sodium levels in the human body is carried out by means of laborious and bulky laboratory equipmentand/or by offline analysis of various body fluids. Herein, an innovative stretchable, skin‐conformal and fast‐response microneedle extended‐gate FET biosensor is reported for real‐time detection of sodium in interstitial fluids for minimally invasive health monitoring along with high sensitivity, low limit of detection, excellent biocompatibility, and on‐body mechanical stability. The integration of the reported device with a wireless‐data transmitter and the Internet‐of‐Things cloud for real‐time monitoring and long‐term analysis is reported and discussed. This platform would eventually help bringing unlimited possibilities for effecient medical care and accurate clinical decision‐making. | 44621 | 2022
T2ALK3AC | 10.1016/j.talanta.2023.124362 | https://linkinghub.elsevier.com/retrieve/pii/S0039914023001133 | YES | patch |  | NO | non-invasive monitoring of glucose and pH levels in human perspiration (sweat) | electrochemical | Electrochemical | Glucose Sensor: Sensitivity: The glucose sensor exhibited a high sensitivity of 69.64 μA mM− 1 cm− 2. This means that for every 1 millimolar (mM) change in glucose concentration, the sensor produces a current response of 69.64 microamperes (μA) per square centimeter (cm− 2) of the electrode's surface area. Limit of Detection (LOD): The LOD for glucose was calculated to be 0.23 μM (micromolar), signifying the sensor's ability to detect very low concentrations of glucose in sweat. pH Sensor: Sensitivity: The pH sensor demonstrated a sensitivity of − 72.4 mV/pH. This indicates that for every unit change in pH, the sensor's open circuit potential (OCP) shifts by 72.4 millivolts (mV). The negative sign implies an inverse relationship between pH and OCP. The source highlights these values as indicators of the device's excellent performance, showcasing its ability to detect both glucose and pH with high sensitivity and accuracy. The low LOD for glucose, in particular, suggests its potential for applications requiring precise measurements of even minute glucose fluctuations in sweat. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [5 empty] | yes |  | glucose, pH |  |  |  |  | 1 | 2 | Proof-of-concept | Human | laboratory testing + human studies | NA | NA | 3D Porous Laser-Scribed Graphene (LSG): a unique architecture with a high surface area and interconnected pores enhances analyte interaction, sensitivity, and response time. Platinum Nanoparticle Integration: simultaneous incorporation of PtNPs boosts conductivity, electrocatalytic activity, and electron transfer, enabling precise glucose and pH detection. Hydroxyethyl Cellulose (HEC) Enhancement: improves electrode hydrophilicity, ensuring effective sweat contact and better analyte detection. Flexible Substrate: PI film provides a comfortable, body-conforming design for reliable on-body measurements. Efficient Fabrication: the one-step laser method simplifies electrode production while combining advanced material properties. Non-Invasive Monitoring: offers a biocompatible, high-performance solution for real-time sweat analysis of physiological compounds. | NO | NA | NA | NA | 2023 | Wang, Ying; Guo, Haojie; Yuan, Min; Yu, Jiabing; Wang, Zeping; Chen, Xianping | One-step laser synthesis platinum nanostructured 3D porous graphene: A flexible dual-functional electrochemical biosensor for glucose and pH detection in human perspiration | Talanta |  | 399140 | There has been a recent increase in the demand for wearable sensors for sweat glucose monitoring to facilitate diabetes management in a patient-friendly and non-invasive manner. To address this issue, the key challenge lies in the design of flexible sensors with high conductivity, miniaturized patterning, and environmental friendliness. Herein, we introduce a flexible electrochemical sensing system for glucose and pH detection based on one-step laser-scribed PtNPs nanostructured 3D porous laser-scribed graphene (LSG). The as-prepared nanocomposites can synchronously possess hierarchical porous graphene architectures, whereas PtNPs can significantly enhance their sensitivity and electrocatalytic activity. Benefiting from these advantages, the fabricated Pt-HEC/LSG biosensor exhibited a high sensitivity of 69.64 μA mM 1 cm 2 as well as a low limit of detection (LOD) of 0.23 μM at a detection range of 5–3000 μM (covering the glucose range in sweat). Moreover, the pH sensor was functionalized with polyaniline (PANI) on a Pt-HEC/LSG electrode, and it also exhibited high sensitivity (72.4 mV/pH) in the linear range of pH 4–8. The feasibility of the biosensor was confirmed by analyzing human perspiration during physical exercise. This dual-functional electrochemical biosensor displayed excellent performance, including a low detection limit, high selectivity, and great flexibility. These results confirm that the proposed dual-functional flexible electrode and fabrication process are highly promising for application in human sweat-based electrochemical glucose and pH sensors. | 45047 | 2023
CE4VVVVC | 10.1007/s00216-020-02939-4 | https://www.ncbi.nlm.nih.gov/pubmed/32989512 | YES | skin-attached |  | NO | monitoring sweat | electrochemical | Electrochemical (voltammetry, potentiometry, impedance) | A potentiometric potassium-ion-selective electrode in combination with an electrodeposited Ag/AgCl reference electrode (RE) enabled the detection of potassium ions in the entire physiologically relevant range (1 to 500 mM) with a fast response time, unaffected by the presence of main interfering ions and sweat-collecting materials. Using glucose and lactate oxidases, lower limits of detection of 13.7 ± 0.5 μM for glucose and 28 ± 3 μM for lactate were obtained, respectively. The sensor showed a good performance at different pH, with sweat-collecting tissues, on a model skin system and furthermore in synthetic sweat as well as in artificial tear fluid. Response time for each analytical cycle totals 75 s, and hence allows a quasi-continuous and simultaneous monitoring of all analytes. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | glucose, lactate, K+, electrolytes |  |  |  |  | 1 | 4 | Prototype |  | In-vitro tests | NA | NA | Multi-analyte sensing using exclusively laser-induced graphene (LIG)-based planar electrode systems was developed for sweat analysis. LIG provides 3D structures of graphene, can be manufactured easier than any other carbon electrode also on large scale, and in form of electrodes: hence, it is predestinated for affordable, wearable point-of-care sensors. | NO | NA | NA | NA | 2021 | Bauer, M.; Wunderlich, L.; Weinzierl, F.; Lei, Y.; Duerkop, A.; Alshareef, H. N.; Baeumner, A. J. | Electrochemical multi-analyte point-of-care perspiration sensors using on-chip three-dimensional graphene electrodes | Anal Bioanal Chem |  | 1618-2650 (Electronic) 1618-2642 (Print) 1618-2642 (Linking) | Multi-analyte sensing using exclusively laser-induced graphene (LIG)-based planar electrode systems was developed for sweat analysis. LIG provides 3D structures of graphene, can be manufactured easier than any other carbon electrode also on large scale, and in form of electrodes: hence, it is predestinated for affordable, wearable point-of-care sensors. Here, it is demonstrated that LIG facilitates all three electrochemical sensing strategies (voltammetry, potentiometry, impedance) in a multi-analyte system for sweat analysis. A potentiometric potassium-ion-selective electrode in combination with an electrodeposited Ag/AgCl reference electrode (RE) enabled the detection of potassium ions in the entire physiologically relevant range (1 to 500 mM) with a fast response time, unaffected by the presence of main interfering ions and sweat-collecting materials. A kidney-shaped interdigitated LIG electrode enabled the determination of the overall electrolyte concentration by electrochemical impedance spectroscopy at a fixed frequency. Enzyme-based strategies with amperometric detection share a common RE and were realized with Prussian blue as electron mediator and biocompatible chitosan for enzyme immobilization and protection of the electrode. Using glucose and lactate oxidases, lower limits of detection of 13.7 +/- 0.5 muM for glucose and 28 +/- 3 muM for lactate were obtained, respectively. The sensor showed a good performance at different pH, with sweat-collecting tissues, on a model skin system and furthermore in synthetic sweat as well as in artificial tear fluid. Response time for each analytical cycle totals 75 s, and hence allows a quasi-continuous and simultaneous monitoring of all analytes. This multi-analyte all-LIG system is therefore a practical, versatile, and most simple strategy for point-of-care applications and has the potential to outcompete standard screen-printed electrodes. Graphical abstract. | 44197 | 2021
PI3SG7T3 | 10.1016/j.aca.2022.339843 | https://linkinghub.elsevier.com/retrieve/pii/S0003267022004147 | YES | patch |  | NO | real-time detection and monitoring of uric acid (UA) levels in human sweat | electrochemical | Electrochemical | The wearable sweat UA sensor exhibits a high sensitivity for UA detection with a low detection limit of 0.34 μM and a linear range of 2 μM to 70 μM. The researchers validated the accuracy of the sensor by comparing its measurements to those obtained using a commercially available colorimetric assay kit. The results showed a strong correlation, with a Pearson correlation coefficient of 0.9897. | Metabolic Syndrome | management of gout, CVD and kidney diseases | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | Uric acid (UA) |  |  |  |  | 1 | 1 | Prototype/Validation | Human | In-vitro tests + human studies | NA | NA | Hydrophilic MOFs (MAF-7): protects uricase enzyme from degradation and enhances sensitivity by creating channels for efficient UA transport. Flexible Microfluidic Chip: enables efficient sweat collection while reducing errors from evaporation or contamination. Flexible Polyimide Substrate: provides conformal wearability, ensuring comfort, reliability, and durability under mechanical stress. Wireless Communication: integrates Bluetooth for real-time monitoring and data logging on smartphones or computers. | NO | NA | NA | NA | 2022 | Xiao, Jingyu; Luo, Yong; Su, Lei; Lu, Jiafa; Han, Wei; Xu, Tailin; Zhang, Xueji | Hydrophilic metal-organic frameworks integrated uricase for wearable detection of sweat uric acid | Analytica Chimica Acta |  | 32670 | The chemistry of the metal-organic frameworks (MOFs) coating may affect the biological functionality of the encapsulated biomacromolecules in harsh environment. Enzymes encapsulated in hydrophilic MAF-7 can retain high activity in harsh environment. We conducted this study to prepare a non-invasive wearable uircase@MAF7-based electrochemical sensor that can achieve accurate and sensitive detection of UA levels in sweat by integrating a flexible microfluidic chip and wireless electronic readout device. The flexible microfluidic chip enabled an easy and effective collection of sweat samples. MAF-7 protected enzyme activity by encapsulating uricase. The uricase@MAF-7-based electrochemical sensor enabled the highly sensitive detection of UA in the concentration range of 2 μM–70 μM with a detection limit of as low as 0.34 μM. Additionally, we evaluated the utility of the sensor for monitoring UA levels in real sweat samples by means of a high purine dietary challenge. This personalized wearable sweat sensing device has a potential to be used for monitoring disease-related metabolites in daily life. | 44682 | 2022
2UXW4UX5 | 10.1021/acssensors.3c00708 | https://www.ncbi.nlm.nih.gov/pubmed/37289663 | YES | skin-attached |  | NO | Continuous sweat lactate monitoring in sports | electrochemical | Electrochemical | The sensor covering the range expected for lactate in sweat (1-20 mM), with appropriate sensitivity (-12.5 ± 0.53 nA mM-1), shows an acceptable response time (<90 s), and the influence of changes in pH, temperature, and flow rate are neglectable. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate |  |  |  |  | 1 | 1 | Validation | Human | Human studies | NA | NA | The novelty of the system is threefold: advanced microfluidics design for sweat collection and analysis, an analytically validated lactate biosensor based on a rational design of an outer diffusion-limiting membrane, and an integrated circuit for signal processing with a custom smartphone application. | NO | NA | YES | NA | 2023 | Xuan, X.; Chen, C.; Molinero-Fernandez, A.; Ekelund, E.; Cardinale, D.; Swaren, M.; Wedholm, L.; Cuartero, M.; Crespo, G. A. | Fully Integrated Wearable Device for Continuous Sweat Lactate Monitoring in Sports | ACS Sens |  | 2379-3694 (Electronic) 2379-3694 (Linking) | The chemical digitalization of sweat using wearable sensing interfaces is an attractive alternative to traditional blood-based protocols in sports. Although sweat lactate has been claimed to be a relevant biomarker in sports, an analytically validated wearable system to prove that has not yet been developed. We present a fully integrated sweat lactate sensing system applicable to in situ perspiration analysis. The device can be conveniently worn in the skin to monitor real-time sweat lactate during sports, such as cycling and kayaking. The novelty of the system is threefold: advanced microfluidics design for sweat collection and analysis, an analytically validated lactate biosensor based on a rational design of an outer diffusion-limiting membrane, and an integrated circuit for signal processing with a custom smartphone application. The sensor covering the range expected for lactate in sweat (1-20 mM), with appropriate sensitivity (-12.5 +/- 0.53 nA mM(-1)), shows an acceptable response time (<90 s), and the influence of changes in pH, temperature, and flow rate are neglectable. Also, the sensor is analytically suitable with regard to reversibility, resilience, and reproducibility. The sensing device is validated through a relatively high number of on-body tests performed with elite athletes cycling and kayaking in controlled environments. Correlation outcomes between sweat lactate and other physiological indicators typically accessible in sports laboratories (blood lactate, perceived exhaustion, heart rate, blood glucose, respiratory quotient) are also presented and discussed in relation to the sport performance monitoring capability of continuous sweat lactate. | 45100 | 2023
TFJLJMC5 | 10.1016/j.bios.2024.116697 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324007036 | YES | patch |  | minimally | real-time, continuous monitoring of fitness-related biomarkers in interstitial fluid | electrochemical | Electrochemical | Good Correlation with Standards: on-body trials show "reliable performance" when compared to commercial glucose and lactate blood strips and a breathalyzer, suggesting reasonable accuracy. Linear Response Characteristics: in vitro tests reveal a consistent linear correlation between sensor-generated electrical current and biomarker concentrations (glucose, lactate, alcohol), indicating responsiveness and predictability. Reproducibility: experiments confirm high reproducibility, with consistent electrochemical performance and chronoamperometry curves across different microneedle arrays and repeated measurements, enhancing reliability. |  | NA | - | - | yes | interstitial fluid |  |  |  |  | yes | yes | [6 empty] | Glucose, Lactate, Alcohol |  |  |  |  | 1 | 3 | Validation | Human | in-vitro tests + animal models + human studies | NA | NA | The key technological innovation in the wearable microneedle sensor is the integration of 3D-printed microneedle arrays with electrochemical sensing technology. This combination allows for continuous, real-time monitoring of multiple biomarkers in interstitial fluid. | NO | NA | NA | NA | 2024 | Zhong, Geng; Liu, Qingzhou; Wang, Qiyu; Qiu, Haoji; Li, Hanlin; Xu, Tailin | Fully integrated microneedle biosensor array for wearable multiplexed fitness biomarkers monitoring | Biosensors and Bioelectronics |  | 9565663 | Fitness monitoring has become increasingly important in modern lifestyles; the current fitness monitoring always relies on physical sensors, making it challenging to detect pertinent issues at a deeper level when exercising. Here, we report a fully integrated wearable microneedle sensor that simultaneously measures fitness related biomarkers (e.g., glucose, lactate, and alcohol) during physical exercise. Such a sensor integrates a biocompatible 3D-printed microneedle array that can comfortably access skin interstitial fluid and a small circuit for signal processing and calibration, and wireless communication. The microneedle array features good biocompatibility and highly sensitive biochemical sensors that can detect even the slightest variations within the biomarkers of this fluid. On-body experimental results indicate that such a sensor can monitor fitness-related biomarkers across multiple subjects and support multi-day monitoring, with results showing a good correlation with commercial devices. The data was transmitted to a smartphone via Bluetooth and uploaded to cloud platforms for further health assessment. This study has the potential to boost intelligent wearable devices in sports health. | 45627 | 2024
INWQM67K | 10.3390/s19122713 | https://www.ncbi.nlm.nih.gov/pubmed/31212895 | YES | skin-attached |  | NO | human stress monitoring | electrochemical | Electrochemical | All ISFETs demonstrate a linear response to nitrate activity within the concentration range of 10^(−4) M to 10^(−1) M. The concentration range of nitrate in saliva is approximately 10^(−4) M to 10^(−3) M. Consequently, NO3-ISFETs are well-suited for nitrate measurement in saliva. |  | NA | - | yes | - | saliva | [12 empty] | nitrate |  |  |  |  | 1 | 1 | prototype |  | Laboratory testing | NA | NA | Polyurethane matrix membranes for the salivary nitrate ISFETs. The research focuses on investigating specific polyurethane materials like KP-13, Pellethane®, and P7281-PU to mitigate the initial drift commonly observed in ISFET measurements. | NO | NA | NA | NA | 2019 | Osaki, S.; Kintoki, T.; Moriuchi-Kawakami, T.; Kitamura, K.; Wakida, S. I. | Investigation of Polyurethane Matrix Membranes for Salivary Nitrate ISFETs to Prevent the Drift | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | We have investigated human-stress monitoring by making use of salivary nitrate, which can be a candidate for stress markers, with ion-selective field-effect transistors (ISFETs). ISFETs are suitable for on-site single-drop analysis of salivary nitrate within 10 s. However, when ISFETs are used for salivary nitrate, ISFETs have a problem that is called the initial drift. The initial drift makes accurate nitrate monitoring difficult. Thus, the purpose of this study is to prevent the initial drift and to search for a new, simple polymer to possess a better performance of sensor responses than conventional matrix membranes, such as PVC. In this research, we investigated ISFETs using specific matrix membranes, for example KP-13, Pellethane((R)-)(-), and P7281-PU. The initial drift was evaluated from the fluctuations of the response values generated by the ISFETs when immersed in saliva or aqueous solution. As a result, P7281-PU showed a prevention effect on the initial drift, both in the whole saliva and in various solutions. Furthermore, the cause of drift may be H(+) diffusion, and the drift prevention effect of P7281-PU may be affected by urethane bond capturing H(+) in the ion-selective membrane. This result suggests that a continuous nitrate monitoring is feasible and may be applied to wearable sensors. | 43633 | 2019
LQ7BWVJL | 10.3390/bios13070684 | https://www.ncbi.nlm.nih.gov/pubmed/37504083 | YES | textile |  | NO | noninvasive monitoring of glucose levels in sweat | electrochemical | Electrochemical | The screen-printed textile-based glucose biosensor achieved a linear response in the range of 20-1000 µM of glucose concentration and high sensitivity (18.41 µA mM-1 cm-2, R2 = 0.996). In addition, the biosensors show high selectivity toward glucose among other interfering analytes and excellent stability over 30 days of storage. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation | Human | In-vitro + in-vivo (human) | NA | NA | A screen-printed electrochemical sensor on a textile substrate for monitoring glucose levels in sweat. The fabrication method is cost-effective and can be easily integrated with everyday clothing using existing clothing manufacturing processes at an industrial scale. The sensors are fabricated layer-by-layer with a screen-printing method that enables the fabrication of customized electrode geometries and configurations. | NO | NA | NA | NA | 2023 | Khosravi, S.; Soltanian, S.; Servati, A.; Khademhosseini, A.; Zhu, Y.; Servati, P. | Screen-Printed Textile-Based Electrochemical Biosensor for Noninvasive Monitoring of Glucose in Sweat | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | Wearable sweat biosensors for noninvasive monitoring of health parameters have attracted significant attention. Having these biosensors embedded in textile substrates can provide a convenient experience due to their soft and flexible nature that conforms to the skin, creating good contact for long-term use. These biosensors can be easily integrated with everyday clothing by using textile fabrication processes to enhance affordable and scalable manufacturing. Herein, a flexible electrochemical glucose sensor that can be screen-printed onto a textile substrate has been demonstrated. The screen-printed textile-based glucose biosensor achieved a linear response in the range of 20-1000 microM of glucose concentration and high sensitivity (18.41 microA mM(-1) cm(-2), R(2) = 0.996). In addition, the biosensors show high selectivity toward glucose among other interfering analytes and excellent stability over 30 days of storage. The developed textile-based biosensor can serve as a platform for monitoring bio analytes in sweat, and it is expected to impact the next generation of wearable devices. | 45104 | 2023
4KILEFWL | 10.1016/j.bios.2022.114955 | https://linkinghub.elsevier.com/retrieve/pii/S0956566322009952 | YES | not specified | microneedle array | minimally | transdermal monitoring of pH levels in interstitial fluid (ISF). | electrochemical | Electrochemical | The wearable pH sensor operates from pH 4.0 to 8.6 with a sensitivity of 62.9 mV per pH unit and an accuracy of ±0.036 pH units. |  | NA | - | - | yes | interstitial fluid | [10 empty] | yes |  | pH |  |  |  |  | 1 | 1 | Proof-of-concept |  | In-vitro, laboratory testing, ex-vivo (animal models) | NA | NA | The High-Density Polymeric Microneedle Array (PMNA) enables painless, minimally invasive access to interstitial fluid (ISF) through microscopically small needles, avoiding pain receptors and blood vessels. Fabricated using cost-effective soft lithography with PDMS molds, it simplifies production compared to silicon-based microneedles. Tailored to a height of ~250 μm, the microneedles ensure accurate pH monitoring by preventing blood contamination. An OrmoComp® insulating layer enhances selectivity by blocking sweat interference, while a polyaniline (PA) sensing layer ensures high-sensitivity pH detection. Integrated into a flexible PEN substrate, the wearable patch offers comfortable, continuous ISF monitoring with resilient performance during movement. | NO | NA | NA | NA | 2023 | Dervisevic, Muamer; Dervisevic, Esma; Esser, Lars; Easton, Christopher D.; Cadarso, Victor J.; Voelcker, Nicolas H. | Wearable microneedle array-based sensor for transdermal monitoring of pH levels in interstitial fluid | Biosensors and Bioelectronics |  | 9565663 | Microneedle-based wearable sensors offer an alternative approach to traditional invasive blood-based health monitoring and disease diagnostics techniques. Instead of blood, microneedle-based sensors target the skin interstitial fluid (ISF), in which the biomarker type and concentration profile resemble the one found in the blood. However, unlike blood, interstitial fluid does not have the same pH-buffering capacity causing deviation of pH levels from the physiological range. Information about the skin ISF pH levels can be used as a biomarker for a wide range of pathophysiological conditions and as a marker for the calibration of a wearable sensor. The ISF pH can significantly affect the detection accuracy of other biomarkers as it influences enzyme activity, aptamer affinity, and antibody-antigen interaction. Herein, we report the fabrication of a high-density polymeric microneedle array-based (PMNA) sensing patch and its optimization for the potentiometric transdermal monitoring of pH levels in ISF. The wearable sensor utilizes a polyaniline-coated PMNA having a density of ~10,000 microneedles per cm2, containing individual microneedles with a height of ~250 μm, and a tip diameter of ~2 μm. To prevent interference from other body fluids like sweat, an insulating layer is deposited at the base of the PMNA. The wearable pH sensor operates from pH 4.0 to 8.6 with a sensitivity of 62.9 mV per pH unit and an accuracy of ±0.036 pH units. Furthermore, testing on a mouse demonstrates the ability of the PMNA to provide a real-time reading of the transdermal pH values. This microneedle-based system will significantly contribute to advancing transdermal wearable sensors technology, simplifying the fabrication process, and improving the costeffectiveness of such devices. | 44958 | 2023
UGXD7F26 | 10.3390/bios9010013 | https://www.ncbi.nlm.nih.gov/pubmed/30654428 | YES | not specified |  | NO | simultaneous detection of multiple biomarkers in sweat | electrochemical | Electrochemical | The device can detect alcohol, glucose, and lactate within physiologically relevant ranges. For instance, it can measure alcohol levels from 0.1 to 100 mg/dL, glucose from 0.1 to 50 mg/dL, and lactate from 1 to 100 mM.The limit of detection for biomarker detection in human sweat was established to be one logarithmic concentration lower and the dynamic range was established to be one logarithmic concentration higher than the physiological relevant range of the biomarkers. |  | NA | yes | - | - | sweat |  |  |  |  | yes | yes | [6 empty] | alcohol, glucose, lactate |  |  |  |  | 1 | 3 | prototype |  | Laboratory testing | NA | NA | Non-faradaic biosensor. Nanotextured ZnO Films for Enhanced Sensing. | NO | NA | NA | NA | 2019 | Bhide, A.; Cheeran, S.; Muthukumar, S.; Prasad, S. | Enzymatic Low Volume Passive Sweat Based Assays for Multi-Biomarker Detection | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | Simultaneous detection of correlated multi-biomarkers on a single low-cost platform in ultra-low fluid volumes with robustness is in growing demand for the development of wearable diagnostics. A non-faradaic biosensor for the simultaneous detection of alcohol, glucose, and lactate utilizing low volumes (1(-)5 muL) of sweat is demonstrated. Biosensing is implemented using nanotextured ZnO films integrated on a flexible porous membrane to achieve enhanced sensor performance. The ZnO sensing region is functionalized with enzymes specific for the detection of alcohol, glucose, and lactate in the ranges encompassing their physiologically relevant levels. A non-faradaic chronoamperometry technique is used to measure the current changes associated with interactions of the target biomarkers with their specific enzyme. The specificity performance of the biosensing platform was established in the presence of cortisol as the non-specific molecule. Biosensing performance of the platform in a continuous mode performed over a 1.5-h duration showed a stable current response to cumulative lifestyle biomarker concentrations with capability to distinguish reliably between low, mid, and high concentration ranges of alcohol (0.1, 25, 100 mg/dL), glucose (0.1, 10, 50 mg/dL), and lactate (1, 50, 100 mM). The low detection limits and a broader dynamic range for the lifestyle biomarker detection are quantified in this research demonstrating its suitability for translation into a wearable device. | 43481 | 2019
RRQHZWAL | 10.1016/j.bios.2023.115360 | https://linkinghub.elsevier.com/retrieve/pii/S0956566323003020 | YES | skin-attached |  | NO | simultaneous and accurate monitoring of pH and tyrosine (Tyr) levels in sweat | electrochemical | Electrochemical | High pH Sensitivity: demonstrates a sensitivity of -71.86 mV/pH with excellent linearity (R² = 0.9975), ensuring precise pH detection. Reproducible pH Measurements: achieves consistent results with a low relative standard deviation (RSD) of 1.6% across multiple sensors. Low Tyrosine Detection Limit: detects tyrosine at concentrations as low as 3.3 μM, ensuring high sensitivity. Broad Tyrosine Linear Range: provides accurate quantification of tyrosine levels between 10 μM and 200 μM (R² = 0.9985). Integrated pH Calibration: uses pH readings to calibrate tyrosine measurements, ensuring accuracy comparable to commercial ELISA kits. |  | NA | yes | - | - | sweat | [10 empty] | yes |  | Tyrosine, pH |  |  |  |  | 1 | 2 | Prototype/Validation | Human | in-vitro tests + laboratory testing + human studies (preliminary) | NA | NA | Multifunctional Conductive Hydrogel (TA-Ag-CNT-PANI): combines antibacterial silver nanoparticles (TA-Ag), conductive and flexible carbon nanotubes (CNTs), and pH/tyrosine-responsive polyaniline (PANI) for precise sensing. Integrated Iontophoresis System: enables controlled sweat stimulation, bypassing reliance on external factors. pH-Calibrated Tyrosine Detection: simultaneous pH monitoring enhances electrochemical accuracy. Screen-Printed Electrode Array: Facilitates cost-effective, scalable production for commercialization. | NO | NA | NA | NA | 2023 | Xu, Zhenying; Qiao, Xiujuan; Tao, Runzhang; Li, Yanxin; Zhao, Shuju; Cai, Yuchen; Luo, Xiliang | A wearable sensor based on multifunctional conductive hydrogel for simultaneous accurate pH and tyrosine monitoring in sweat | Biosensors and Bioelectronics |  | 9565663 | Flexible and wearable sweat sensors have drawn extensive attention by virtue of their continuous and real-time monitoring of molecular level information. However, current sweat-based sensors still pose several challenges, such as low accuracy for analytes detection, susceptibility to microorganism and poor mechanical performance. Herein, we demonstrated a noninvasive wearable sweat sensing patch composed of an electrochemical sensing system, and a pilocarpine-based iontophoretic system to stimulate sweat secretion. The electrochemical sensor based on tannic acid-Ag-carbon nanotube-polyaniline (TA-Ag-CNT-PANI) composite hydrogel was designed for on-body detection of pH and tyrosine (Tyr), a disease marker associated with multiple disorders, such as tyrosinemia and bulimia nervosa. The wearable sweat sensor can not only monitor the pH and Tyr in sweat simultaneously, but also further calibrate Tyr detection results with the measured pH value, so as to eliminate the effect of Tyr response variance at different pH and enhance the accuracy of the sensor. Furthermore, the presence of tannic acid chelated-Ag nanoparticles (TA-Ag NPs) and carbon nanotubes (CNTs) significantly improved the conductivity and flexibility of the hydrogel and endowed the composite hydrogel with antibacterial capability. Of note, the constructed wearable sensor was capable of monitoring Tyr with enhanced accuracy in various sweats. | 45139 | 2023
7BMN76BK | 10.3390/bios13020157 | https://www.ncbi.nlm.nih.gov/pubmed/36831923 | YES | chip | microfluidic | NO | exercise training monitoring | Colorimetric | Colorimetric | Under optimized conditions, this proposed chip can detect glucose with low concentrations from 0.05 mM to 0.40 mM, with a pH range of 4.0 to 6.5 for human sweat. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [5 empty] | yes |  | glucose, pH |  |  |  |  | 1 | 2 | Prototype/Validation | Human | Human studies | NA | NA | a wearable microfluidic sweat chip combined with smartphone image processing to realize non-invasive in situ analysis of epidermal sweat for sports practitioners. The polydimethylsiloxane (PDMS) based chip was modified with nonionic surfactants to ensure good hydrophilicity for the automatic collection of sweat. | NO | NA | YES | NA | 2023 | Liu, D.; Liu, Z.; Feng, S.; Gao, Z.; Chen, R.; Cai, G.; Bian, S. | Wearable Microfluidic Sweat Chip for Detection of Sweat Glucose and pH in Long-Distance Running Exercise | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | Traditional exercise training monitoring is based on invasive blood testing methods. As sweat can reveal abundant blood-related physiological information about health, wearable sweat sensors have received significant research attention and become increasingly popular in the field of exercise training monitoring. However, most of these sensors are used to measure physical indicators such as heart rate, blood pressure, respiration, etc., demanding a versatile sensor that can detect relevant biochemical indicators in body fluids. In this work, we proposed a wearable microfluidic sweat chip combined with smartphone image processing to realize non-invasive in situ analysis of epidermal sweat for sports practitioners. The polydimethylsiloxane (PDMS) based chip was modified with nonionic surfactants to ensure good hydrophilicity for the automatic collection of sweat. Besides, a simple, reliable, and low-cost paper-based sensor was prepared for high-performance sensing of glucose concentration and pH in sweat. Under optimized conditions, this proposed chip can detect glucose with low concentrations from 0.05 mM to 0.40 mM, with a pH range of 4.0 to 6.5 for human sweat. The ability of this microfluidic chip for human sweat analysis was demonstrated by dynamically tracking the changes in glucose concentration and pH in long-distance running subjects. | 44945 | 2023
6UL3HGJQ | 10.3390/molecules28197000 | https://www.ncbi.nlm.nih.gov/pubmed/37836843 | YES | not specified |  | NO | selective quantification of potassium ion (K+) levels in human sweat | electrochemical | Electrochemical | The sensor exhibits a linear range from 10 μM to 100 mM for potassium ion detection, with a sensitivity of 54.7 mV/decade. The sensor demonstrates good reproducibility over 10 cycles in 10 mM KCl. The sensor exhibits minimal interference from 100 mM Na+ and other common sweat constituents. |  | NA | yes | - | - | sweat | [12 empty] | K+ |  |  |  |  | 1 | 1 | Validation | Human | Human studies | NA | NA | Printed electrodes using β-cyclodextrin functionalized reduced graphene oxide (β-CD-RGO) for selective K+ quantification in sweat. The β-CD prevents the aggregation of RGO sheets while also providing selective binding sites for K+ capture. Electrodes were fabricated by screen printing the β-CD-RGO ink onto conductive carbon substrates. | NO | NA | NA | NA | 2023 | Liu, R.; Shi, X. | Preparation of beta-Cyclodextrin Functionalized Platform for Monitoring Changes in Potassium Content in Perspiration | Molecules |  | 1420-3049 (Electronic) 1420-3049 (Linking) | The monitoring of potassium ion (K(+)) levels in human sweat can provide valuable insights into electrolyte balance and muscle fatigue non-invasively. However, existing laboratory techniques for sweat testing are complex, while wearable sensors face limitations like drift, fouling and interference from ions such as Na(+). This work develops printed electrodes using beta-cyclodextrin functionalized reduced graphene oxide (beta-CD-RGO) for selective K(+) quantification in sweat. The beta-CD prevents the aggregation of RGO sheets while also providing selective binding sites for K(+) capture. Electrodes were fabricated by screen printing the beta-CD-RGO ink onto conductive carbon substrates. Material characterization confirmed the successful functionalization of RGO with beta-CD. Cyclic voltammetry (CV) showed enhanced electrochemical behavior for beta-CD-RGO-printed electrodes compared with bare carbon and RGO. Sensor optimization resulted in a formulation with 30% beta-CD-RGO loading. The printed electrodes were drop-casted with an ion-selective polyvinyl chloride (PVC) membrane. A linear range from 10 muM to 100 mM was obtained along with a sensitivity of 54.7 mV/decade. The sensor showed good reproducibility over 10 cycles in 10 mM KCl. Minimal interference from 100 mM Na(+) and other common sweat constituents validated the sensor's selectivity. On-body trials were performed by mounting the printed electrodes on human subjects during exercise. The K(+) levels measured in sweat were found to correlate well with serum analysis, demonstrating the sensor's ability for non-invasive electrolyte monitoring. Overall, the facile synthesis of stable beta-CD-RGO inks enables the scalable fabrication of wearable sensors for sweat potassium detection. | 45208 | 2023
XL9WGNP4 | 10.1109/TBME.2017.2764507 | https://www.ncbi.nlm.nih.gov/pubmed/29989933 | YES | not specified |  | NO | quantify stress level through ECG, EDA, and EEG analysis (in comparison with saliva cortisol levels). | Electrical | Electrophysiology | The classification algorithm, based on significant features, provided satisfactory outcomes with 86% accuracy. Furthermore, correlation analysis showed that the observed changes in physiological features were consistent with the trend of salivary cortisol levels (R 2 = 0.714). |  | NA | - | yes | - | saliva | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | Validation | Human | Human studies | NA | NA | Wearable physiological sensors system, based on ECG, EDA, and EEG, to capture human stress | NO | NA | YES | NA | 2018 | Betti, S.; Lova, R. M.; Rovini, E.; Acerbi, G.; Santarelli, L.; Cabiati, M.; Del Ry, S.; Cavallo, F. | Evaluation of an Integrated System of Wearable Physiological Sensors for Stress Monitoring in Working Environments by Using Biological Markers | IEEE Trans Biomed Eng |  | 1558-2531 (Electronic) 0018-9294 (Linking) | OBJECTIVE: The objectives of this paper are to develop and test the ability of a wearable physiological sensors system, based on ECG, EDA, and EEG, to capture human stress and to assess whether the detected changes in physiological signals correlate with changes in salivary cortisol level, which is a reliable, objective biomarker of stress. METHODS: 15 healthy participants, eight males and seven females, mean age 40.8 +/- 9.5 years, wore a set of three commercial sensors to record physiological signals during the Maastricht Acute Stress Test, an experimental protocol known to elicit robust physical and mental stress in humans. Salivary samples were collected throughout the different phases of the test. Statistical analysis was performed using a support vector machine (SVM) classification algorithm. A correlation analysis between extracted physiological features and salivary cortisol levels was also performed. RESULTS: 15 features extracted from heart rate variability, electrodermal, and electroencephalography signals showed a high degree of significance in disentangling stress from a relaxed state. The classification algorithm, based on significant features, provided satisfactory outcomes with 86% accuracy. Furthermore, correlation analysis showed that the observed changes in physiological features were consistent with the trend of salivary cortisol levels (R(2) = 0.714). CONCLUSION: The tested set of wearable sensors was able to successfully capture human stress and quantify stress level. SIGNIFICANCE: The results of this pilot study may be useful in designing portable and remote control systems, such as medical devices used to turn on interventions and prevent stress consequences. | 43313 | 2018
9ZAPL828 | 10.1016/j.talanta.2020.121484 | https://linkinghub.elsevier.com/retrieve/pii/S003991402030775X | YES | textile |  | NO | real-time, non-invasive lactate monitoring in sweat | electrochemical | Electrochemical | The wearable textile electrochemical biosensor exhibits high sensitivity for lactate detection, achieving a sensitivity of 19.13 μA/mM cm2 in a phosphate-buffered solution (PBS) and 14.6 μA/mM cm2 in artificial sweat. This sensitivity remains consistent even when the device is stretched, showcasing its robustness and reliability for real-world use5. Here's a further breakdown of the sensor's sensitivity: PBS: The sensor demonstrated a higher sensitivity in the low lactate concentration range of 0-5 mM (R2=0.927) with a sensitivity of 19.13 μA/mM cm2. In the higher concentration range of 10-30 mM (R2=0.91), the sensitivity was 2.9 μA/mM cm2. Artificial Sweat: In the artificial sweat solution, the sensitivity was slightly lower but still notable. For the 0-5 mM range (R2=0.986), the sensitivity was 14.6 μA/mM cm2. For the broader range of 0-30 mM (R2=0.931), it was 2.5 μA/mM cm2. The lower sensitivity in artificial sweat compared to PBS is likely attributed to the presence of various interfering chemical species in the artificial sweat, which could potentially affect the sensor's performance. Nevertheless, the device demonstrates a high sensitivity for lactate detection in both controlled and simulated real-world conditions. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate |  |  |  |  | 1 | 1 | Proof-of-concept | Human | laboratory testing + in-vitro tests + human study (preliminary) | NA | NA | Stretchable Gold Fiber Electrodes: highly conductive, strain-insensitive gold fibers, created via dry-spinning, integrate seamlessly into textiles for enhanced wearability. Wrinkled Gold Film Coating: increased electrochemically active surface area boosts enzyme immobilization and lactate detection sensitivity. Flexible Counter Electrode: elastomeric gold fibers replace rigid materials, enhancing flexibility and user comfort. Prussian Blue Layer: acts as both artificial peroxidase and electron mediator, improving sensitivity and efficient hydrogen peroxide decomposition. Chitosan Enzyme Immobilization: biocompatible chitosan stabilizes lactate oxidase, ensuring reliable, long-term sensor performance. Textile Integration: a fully woven three-electrode system enables continuous, non-invasive lactate monitoring during everyday activities. | NO | NA | NO | NA | 2021 | Wang, Ren; Zhai, Qingfeng; An, Tiance; Gong, Shu; Cheng, Wenlong | Stretchable gold fiber-based wearable textile electrochemical biosensor for lactate monitoring in sweat | Talanta |  | 399140 | Past several years have witnessed growing interest in developing wearable biosensors for non-invasive monitoring vital signs of chemical/biological markers such as lactate. In this context, textiles can be seen as a promising platform for the integration of wearable chemical sensors due to their inherent breathability, flexibility, softness and comfortableness. Gold is regarded as a preferred active sensing material due to its excellent biocompatibility, chemical inertness and wide electrochemical window. Here, a dry-spinning method was used to fabricate stretchable, strain-insensitive and highly conductive gold fibers. Such gold fibers could be used to fabricate lactate-sensing working electrodes, reference electrode, counter electrodes and further weaved into textiles in a standard three-electrode system with a planar layout. The textile lactate biosensors showed a high sensitivity of 19.13 μA/mM cm2 in phosphate-buffered solution (PBS) and 14.6 μA/mM cm2 in artificial sweat. This sensitivity could be maintained under high tensile strain up to 100% without external structural design. The results presented here indicate the potential application of wearable smart textile towards non-invasive lactate monitoring. | 44197 | 2021
HCPMFGBN | 10.1038/s41551-023-01059-5 | https://www.ncbi.nlm.nih.gov/pubmed/37349389 | YES | patch | wireless | NO | monitoring C-reactive protein (CRP) in sweat | electrochemical | Electrochemical | This technology realizes highly sensitive detection of ultra-low-level inflammatory proteins in situ with a 6 orders-of-magnitude (picomolar level) improvement in sensitivity | Inflammation | Inflammatory conditions (particularly chronic diseases) management | yes | - | - | sweat | [12 empty] | C-reactive protein |  |  |  |  | 1 | 1 | Validation | Human | Human studies | NA | NA | The device integrates iontophoretic sweat extraction, microfluidic channels for sweat sampling and for reagent routing and replacement, and a graphene-based sensor array for quantifying CRP (via an electrode functionalized with anti-CRP capture antibodies-conjugated gold nanoparticles), ionic strength, pH and temperature for the real-time calibration of the CRP sensor. | NO | NA | NA | NA | 2023 | Tu, J.; Min, J.; Song, Y.; Xu, C.; Li, J.; Moore, J.; Hanson, J.; Hu, E.; Parimon, T.; Wang, T. Y.; Davoodi, E.; Chou, T. F.; Chen, P.; Hsu, J. J.; Rossiter, H. B.; Gao, W. | A wireless patch for the monitoring of C-reactive protein in sweat | Nat Biomed Eng |  | 2157-846X (Electronic) 2157-846X (Linking) | The quantification of protein biomarkers in blood at picomolar-level sensitivity requires labour-intensive incubation and washing steps. Sensing proteins in sweat, which would allow for point-of-care monitoring, is hindered by the typically large interpersonal and intrapersonal variations in its composition. Here we report the design and performance of a wearable and wireless patch for the real-time electrochemical detection of the inflammatory biomarker C-reactive (CRP) protein in sweat. The device integrates iontophoretic sweat extraction, microfluidic channels for sweat sampling and for reagent routing and replacement, and a graphene-based sensor array for quantifying CRP (via an electrode functionalized with anti-CRP capture antibodies-conjugated gold nanoparticles), ionic strength, pH and temperature for the real-time calibration of the CRP sensor. In patients with chronic obstructive pulmonary disease, with active or past infections or who had heart failure, the elevated concentrations of CRP measured via the patch correlated well with the protein's levels in serum. Wearable biosensors for the real-time sensitive analysis of inflammatory proteins in sweat may facilitate the management of chronic diseases. | 45200 | 2023
U2NIYAGN | 10.1177/2472630319882003 | https://www.ncbi.nlm.nih.gov/pubmed/31617455 | YES | not specified |  | NO | analyzing biomarkers in sweat | electrochemical | Electrochemical (non-faradaic electrochemical impedance spectroscopy (EIS)). | A strong linear relationship between the sensor response and the analyte concentration for both biomarkers. This linearity is reflected in the high R-squared values (R² > 0.95) reported for both cortisol (5-200 ng/mL) and chloride (10-100 mM) detection in both types of samples. The researchers tested the device using pooled human sweat samples, demonstrating its applicability in a real-world biological fluid. The fact that the device achieved an R² > 0.95 in human sweat further supports its reliability and suggests that the device can accurately measure these biomarkers in a complex biological matrix. |  | NA | yes | - | - | sweat | [6 empty] | yes | [5 empty] | Chloride, cortisol |  |  |  |  | 1 | 2 | Validation |  | In-vitro tests + laboratory testing | NA | NA | Label-free affinity biosensing is achieved by using a two-electrode electrochemical system on a flexible substrate suitable for wearable applications. The electrochemical stability of the fabricated electrodes for biosensing applications was studied by open-circuit potential measurements. | EnLiSense LLC | USA | NA | NA | 2020 | Ganguly, A.; Rice, P.; Lin, K. C.; Muthukumar, S.; Prasad, S. | A Combinatorial Electrochemical Biosensor for Sweat Biomarker Benchmarking | SLAS Technol |  | 2472-6311 (Electronic) 2472-6303 (Linking) | Misclassification of an acute disease condition as chronic and vice versa by electrochemical sweat biomarker sensors can cause significant psychological, emotional, and financial stress among patients. To achieve higher accuracy in distinguishing between a chronic condition and an acute condition, there is a need to establish a reference biomarker to index the actual chronic disease biomarker of interest by combinatorial sensing. This work provides the first technological proof of leveraging the chloride ion content in sweat for a combinatorial sweat biomarker benchmarking scheme. In this scheme, the sweat chloride ion has been demonstrated as the reference/indexing biomarker, while sweat cortisol has been studied as the disease biomarker of interest. Label-free affinity biosensing is achieved by using a two-electrode electrochemical system on a flexible substrate suitable for wearable applications. The electrochemical stability of the fabricated electrodes for biosensing applications was studied by open-circuit potential measurements. Attenuated total reflectance-Fourier transform infrared spectroscopy spectra validate the crosslinker-antibody binding chemistry. Concentration-dependent analyte-capture probe binding induces a modulation in the electrical properties (charge transfer resistance and double-layer capacitance) at the electrode-sweat buffer interface, which are transduced by nonfaradaic electrochemical impedance spectroscopy (EIS). Calibration dose responses for the sensor for cortisol (5-200 ng/mL) and chloride (10-100 mM) detection were evaluated in synthetic (pH 6) and pooled human sweat (R(2) > 0.95). The variation in the cortisol sensor response due to fluctuations in sweat chloride levels and the significance of reporting normalized biomarker levels were demonstrated to further emphasize the need for biomarker benchmarking in electrochemical sensors. | 43862 | 2020
EZSP36YR | 10.1109/TBCAS.2023.3286528 | https://www.ncbi.nlm.nih.gov/pubmed/37318976 | YES | patch |  | NO | monitoring of sweat biomarkers during exercise | electrochemical | Electrochemical | Analysis of the sweat biomarkers collected from the printed patch prototype shows that their real-time measurements correlate well (correlation coefficient ≥ 0.65) with other physiological biomarkers such as heart rate and regional sweat rate collected in the same session. We show for the first time, that the real-time sweat sodium and potassium concentration biomarker measurements from the printed sensors can be used to predict the core body temperature with root mean square error (RMSE) of 0.02 °C which is 71% lower compared to the use of only the physiological biomarkers. |  | NA | yes | - | - | sweat | [12 empty] | Na+, K+ |  |  |  |  | 1 | 2 | prototype | Human | Human studies | NA | NA | A wearable sweat biomonitoring patch incorporating printed electrochemical sensors into a plastic microfluidic sweat collector and data analysis that shows the real-time recorded sweat biomarkers can be used to predict a physiological biomarker. The system was placed on subjects carrying out an hour-long exercise session and results were compared to a wearable system using potentiometric robust silicon-based sensors and to commercially available HORIBA-LAQUAtwin devices. | NO | NA | NA | NA | 2023 | Wang, S.; Rovira, M.; Demuru, S.; Lafaye, C.; Kim, J.; Kunnel, B. P.; Besson, C.; Fernandez-Sanchez, C.; Serra-Graells, F.; Margarit-Taule, J. M.; Aymerich, J.; Cuenca, J.; Kiselev, I.; Gremeaux, V.; Saubade, M.; Jimenez-Jorquera, C.; Briand, D.; Liu, S. C. | Multisensing Wearables for Real-Time Monitoring of Sweat Electrolyte Biomarkers During Exercise and Analysis on Their Correlation With Core Body Temperature | IEEE Trans Biomed Circuits Syst |  | 1940-9990 (Electronic) 1932-4545 (Linking) | Sweat secreted by the human eccrine sweat glands can provide valuable biomarker information during exercise. Real-time non-invasive biomarker recordings are therefore useful for evaluating the physiological conditions of an athlete such as their hydration status during endurance exercise. This work describes a wearable sweat biomonitoring patch incorporating printed electrochemical sensors into a plastic microfluidic sweat collector and data analysis that shows the real-time recorded sweat biomarkers can be used to predict a physiological biomarker. The system was placed on subjects carrying out an hour-long exercise session and results were compared to a wearable system using potentiometric robust silicon-based sensors and to commercially available HORIBA-LAQUAtwin devices. Both prototypes were applied to the real-time monitoring of sweat during cycling sessions and showed stable readings for around an hour. Analysis of the sweat biomarkers collected from the printed patch prototype shows that their real-time measurements correlate well (correlation coefficient >/= 0.65) with other physiological biomarkers such as heart rate and regional sweat rate collected in the same session. We show for the first time, that the real-time sweat sodium and potassium concentration biomarker measurements from the printed sensors can be used to predict the core body temperature with root mean square error (RMSE) of 0.02 degrees C which is 71% lower compared to the use of only the physiological biomarkers. These results show that these wearable patch technologies are promising for real-time portable sweat monitoring analytical platforms, especially for athletes performing endurance exercise. | 45139 | 2023
HN5YMJ9X | 10.1016/j.bios.2022.114005 | https://linkinghub.elsevier.com/retrieve/pii/S0956566322000458 | YES | skin-attached |  | NO | detection and quantification of copper excreted in sweat | electrochemical | Electrochemical detection + impedance-based measurements | Copper Sensor (SPCE) The copper sensor, based on a screen-printed carbon electrode (SPCE) and SWASV, demonstrated a limit of detection (LOD) of 396 ppb. Its linear range extends up to 2500 ppb. The sensor exhibits a sensitivity of 2.3 nA/ppb. A recovery test with 1500 ppb of copper yielded a calculated concentration of 1675 ppb, resulting in a recovery of 111.6%. The researchers acknowledge that the LOD is relatively high compared to other copper sweat sensors reported in the literature. They attribute this to the use of bare carbon paste, which has limited affinity for copper. Despite the higher LOD, the device's performance falls within the clinically relevant range for copper concentrations in sweat (240-2400 ppb). Conductivity Nanosensor (AgNP) The conductivity nanosensor, utilizing interdigitated AgNP electrodes, showed a clear response to varying NaCl concentrations (5 to 156 mM), which correspond to conductivities ranging from 0.60 to 17.03 mS/cm. The impedance measurements at 10 kHz revealed a distinct change from 100 kΩ (empty channel) to 30 kΩ when a 5 mM NaCl solution reached the electrodes. The measured admittance exhibited a perfect linear relationship with NaCl concentration: Admittance = 2.95 * 10^-6 * [NaCl] + 1.17 * 10^-5, R² = 0.99. Volume Nanosensor (AgNP) The volume nanosensor, employing parallel AgNP electrodes along the microfluidic channel, effectively measured changes in impedance corresponding to the volume of sweat present. The initial impedance of the empty microchannel (approximately 2*10⁶ Ω) dropped to less than 100 Ω when filled with artificial sweat, demonstrating a substantial response to volume changes. The volume nanosensor's admittance exhibited a linear relationship with the injected volume: ΔVolume = 0.00169 * Volume, R² = 0.99. Overall, the wearable microfluidic nanosensor demonstrates promising accuracy and sensitivity for monitoring copper levels, sweat conductivity, and sweat volume. |  | NA | yes | - | - | sweat | [12 empty] | Copper ions (+ conductivity) |  |  |  |  | 1 | 1 | Prototype |  | In-vitro tests (with artificial sweat) | NA | NA | All-Printed, Flexible Design: Incorporates inkjet-printed silver nanoparticle (AgNP) electrodes, screen-printed carbon electrodes, and laser-cut microfluidic channels for lightweight, body-conforming wearability. Integrated Multi-Sensor System: Features sensors for copper detection (SWASV), conductivity (electrolyte balance), and sweat volume (impedance changes), allowing detailed and real-time sweat analysis. Active Sweat Stimulation: Utilizes reverse iontophoresis with inkjet-printed electrodes and conductive hydrogels for controlled sweat production, overcoming passive collection limits. Sweat Rate Normalization: Real-time adjustment of biomarker concentration to sweat rate ensures accuracy despite physiological variations. Wearable Potentiostat and Smartphone Integration: A flexible potentiostat with Bluetooth connectivity streams data to a smartphone app for easy control and visualization. Non-Enzymatic Glucose Detection: Pioneers enzyme-free sensing, enhancing sensor stability, longevity, and safety. | NO | NA | YES | NA | 2022 | Yang, Qiuyue; Rosati, Giulio; Abarintos, Vernalyn; Aroca, Miguel Angel; Osma, Johann F.; Merkoçi, Arben | Wearable and fully printed microfluidic nanosensor for sweat rate, conductivity, and copper detection with healthcare applications | Biosensors and Bioelectronics |  | 9565663 | Wearables are becoming pervasive in our society, but they are still mainly based on physical sensors with just few optical and electrochemical exceptions. Sweat, amongst other body fluids, is easily and non-invasively accessible, abundant, and relatively poor of interfering species. The biomarkers of interest in sweat space from ions and small molecules to whole organisms. Heavy metals have been found being biomarkers of several diseases and pathological conditions. Copper in particular is correlated to Wilson’s disease and liver cirrhosis among others. Nevertheless, several issues such as sampling conditions, sweat rate normalization, reliable continuous monitoring, and typically expensive fabrication methods still needs to be addressed in sweat analysis with wearables. Herein, we propose a fully printed wearable microfluidic nanosensor with an integrated wireless smartphonebased readout. Our system can easily be applied on the skin and actively stimulate perspiration, normalizing the heavy metals concentration with respect to the volume of the sample and the sweat rate. The system has a limit of detection of 396 ppb, a linear range up to 2500 ppb and a sensitivity of 2.3 nA/ppb. | 44652 | 2022
XEBUXF69 | 10.1002/adma.202201768 | https://onlinelibrary.wiley.com/doi/10.1002/adma.202201768 | YES | skin-attached |  | NO | real-time, non-invasive monitoring of specific metabolites in sweat, particularly vitamin C and uric acid. | electrochemical | Electrochemical | Vitamin C Detection: Linear Range: the sensor demonstrates a good linear relationship between current intensity and vitamin C concentration over a wide range of 10–1190 µM. This indicates its ability to accurately quantify vitamin C levels across a broad spectrum of concentrations. Correlation Coefficient: the correlation coefficient for vitamin C detection is reported as 0.9997, indicating a very strong linear correlation between the measured current and the actual vitamin C concentration. This high correlation coefficient suggests excellent accuracy in quantifying vitamin C levels. Low Operating Potential: the sensor requires a relatively low operating potential (approximately 0.001 V) for chronoamperometric measurements of vitamin C. This low potential minimizes the interference from other electroactive species in sweat, enhancing the sensor's selectivity and accuracy. Comparison with HPLC: the researchers validated the sensor's accuracy by comparing its measurements of vitamin C concentration in human sweat with results obtained using high-performance liquid chromatography (HPLC), a gold-standard analytical technique. The average percentage error between the sensor and HPLC measurements was 13.4% across five sweat samples. This comparison indicates a reasonable level of accuracy for the wearable sensor in real-world applications. Uric Acid Detection: Differential Pulse Voltammetry (DPV): the source mentions the use of DPV for uric acid detection, highlighting its ability to effectively distinguish the uric acid signal from those of other electroactive species in sweat. This suggests good selectivity and accuracy for uric acid measurements. Linear Relationship: the DPV responses show a linear increase in current with increasing uric acid concentration, indicating the sensor's ability to accurately quantify uric acid levels. | Metabolic Syndrome | metabolic conditions and nutritional monitoring | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | ascorbic acid, uric acid |  |  |  |  | 1 | 2 | Prototype | Human | Laboratory testing + in-vitro studies + human studies (preliminary) | NA | NA | Conductive Metal-Organic Framework (cMOF): utilizes Ni3HHTP2 for its inherent conductivity, porous structure, and catalytic properties, enabling sensitive, enzyme-free detection of vitamin C and uric acid. Flexible Nanocellulose Substrate: features a breathable bacterial nanocellulose (BNC) substrate for conformal skin adhesion, enhanced comfort, and durability during movement. Wireless Data Transmission: incorporates Bluetooth connectivity for real-time monitoring and smartphone-based health tracking. Multichannel Sensing: simultaneously detects multiple biomarkers, providing comprehensive health insights. | NO | NA | YES | NA | 2022 | Yang, Xue; Yi, Junqi; Wang, Ting; Feng, Yanan; Wang, Jianwu; Yu, Jing; Zhang, Feilong; Jiang, Zhi; Lv, Zhisheng; Li, Haicheng; Huang, Tao; Si, Duanhui; Wang, Xiaoshi; Cao, Rong; Chen, Xiaodong | Wet‐Adhesive On‐Skin Sensors Based on Metal–Organic Frameworks for Wireless Monitoring of Metabolites in Sweat | Advanced Materials |  | 0935-9648, 1521-4095 | Abstract Metal–organic frameworks (MOFs) with well‐defined porous structures and tailored functionalities have been widely used in chemical sensing. However, the integration of MOFs with flexible electronic devices for wearable sensing is challenging because of their low electrical conductivity and fragile mechanical properties. Herein, a wearable sweat sensor for metabolite detection is presented by integrating an electrically conductive Ni‐MOF with a flexible nanocellulose substrate. The MOF‐based layered film sensor with inherent conductivity, highly porous structure, and active catalytic properties enables the selective and accurate detection of vitamin C and uric acid. More importantly, the lightweight sensor can conformably self‐adhere to sweaty skin and exhibits high water‐vapor permeability. Furthermore, a wireless epidermal nutrition tracking system for the in situ monitoring of the dynamics of sweat vitamin C is demonstrated, the results of which are comparable to those tested by high‐performance liquid chromatography. This study opens a new avenue for integrating MOFs as the active layer in wearable electronic devices and holds promise for the future development of high‐performance electronics with enhanced sensing, energy production, and catalytic capabilities through the implementation of multifunctional MOFs. | 44866 | 2022
JJ48E3SD | 10.1007/s00216-020-02430-0 | http://link.springer.com/10.1007/s00216-020-02430-0 | YES | skin-attached |  | NO | non-invasive detection of cortisol levels in human sweat. | electrochemical | Electrochemical | Limit of Detection (LOD): 2.0 ± 0.4 ng/mL. An LOD of 2.0 ng/mL signifies the sensor's capability to detect minute traces of cortisol in sweat, making it suitable for tracking subtle changes in cortisol levels. Dynamic Range: 6–60 ng/mL. The reported dynamic range falls within the physiologically relevant range of cortisol levels in human sweat (8–50 ng/mL), demonstrating the sensor's suitability for real-world applications. Reproducibility: 2.6% Relative Standard Deviation (RSD). An RSD of 2.6% indicates high reproducibility, suggesting that the sensor provides stable and reliable readings across multiple measurements. Real Sample Validation: The accuracy of the sensor was validated by comparing its measurements to those obtained using high-performance liquid chromatography (HPLC) with ultraviolet (UV) detection, a gold-standard analytical technique. The sensor's determination of cortisol content in a real sweat sample (33 ± 2 ng/mL) agreed with the HPLC result (29.0 ± 0.5 ng/mL) to within 5%. This close agreement provides strong evidence for the sensor's accuracy in real-world scenarios. Sensitivity Comparison with Non-Imprinted Polymer (NIP): The molecularly imprinted polymer (MIP) sensor was found to be four orders of magnitude more sensitive than its non-imprinted (NIP) counterpart. This highlights the crucial role of molecular imprinting in enhancing the sensor's sensitivity and selectivity for cortisol, enabling precise measurements even at low concentrations. | Stress | Stress monitoring and management | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | Validation | Human | laboratory experiments + a human studies (a pilot human study) | NA | NA | Flexible and Stretchable Substrate: Built on micropillared PDMS with Ecoflex, it conforms to skin contours, stretches up to 100% without performance loss, ensuring comfort and reliability for sweat collection. Nanoporous Conductive Film: Combines carbon nanotubes (CNTs) and cellulose nanocrystals (CNCs) for high conductivity and sensitivity, creating a robust scaffold for molecularly imprinted polymers (MIP). Molecular Imprinting Technology: Features highly selective cortisol detection via MIP cavities, achieving sensitivity four orders of magnitude greater than non-imprinted systems. Future Microneedle Integration: Potential addition of microneedles for minimally invasive interstitial fluid sampling could expand biomarker analysis capabilities. Wireless Smartphone Data Transmission: Includes a wireless potentiostat for real-time electrochemical data transfer, enhancing portability and ease of use. | NO | NA | NA | NA | 2020 | Mugo, Samuel M.; Alberkant, Jonathan | Flexible molecularly imprinted electrochemical sensor for cortisol monitoring in sweat | Analytical and Bioanalytical Chemistry |  | 1618-2642, 1618-2650 | A selective cortisol sensor based on molecularly imprinted poly(glycidylmethacrylate-co ethylene glycol dimethacrylate) (poly(GMA-co-EGDMA)) has been demonstrated for detection of cortisol in human sweat. The non-enzymatic biomimetric flexible sweat sensor was fabricated inexpensively by layer by layer (LbL) assembly. The sensor layers comprised a stretchable polydimethylsiloxane (PDMS) base with carbon nanotubes-cellulose nanocrystals (CNC/CNT) conductive nanoporous nanofilms. The imprinted (MIP) poly(GMA-co-EGDMA) deposited on the CNC/CNT was the cortisol biomimetric receptor. Rapid in analyte response (3 min), the cortisol MIP sensor demonstrated excellent performance. The sensor has a limit of detection (LOD) of 2.0 ng/mL ± 0.4 ng/mL, dynamic range of 10–66 ng/mL, and a sensor reproducibility of 2.6% relative standard deviation (RSD). The MIP sensor also had high cortisol specificity and was inherently blind to selected interfering species including glucose, epinephrine, β-estradiol, and methoxyprogestrone. The MIP was four orders of magnitude more sensitive than its non-imprinted (NIP) counterpart. The MIP sensor remains stable over time, responding proportionately to doses of cortisol in human sweat. | 43891 | 2020
CNL5PLGL | 10.3390/s21051642 | https://www.ncbi.nlm.nih.gov/pubmed/33652955 | YES | not specified |  | NO | saliva monitoring for early diagnosis of diseases | electrochemical | Electrochemical | The sensor exhibits good sensitivity, measured as 58.9 mV/decade. The sensor demonstrates selectivity for sodium ions, with a selectivity coefficient (log K) of -2.68 for potassium. The sensor has a broad detection range of 5 × 10^-5 ≈ 1 M and a low detection limit of 4.27 × 10^-5 M. |  | NA | - | yes | - | saliva | [12 empty] | sodium |  |  |  |  | 1 | 1 | Prototype/Validation | Human | Human studies | NA | NA | A flexible system, including electrochemical sensors and a Bluetooth-enabled wireless circuit, has an exceptionally small form factor,showing an excellent possibility for non invasive,portable disease diagnosis. | NO | NA | NA | NA | 2021 | Lim, H. R.; Lee, S. M.; Mahmood, M.; Kwon, S.; Kim, Y. S.; Lee, Y.; Yeo, W. H. | Development of Flexible Ion-Selective Electrodes for Saliva Sodium Detection | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | Saliva can be used for health monitoring with non-invasive wearable systems. Such devices, including electrochemical sensors, may provide a safe, fast, and cost-efficient way of detecting target ions. Although salivary ions are known to reflect those in blood, no available clinical device can detect essential ions directly from saliva. Here, we introduce an all-solid-state, flexible film sensor that allows highly accurate detection of sodium levels in saliva, comparable to those in blood. The wireless film sensor system can successfully measure sodium ions from a small volume of infants' saliva (<400 microL), demonstrating its potential as a continuous health monitor. This study includes the structural characterization and error analysis of a carbon/elastomer-based ion-selective electrode and a reference electrode to confirm the signal reliability. The sensor, composed of a pair of the electrodes, shows good sensitivity (58.9 mV/decade) and selectivity (log K = -2.68 for potassium), along with a broad detection range of 5 x 10(-5) approximately 1 M with a low detection limit of 4.27 x 10(-5) M. The simultaneous comparison between the film sensor and a commercial electrochemical sensor demonstrates the accuracy of the flexible sensor and a positive correlation in saliva-to-blood sodium levels. Collectively, the presented study shows the potential of the wireless ion-selective sensor system for a non-invasive, early disease diagnosis with saliva. | 44253 | 2021
D8GPBM8B | 10.1016/j.bios.2023.115412 | https://linkinghub.elsevier.com/retrieve/pii/S0956566323003548 | YES | skin-attached |  | NO | on-demand sweat bioanalysis | electrochemical | Electrochemical | The glucose sensor exhibits a linear response to glucose concentrations ranging from 5.0 to 50 μM, as demonstrated by the chronoamperometric measurements. The accuracy of the glucose sensor was assessed by comparing its measurements with those obtained from a commercial glucometer using blood samples. There is a close correlation between the sweat glucose levels measured by the wearable sensor and the blood glucose levels, indicating a reasonable level of accuracy. The cortisol sensor demonstrates a linear relationship between the charge transfer resistance and cortisol concentration within a range of 1.0 to 15 ng/mL. There is a good correlation (CV<10%) between the cortisol levels determined by the wearable sensor and the ELISA kit. | Diabetes | monitoring and assisting with the management of diabetes and stress-related conditions. | yes | - | - | sweat |  |  |  |  | yes |  | yes | [5 empty] | glucose, cortisol |  |  |  |  | 1 | 2 | Proof-of-concept | Human | Human studies | NA | NA | The device utilizes a thermoresponsive hydrogel made from poly(N-isopropylacrylamide) (pNIPAM) to passively collect sweat without requiring external stimulation like exercise or heat.1234 This is a significant advancement compared to conventional sweat collection methods that can distort biomarker levels and may not be suitable for all individuals. | NO | NA | NA | NA | 2023 | Chen, Yichen; Ma, Biao; Zuo, Yinxiu; Chen, Gangsheng; Hao, Qing; Zhao, Chao; Liu, Hong | Versatile sweat bioanalysis on demand with hydrogel-programmed wearables | Biosensors and Bioelectronics |  | 9565663 | Wearable sweat bioanalysis is promising for non-invasive diagnostics of diseases. However, collection of representative sweat samples without disturbing daily life and wearable bioanalysis of targets that are clinically significant are still challenging. In this work, we report on a versatile method for the sweat bioanalysis. The method is based on a thermoresponsive hydrogel which can imperceptibly absorb slowly secreted sweat without stimulation such as heat or sport exercise. The wearable bioanalysis is accomplished by programmed electric heating of hydrogel modules to 42◦C to release absorbed sweat or preloaded reagents into a microfluidic detection channel. Using our method, not only one-step detection of glucose but also multi-step immunoassay of cortisol is accomplished within 1 h, even at a very low sweat rate. Our test results are also compared with those obtained with conventional blood samples and stimulated sweat samples to evaluate the applicability of our method to non-invasive clinical practice. | 45170 | 2023
L7EPPAFX | 10.1002/adma.202008465 | https://onlinelibrary.wiley.com/doi/10.1002/adma.202008465 | YES | patch |  | NO | rapid, stress-free detection of cortisol levels in sweat | electrochemical | Electrochemical | Sensitivity in Buffer Solution (PBS): 38.8 nA log[nM]⁻¹. Sensitivity in Artificial Sweat (AS): 60.31 nA log[nM]⁻¹, indicating enhanced performance due to specific interactions with the sweat matrix. Detection Threshold: capable of detecting changes at 1 nM cortisol concentration, enabling precise monitoring of subtle fluctuations. The device's accuracy is validated by a strong correlation with standard immunoassay results, confirming its reliability for cortisol detection. | Stress | monitoring stress levels | yes | - | - | sweat | [6 empty] | yes | [5 empty] | cortisol |  |  |  |  | 1 | 1 | validation | Human | in-vitro tests + human studies | NA | NA | Molecularly Imprinted Polymer (MIP) Technology: a highly selective "lock-and-key" mechanism ensures precise cortisol detection by binding specifically to cortisol molecules. Integrated Prussian Blue Redox Probes: embedded in the MIP layer, these probes enable direct electrochemical signaling without external labels, simplifying the detection process. Rapid Chronoamperometric Detection: optimized for a fast assay time of 3.5 minutes, enabling near real-time cortisol monitoring. Porous PVA Hydrogel: effortlessly absorbs sweat through natural perspiration, avoiding stress-inducing external stimulation methods. Stretchable Epidermal Patch: provides continuous monitoring with a flexible, skin-conforming design suitable for use during physical activity. | NO | NA | NA | NA | 2021 | Tang, Wanxin; Yin, Lu; Sempionatto, Juliane R.; Moon, Jong‐Min; Teymourian, Hazhir; Wang, Joseph | Touch‐Based Stressless Cortisol Sensing | Advanced Materials |  | 0935-9648, 1521-4095 | Abstract Tracking fluctuations of the cortisol level is important in understanding the body's endocrine response to stress stimuli. Traditional cortisol sensing relies on centralized laboratory settings, while wearable cortisol sensors are limited to slow and complex assays. Here, a touch‐based non‐invasive molecularly imprinted polymer (MIP) electrochemical sensor for rapid, simple, and reliable stress‐free detection of sweat cortisol is described. The sensor readily measures fingertip sweat cortisol via highly selective binding to the cortisol‐imprinted electropolymerized polypyrrole coating. The MIP network is embedded with Prussian blue redox probes that offer direct electrical signaling of the binding event to realize sensitive label‐free amperometric detection. Using a highly permeable sweat‐wicking porous hydrogel, instantaneously secreted fingertip sweat can be conveniently and rapidly collected without any assistance. By eliminating time lags, such rapid (3.5 min) fingertip assay enables the capture of sharp variations in cortisol levels, compared to previous methods. Such advantages are demonstrated by tracking cortisol response in short cold‐pressor tests and throughout day‐long circadian rhythm, along with gold‐standard immunoassay validation. A stretchable epidermal MIP sensor is also described for directly tracking cortisol in exercise‐induced sweat. The rapid touch‐based cortisol sensor offers an attractive, accessible, stressless avenue for quantitative stress management. | 44317 | 2021
YUUB64D6 | 10.3390/bios12121095 | https://www.ncbi.nlm.nih.gov/pubmed/36551062 | YES | skin-attached |  | NO | measuring prognostic TBI markers relevant to monitoring of a TBI post-injury | electrochemical | Electrochemical | The LOD for GFAP detection in the sweat analog was estimated to be 14 pg/mL using EIS. The LOD for IL-6 was estimated to be 10 pg/mL using EIS | Neuro | Traumatic Brain Injuries management | yes | - | - | sweat | [12 empty] | glial fibrillary acidic protein (GFAP) and Interleukin-6 (IL6) |  |  |  |  | 1 | 2 | Proof-of-concept |  | In-vitro tests | NA | NA | SWEATSENSER is a noninvasive enabling technology that enables baseline inflammatory response tracking (Flexible Substrates + High-Sensitivity Sensors + Multiplexed Sensing). | EnLiSense LLC | USA | NA | NA | 2022 | Shahub, S.; Lin, K. C.; Muthukumar, S.; Prasad, S. | A Proof-of-Concept Electrochemical Skin Sensor for Simultaneous Measurement of Glial Fibrillary Acidic Protein (GFAP) and Interleukin-6 (IL-6) for Management of Traumatic Brain Injuries | Biosensors (Basel) |  | 2079-6374 (Electronic) 2079-6374 (Linking) | This work demonstrates the use of a noninvasive, sweat-based dual biomarker electrochemical sensor for continuous, prognostic monitoring of a Traumatic Brain Injury (TBI) with the aim of enhancing patient outcomes and reducing the time to treatment after injury. A multiplexed SWEATSENSER was used for noninvasive continuous monitoring of glial fibrillary acidic protein (GFAP) and Interleukin-6 (IL-6) in a human sweat analog and in human sweat. Electrochemical impedance spectroscopy (EIS) and chronoamperometry (CA) were used to measure the sensor response. The assay chemistry was characterized using Fourier Transform Infrared Spectroscopy (FTIR). The SWEATSENSER was able to detect GFAP and IL-6 in sweat over a dynamic range of 3 log orders for GFAP and 2 log orders for IL-6. The limit of detection (LOD) for GFAP detection in the sweat analog was estimated to be 14 pg/mL using EIS and the LOD for IL-6 was estimated to be 10 pg/mL using EIS. An interference study was performed where the specific signal was significantly higher than the non-specific signal. Finally, the SWEATSENSER was able to distinguish between GFAP and IL-6 in simulated conditions of a TBI in human sweat. This work demonstrates the first proof-of-feasibility of a multiplexed TBI marker combined with cytokine and inflammatory marker detection in passively expressed sweat in a wearable form-factor that can be utilized toward better management of TBIs. This is the first step toward demonstrating a noninvasive enabling technology that can enable baseline tracking of an inflammatory response. | 44895 | 2022
DWD9G8IH | 10.1016/j.talanta.2022.124122 | https://linkinghub.elsevier.com/retrieve/pii/S0039914022009183 | YES | skin-attached |  | NO | non-invasive monitoring of lactate levels in interstitial fluid (ISF) | electrochemical | Electrochemical | The in-vitro characterization of the lactate biosensor revealed a limit of detection of 0.15 mM. This means the sensor can reliably detect lactate concentrations at or above this value. The sensor exhibited a linear response to lactate concentrations ranging from 1 mM to 5 mM1. This suggests the sensor's output signal is directly proportional to the lactate concentration within this range. The calibration plot of the sensor, derived from the chronoamperometric signal against lactate concentration, fitted a linear equation: y = - 4.95E−07[Lactate] – 1.34E−07, with an R2 of 0.99. The high R2 value indicates a strong linear relationship between the sensor's output and the lactate concentration. |  | NA | - | - | yes | interstitial fluid | [5 empty] | yes | [6 empty] | Lactate |  |  |  |  | 1 | 1 | Prototype/Validation | Human | Human studies | NA | NA | The most significant innovation is the all-in-one platform that combines reverse iontophoresis (RI) for ISF extraction with an enzymatic amperometric biosensor for lactate detection on a single, flexible, and adhesive patch.This integration eliminates the need for separate extraction and analysis steps, enabling rapid, point-of-care lactate monitoring | NO | NA | NA | NA | 2023 | De La Paz, Ernesto; Saha, Tamoghna; Del Caño, Rafael; Seker, Sumeyye; Kshirsagar, Nikhil; Wang, Joseph | Non-invasive monitoring of interstitial fluid lactate through an epidermal iontophoretic device | Talanta |  | 399140 | The development of a non-invasive sensing technology that allows collection of interstitial fluid (ISF) lactate and its subsequent analysis without exertion requirement, could enable lactate monitoring from rested individuals. Here, we describe a wearable, soft epidermal adhesive patch that integrates a reverse iontophoretic (RI) system, and an amperometric lactate biosensor placed on the anodic electrode with a porous hydrogel reservoir, for simultaneous ISF lactate extraction and quantification via electrochemical sensing, respectively. The iontophoretic system includes agarose hydrogels for preventing skin electrocution, while a porous polyvinyl alcohol-based hydrogel facilitates the effective transport of lactate from skin to the biosensor. The flexible skin-worn device tested on healthy individuals at rest showed rapid lactate collection from the ISF after 10 min of reverse iontophoresis with no evidence of discomfort or irritation to the skin. Detailed characterization of the enzymatic biosensor before and during on-body trials along with relevant control experiments confirmed the efficient extraction and selective detection of ISF lactate. Such an epidermal technology represents the first demonstration of an all-in-one platform that integrates non-invasive collection and subsequent analysis of lactate from iontophoretically extracted ISF toward point-of-care operation. | 44986 | 2023
T9NTIH34 | 10.1002/adma.202403758 | https://www.ncbi.nlm.nih.gov/pubmed/38733567 | YES |  |  | Minimally | transdermal sensing of enzyme biomarkers. | electrochemical | Electrochemical | The epidermal microneedle sensor also responded with good linearity, and sensi-tivity against increasing concentrations of Tyr (between 0.3 and0.7 mg mL−1 ; R2 > 0.99, 5.52 ± 0.58 mA g mL−1 cm−2 , with avery low limit of detection (0.06 mg mL−1 ; S/N = 3). The reproducibility of the epidermal patch sensor is also very high (RSD< 5% (n = 5) for both device-to-device or repetitive measurements using a single device | Cancer | Melanoma detection | - | - | yes | interstitial fluid | [12 empty] | Tyrosinase (Tyr) |  |  |  |  | 1 | 1 | Proof-of-concept |  | Ex-vivo | NA | NA | A wearable bioelectronic patch is established for transdermal sensing of the melanoma biomarker, tyrosinase (Tyr), using a microneedle array integrated with a surface-bound chemo-responsive smart probe to enable target-specific electrochemical detection of Tyr directly from human skin tissue. | NO | NA | NA | NA | 2024 | Poursharifi, N.; Hassanpouramiri, M.; Zink, A.; Ucuncu, M.; Parlak, O. | Transdermal Sensing of Enzyme Biomarker Enabled by Chemo-Responsive Probe-Modified Epidermal Microneedle Patch in Human Skin Tissue | Adv Mater |  | 1521-4095 (Electronic) 0935-9648 (Linking) | Wearable bioelectronics represents a significant breakthrough in healthcare settings, particularly in (bio)sensing which offers an alternative way to track individual health for diagnostics and therapy. However, there has been no notable improvement in the field of cancer, particularly for skin cancer. Here, a wearable bioelectronic patch is established for transdermal sensing of the melanoma biomarker, tyrosinase (Tyr), using a microneedle array integrated with a surface-bound chemo-responsive smart probe to enable target-specific electrochemical detection of Tyr directly from human skin tissue. The results presented herein demonstrate the feasibility of a transdermal microneedle sensor for direct quantification of enzyme biomarkers in an ex vivo skin model. Initial performance analysis of the transdermal microneedle sensor proves that the designed methodology can be an alternative for fast and reliable diagnosis of melanoma and the evaluation of skin moles. The innovative approach presented here may revolutionize the landscape of skin monitoring by offering a nondisruptive means for continuous surveillance and timely intervention of skin anomalies, such as inflammatory skin diseases or allergies and can be extended to the screening of multiple responses of complementary biomarkers with simple modification in device design. | 45474 | 2024
EGMLIHS2 | 10.1016/j.bios.2024.116012 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324000150 | YES | skin-attached |  | NO | non-invasive detection of specific metabolites in sweat | electrochemical | Electrochemical | Tyrosine (Tyr): the LOD for tyrosine detection is reported as 5.11 μM. This indicates the sensor's ability to detect very low concentrations of tyrosine in sweat. Uric Acid (UA): the LOD for uric acid is even lower, reaching 1.37 μM, showcasing the device's exceptional sensitivity for this metabolite. The researchers attribute this high sensitivity to several factors, including: FeNCs/LIG Electrode: the use of iron nano-catalysts incorporated laser-induced graphene (FeNCs/LIG) as the working electrode material plays a crucial role. The three-dimensional nanostructure of LIG and the electrocatalytic activity of FeNCs enhance electron transfer, leading to a more sensitive response. Hydrogel Chip Design: the dual-channel hydrogel chip facilitates efficient sweat collection and analyte delivery to the sensing electrodes, ensuring a sufficient sample volume for accurate detection even at low concentrations. | Metabolic Syndrome | monitoring metabolic health and nutritional status. | yes | - | - | sweat | [7 empty] | yes |  |  |  |  | Tyrosine, Uric Acid |  |  |  |  | 1 | 2 | Prototype |  | laboratory testing + in-vitro tests | NA | NA | Iron Nano-Catalysts on Laser-Induced Graphene (FeNCs/LIG): combines a 3D porous, conductive LIG substrate with iron nano-catalysts for enhanced electrocatalytic activity and sensitivity. Hydrogel-Based Dual-Channel Microfluidic Chip: ensures efficient sweat collection and wearability, with separate channels enabling simultaneous multi-metabolite detection (e.g., tyrosine and uric acid). Flexible and Biocompatible Materials: the hydrogel's mechanical properties and LIG's flexibility ensure a secure and comfortable fit on the skin. Computational Modeling: density functional theory (DFT) simulations confirm the Fe active sites' role in boosting molecular adsorption and electron transfer. | NO | NA | NA | NA | 2024 | Zhao, Peng; Zhang, Yong; Liu, Yiyi; Huo, Danqun; Hou, Jingzhou; Hou, Changjun | Wearable electrochemical patch based on iron nano-catalysts incorporated laser-induced graphene for sweat metabolites detection | Biosensors and Bioelectronics |  | 9565663 | The development of wearable devices shows great application potential in health management. In this work, we propose the fabrication of a novel wearable electrochemical patch and prove its application in sweat metabolites detection. The patch is developed based on iron nano-catalysts incorporated laser-induced graphene (FeNCs/ LIG), which is a newly integrated sensing electrode with unique three-dimensional nanostructure and good electrocatalytic activity. It shows desirable sensing performances for sweat metabolites including tyrosine (Tyr) and uric acid (UA) molecules. The detection limit of Tyr and UA can reach 5.11 μM and 1.37 μM, respectively. Besides, density functional theory calculation deeply reveals that the Fe active sites of FeNCs play an important role in molecule adsorption and electron transference, thus promoting sensing performance. To realize wearable application, a dual-channel hydrogel chip is designed and assembled with FeNCs/LIG. The developed patch is successfully utilized to accurately determination of Tyr and UA in sweat. This work is expected to provide a new non-invasive strategy for evaluating amino acid intake and metabolic level. | 45383 | 2024
VXTKM3A9 | 10.1016/j.bios.2024.116354 | https://linkinghub.elsevier.com/retrieve/pii/S0956566324003592 | YES | patch |  | NO | continuous, non-invasive monitoring of glucose levels in sweat | electrochemical | Electrochemical | Linear Relationship: the research highlights a strong linear relationship between the glucose concentration and the sensor's current response. This linearity suggests a good degree of accuracy in measuring glucose levels within the specified detection range. Limit of Detection (LOD): the reported LOD of 0.762 μM indicates the sensor's ability to detect very low concentrations of glucose, suggesting high sensitivity. Stability Tests: the researchers conducted stability tests using five independent sensors, demonstrating a relative standard deviation (RSD) of 5.2%. Glucose Recovery Experiments: the article mentions glucose recovery experiments that showed excellent consistency between the sensor's measurements and the actual glucose concentrations. | Diabetes | diabetes management | yes | - | - | sweat |  |  |  |  | yes | [7 empty] | glucose |  |  |  |  | 1 | 1 | Prototype/Validation | Human | laboratory testing + human studies (preliminary) | NA | NA | Single-Atom Catalyst (SAC): utilizes Pt atoms on Co₃O₄ nanorods and rGO for enhanced sensitivity and catalytic activity in glucose detection. Synergistic Catalysis: combines Pt, Co₃O₄, and rGO for superior glucose oxidation, confirmed by DFT calculations showing improved glucose binding and activity. Flexible Substrate: built on polyimide (PI) film for comfortable, body-conforming wear during physical activities. Hydrophilic Microfluidic Chip: features surfactant-modified "S-shaped" PDMS channels for efficient sweat transport and reduced resistance. Wireless Communication: bluetooth integration enables real-time, convenient glucose monitoring on smartphones. Point-of-Care Testing (POCT): facilitates immediate on-site sweat analysis, bypassing laboratory needs for rapid feedback. | NO | NA | YES | NA | 2024 | Zhang, Yong; Yang, Zhihao; Qiao, Cailin; Liu, Yiyi; Wang, Cuncun; Zeng, Xin; Hou, Jingzhou; Huo, Danqun; Hou, Changjun | Synergistic enhancement of wearable biosensor through Pt single-atom catalyst for sweat analysis | Biosensors and Bioelectronics |  | 9565663 | Real-time monitoring of biological markers in sweat is a valuable tool for health assessment. In this study, we have developed an innovative wearable biosensor for precise analysis of glucose in sweat during physical activities. The sensor is based on a single-atom catalyst of platinum (Pt) uniformly dispersed on tricobalt tetroxide (Co3O4) nanorods and reduced graphene oxide (rGO), featuring a unique three-dimensional nanostructure and excellent glucose electrocatalytic performance with a wide detection range of 1–800 μM. Additionally, density functional theory calculations have revealed the synergetic role of Pt active sites in the Pt single-atom catalyst (Co3O4/rGO/Pt) in glucose adsorption and electron transfer, thereby enhancing sensor performance. To enable application in wearable devices, we designed an S-shaped microfluidic chip and a point-of-care testing (POCT) device, both of which were validated for effectiveness through actual use by volunteers. This research provides valuable insights and innovative approaches for analyzing sweat glucose using wearable devices, contributing to the advancement of personalized healthcare. | 45505 | 2024
CXX6K392 | 10.1016/j.talanta.2023.124610 | https://linkinghub.elsevier.com/retrieve/pii/S0039914023003612 | YES | patch |  | NO | real-time, non-invasive monitoring of multiple biomarkers in sweat | electrochemical | Electrochemical | Sodium (Na+) Sensor: the Na+ sensor exhibits a sensitivity of 53.83 mV/decade. pH Sensor: the pH sensor demonstrates a sensitivity of 98.68 mV/decade. Glucose Sensor: the glucose sensor shows a sensitivity of 0.04 μA μM−1. |  | NA | yes | - | - | sweat |  |  |  |  | yes | [5 empty] | yes |  | Na+, pH, Glucose |  |  |  |  | 1 | 3 | Prototype | Human | laboratory testing + human studies (preliminary) | NA | NA | All-Fabric Construction: made from breathable silk fabric, enhancing comfort, flexibility, and wearability over rigid materials. Janus Fabric for Sweat Transport: uses asymmetric wettability for one-way sweat flow, improving sensor performance and efficiency. Silk-Based Carbon Cloth Electrodes: biocompatible, flexible, and conductive electrodes created via carbonizing silk fibroin. Laser-Engraved Sensor Patterns: CO₂ laser cutting enables precise, multi-sensor integration on the fabric substrate. Multiplexed Biomarker Detection: simultaneous detection of Na⁺, pH, and glucose for comprehensive physiological monitoring. | NO | NA | NA | NA | 2023 | Zhang, Yingwen; Liao, Jianjun; Li, Zehao; Hu, Mingxu; Bian, Chao; Lin, Shiwei | All fabric and flexible wearable sensors for simultaneous sweat metabolite detection and high-efficiency collection | Talanta |  | 399140 | Wearable sweat electrochemical sensors have attracted wide attention due to their advantages of non-invasive, portable, real-time monitoring, etc. However, existing sensors still have some problems with efficient sweat collection. Microfluidic channel technology and electrospinning technology are commonly used to collect sweat efficiently, but there are some limitations such as complex channel design and multiple spinning parameters. Besides, existing sensors are mostly based on flexible polymers, such as, PET, PDMS, PI and PI, which have limited wearability and permeability. Based on the above, all fabric and dual-function flexible wearable sweat electrochemical sensor is proposed in this paper. This sensor uses fabric as the raw material to implement the directional transport of sweat and the multi-component integrated detection dual functions. Meanwhile, the high-efficiency collection of sweat is obtained by a Janus fabric, wherein one side of the selected silk is superhydrophobic graft treated and the other side is hydrophilic plasma treated. Therefore, the resulting Janus fabric can effectively transfer sweat from the skin side to the electrode, and the minimum sweat droplet can reach 0.2 μL to achieve micro-volume collection. Besides, the patterned sensor, made of silk-based carbon cloth, is fabricated using a simple laser engraving, which could detect Na+, pH, and glucose instantaneously. As a result, these proposed sensors can achieve good sensing performance and high-efficiency sweat collection dual functionality; moreover, it has good flexibility and comfortable wearability. | 45139 | 2023
LXY688CC | 10.1016/j.bios.2017.06.053 | https://linkinghub.elsevier.com/retrieve/pii/S0956566317304396 | YES | patch |  | NO | quantitatively measuring wound biomarkers in situ | electrochemical | Electrochemical | Sensitivity: Linear response to uric acid concentrations from 0 to 800 µM with a strong correlation coefficient (R² = 0.995), indicating high accuracy. Reproducibility: Standard deviations (SDs) are under 7% across three measurements, demonstrating consistent performance. Stability: Maintains accuracy for up to 7 hours of continuous monitoring with a coefficient of variance (COV = 0.05). Specificity: Negligible signals for interfering analytes (e.g., glucose, lactate), ensuring high selectivity for uric acid detection. |  | NA | - | - | - | Simulated wound fluid | [7 empty] | yes |  |  |  |  | uric acid (UA) |  |  |  |  | 1 | 1 | Proof-of-concept |  | in-vitro + laboratory testing | NA | NA | Embroidered Electrodes: Uses embroidery to integrate conductive threads onto flexible, breathable gauze, overcoming limitations of traditional screen-printing for porous materials. Ink-Coated Threads: Features carbon-coated working/counter electrodes and Ag/AgCl-coated reference electrodes, ensuring strong electrochemical performance and biocompatibility. Flexible Substrate: Gauze provides flexibility, breathability, and absorbency, enhancing comfort and enabling seamless integration with wound dressings. Customization: Allows precise electrode design and potential for multiplexed detection, enabling tailored applications and simultaneous biomarker monitoring. Biocompatibility Potential: Suggests future use of coatings like chitosan to increase lifespan and support long-term wound care. | NO | NA | NA | NA | 2017 | Liu, Xiyuan; Lillehoj, Peter B. | Embroidered electrochemical sensors on gauze for rapid quantification of wound biomarkers | Biosensors and Bioelectronics |  | 9565663 | Electrochemical sensors are an attractive platform for analytical measurements due to their high sensitivity, portability and fast response time. These attributes also make electrochemical sensors well suited for wearable applications which require excellent ﬂexibility and durability. Towards this end, we have developed a robust electrochemical sensor on gauze via a unique embroidery fabrication process for quantitative measurements of wound biomarkers. For proof of principle, this biosensor was used to detect uric acid, a biomarker for wound severity and healing, in simulated wound ﬂuid which exhibits high speciﬁcity, good linearly from 0 to 800 µM, and excellent reproducibility. Continuous sensing of uric acid was also performed using this biosensor which reveals that it can generate consistent and accurate measurements for up to 7 h. Experiments to evaluate the robustness of the embroidered gauze sensor demonstrate that it oﬀers excellent resilience against mechanical stress and deformation, making it a promising wearable platform for assessing and monitoring wound status in situ. | 43070 | 2017
AMMK95CX | 10.1016/j.bios.2021.113270 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321003079 | YES | head-mounted | band | NO | simultaneous detection and monitoring of lactate and sodium levels in human sweat | electrochemical | Electrochemical | Limits of Detection (LODs): Lactate: 3.61 mM Sodium: 0.16 mM Linear Ranges: the linear range is the concentration range over which the sensor's response is directly proportional to the analyte concentration. This is the range where the sensor provides the most accurate and reliable measurements. The reported linear ranges are: Lactate: 0–25 mM Sodium: 0.1–100 mM These ranges cover the clinically relevant concentrations of lactate and sodium in human sweat, indicating the device's suitability for practical applications. Comparison to Standard Tests: the researchers conducted off-body testing of real human sweat using the wearable biosensor and compared the results to standard laboratory tests for lactate and sodium. The findings showed that the biosensor's measurements were "in good agreement" with the standard test results. This strongly suggests that the device is capable of providing accurate and reliable data, comparable to established laboratory methods. Standard Deviation of Measurements: while not a direct measure of accuracy, the reported standard deviations of the measurements provide an indication of the precision and variability of the sensor's output. Lower standard deviations suggest greater precision. |  | NA | yes | - | - | sweat | [5 empty] | yes | [6 empty] | lactate, Na+ |  |  |  |  | 1 | 2 | Prototype/Validation | Human | laboratory testing + in-vitro tests + human studies (on-body and off-body testing) | NA | NA | ZnO Nanowire-Decorated Electrodes: ZnO nanowires enhance sensitivity with high surface area and efficient electron transport, and enable stable enzyme immobilization. Thread-Based Design: uses conductive threads for flexibility and wearability, allowing seamless integration into textiles like clothing or headbands. Integrated 'Smart' Headband: includes a microcontroller-based circuit for real-time data processing and wireless Bluetooth communication for external data transmission. Wireless Monitoring: facilitates real-time sweat biomarker tracking via wireless communication to smartphones or computers. | NO | NA | NA | NA | 2021 | Zhao, Chen; Li, Xiao; Wu, Qiyang; Liu, Xinyu | A thread-based wearable sweat nanobiosensor | Biosensors and Bioelectronics |  | 9565663 | Non-invasive wearable biosensors provide an efficient way of continuously quantifying a person’s biochemical parameters, and are highly valuable for predicting human physiological status and flagging risks and illness. Commercial wearable sensors are available for tracking a user’s physical activities, but few could monitor user’s health conditions through sweat analysis. Electronic textile (e-textile) biosensors enable new applications in this scenario because of its high flexibility/wearability, low cost, high level of electronic integration, and unobtrusiveness. However, challenges in developing e-textile sweat biosensors remain in the production of textile-based biosensing materials, skin interfacing design, and embedded data acquisition/transmission. Here, we propose a novel wearable electrochemical sweat biosensor based on conductive threads decorated with zinc-oxide nanowires (ZnO NWs) and apply it to detecting lactate and sodium in perspiration during physical exercise. The sweat biosensor is fully integrated with signal readout and data communication circuits in a wearable headband and is capable of monitoring human sweat accurately and wirelessly. We achieved the detection of lactate and sodium in linear ranges of 0–25 mM and 0.1–100 mM and limits of detection of 3.61 mM and 0.16 mM, respectively, which cover the clinically-relevant ranges of lactate and sodium in human sweat. We demonstrated accurate lactate and sodium measurements in human sweat from a healthy volunteer, and the results are in good agreement with standard test results. We also conducted on-body measurements on the same human volunteer during exercise and confirmed the robustness of the signal readout during body movements and the excellent accuracy of the testing results. | 44440 | 2021
KHLH2SEZ | 10.1016/j.bios.2021.113650 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321006874 | YES | wrist-worn |  | NO | non-invasive detection of transdermal volatile biomarkers to assess dietary macronutrient intake. | Optical | Optical | The transdermal acetone response from the GCAS sensor correlates well with breath acetone concentration in the range between 0 and 40 ppm and the correlation factor (R2) is as high as 0.8877. | Metabolic Syndrome | weight control, disease and nutrition management | - | - | - | skin gas | [12 empty] | acetone, ammonia, hydrogen, CO2 |  |  |  |  | 1 | 4 | Proof-of-concept/validation | Human | laboratory testing + human studies (preliminary) | NA | NA | Gradient-Based Colorimetric Array Sensor (GCAS): employs a novel chip with gradient sensing channels that detect volatile biomarkers via colorimetric changes proportional to biomarker concentration. Image-Based Signal Processing: uses a miniaturized imaging system with advanced algorithms to analyze real-time color changes, providing quantitative biomarker data. Porous Sensing Matrix: enhances sensitivity and response time by enabling efficient diffusion of skin gas biomarkers. Multiplexed Detection: simultaneously detects multiple biomarkers, such as acetone, ammonia, and CO₂, for comprehensive metabolic monitoring. Wearable Design: compact wristband integrates a gas diffusion sheet for unobtrusive, real-world use. | NO | NA | YES | NA | 2022 | Yu, Jingjing; Wang, Di; Tipparaju, Vishal Varun; Jung, Wonjong; Xian, Xiaojun | Detection of transdermal biomarkers using gradient-based colorimetric array sensor | Biosensors and Bioelectronics |  | 9565663 | Accurate assessment of dietary macronutrients intake is critical for the effective management of multiple diseases, such as obesity, diabetes, cardiovascular disease, metabolic disease, and cancer. Conventional selfreporting method is burdensome, inaccurate, and often biased. Though blood analysis and breath analysis can provide evidence-based information, they are either invasive or subject to human errors. Here we reported a wearable transdermal volatile biomarkers detection system based on novel colorimetric sensing technology for dietary macronutrients intake assessment. This technique quantifies the emission rates of transdermal volatile biomarkers via a gradient-based colorimetric array sensor (GCAS). The optical system of the GCAS device tracks the localized color development associated with the chemical reaction between the volatile biomarkers and the porous sensing probes, and determines the biomarkers emission rates through image processing algorithms. The localized chemical reaction and the image-based signal processing also make the GCAS capable for multiplexed detection of multiple analytes simultaneously. The GCAS sensor has been applied for transdermal acetone detection on 5 subjects in a keto diet intervention. The study indicates that the transdermal acetone increases after the subjects consuming keto diets and it decreases to basal level after intaking carb-rich diets. The transdermal acetone response from the GCAS sensor correlates well with breath acetone concentration in the range between 0 and 40 ppm and the correlation factor (R2) is as high as 0.8877. This method provides a noninvasive, low-cost, and wearable tool for assessing dietary macronutrients intake outside of lab or hospital settings. It could be widely applied in disease management, weight control, and nutrition management. | 44562 | 2022
JJGHS83P | 10.1038/s41467-020-18238-6 | https://www.ncbi.nlm.nih.gov/pubmed/32879320 | YES | wrist-worn | smartwatch-like | NO | measuring of biomarkers in epidermally-retrievable biofluids | electrochemical | Electrochemical (Voltammetry) | The device's data based on a strong correlation with laboratory-based assays (R2 ≈ 0.94). The response of the glucose and lactate sensors were validated within the respective analytes’ physiologically relevant concentration range in sweat. Linear relationships were observed between the measured current responses and target analytes’ concentration levels (R2 = 0.99, for both sensors). |  | NA | yes | yes | - | sweat | saliva | [9 empty] | yes |  | Acetaminophen | Acetaminophen |  |  |  | 2 | 1 | Validation | Human | Human studies | NA | NA | A wearable voltammetric sensor, using a surface-modified boron-doped diamond sensing interface to track the temporal profile of circulating acetaminophen in these fluids. | NO | NA | YES | NA | 2020 | Lin, H.; Tan, J.; Zhu, J.; Lin, S.; Zhao, Y.; Yu, W.; Hojaiji, H.; Wang, B.; Yang, S.; Cheng, X.; Wang, Z.; Tang, E.; Yeung, C.; Emaminejad, S. | A programmable epidermal microfluidic valving system for wearable biofluid management and contextual biomarker analysis | Nat Commun |  | 2041-1723 (Electronic) 2041-1723 (Linking) | Active biofluid management is central to the realization of wearable bioanalytical platforms that are poised to autonomously provide frequent, real-time, and accurate measures of biomarkers in epidermally-retrievable biofluids (e.g., sweat). Accordingly, here, a programmable epidermal microfluidic valving system is devised, which is capable of biofluid sampling, routing, and compartmentalization for biomarker analysis. At its core, the system is a network of individually-addressable microheater-controlled thermo-responsive hydrogel valves, augmented with a pressure regulation mechanism to accommodate pressure built-up, when interfacing sweat glands. The active biofluid control achieved by this system is harnessed to create unprecedented wearable bioanalytical capabilities at both the sensor level (decoupling the confounding influence of flow rate variability on sensor response) and the system level (facilitating context-based sensor selection/protection). Through integration with a wireless flexible printed circuit board and seamless bilateral communication with consumer electronics (e.g., smartwatch), contextually-relevant (scheduled/on-demand) on-body biomarker data acquisition/display was achieved. | 44076 | 2020
MHBA7EZX | 10.1073/pnas.2009979117 | https://www.ncbi.nlm.nih.gov/pubmed/32719130 | YES | wrist-worn | smartwatch-like | NO | therapeutic drug monitoring | electrochemical | Electrochemical (Voltammetry) | The study mentions that the sweat and saliva readings for acetaminophen, obtained using the device, were cross-validated with laboratory-based assays, achieving an R² value of approximately 0.94. This high correlation coefficient suggests a strong agreement between the device's measurements and established laboratory techniques, supporting the accuracy of the collected data. |  | NA | yes | yes | - | sweat | saliva (as control) | [9 empty] | yes |  | Acetaminophen | Acetaminophen |  |  |  | 2 | 1 | Validation | Human | Human studies | NA | NA | A voltammetric sensing interface, featuring an “undistorted potential window,” within which the target electroactive drug’s voltammetric response is dominant and interference is eliminated, rendering reliable target quantification in noninvasively retrievable biofluids (sweat and saliva) | NO | NA | YES | NA | 2020 | Lin, S.; Yu, W.; Wang, B.; Zhao, Y.; En, K.; Zhu, J.; Cheng, X.; Zhou, C.; Lin, H.; Wang, Z.; Hojaiji, H.; Yeung, C.; Milla, C.; Davis, R. W.; Emaminejad, S. | Noninvasive wearable electroactive pharmaceutical monitoring for personalized therapeutics | Proc Natl Acad Sci U S A |  | 1091-6490 (Electronic) 0027-8424 (Print) 0027-8424 (Linking) | To achieve the mission of personalized medicine, centering on delivering the right drug to the right patient at the right dose, therapeutic drug monitoring solutions are necessary. In that regard, wearable biosensing technologies, capable of tracking drug pharmacokinetics in noninvasively retrievable biofluids (e.g., sweat), play a critical role, because they can be deployed at a large scale to monitor the individuals' drug transcourse profiles (semi)continuously and longitudinally. To this end, voltammetry-based sensing modalities are suitable, as in principle they can detect and quantify electroactive drugs on the basis of the target's redox signature. However, the target's redox signature in complex biofluid matrices can be confounded by the immediate biofouling effects and distorted/buried by the interfering voltammetric responses of endogenous electroactive species. Here, we devise a wearable voltammetric sensor development strategy-centering on engineering the molecule-surface interactions-to simultaneously mitigate biofouling and create an "undistorted potential window" within which the target drug's voltammetric response is dominant and interference is eliminated. To inform its clinical utility, our strategy was adopted to track the temporal profile of circulating acetaminophen (a widely used analgesic and antipyretic) in saliva and sweat, using a surface-modified boron-doped diamond sensing interface (cross-validated with laboratory-based assays, R(2) approximately 0.94). Through integration of the engineered sensing interface within a custom-developed smartwatch, and augmentation with a dedicated analytical framework (for redox peak extraction), we realized a wearable solution to seamlessly render drug readouts with minute-level temporal resolution. Leveraging this solution, we demonstrated the pharmacokinetic correlation and significance of sweat readings. | 44054 | 2020
KGA75EJM | 10.1021/acssensors.8b00528 | https://www.ncbi.nlm.nih.gov/pubmed/30095251 | YES | skin-attached |  | NO | lithium drug monitoring | electrochemical | Electrochemical | The sensitivity of the watch-type sweat rate sensor is stated as 0.039 (pF/s)/(g/m2h). | Neuro | Lithium therapy | yes | - | yes | interstitial fluid | sweat | [11 empty] | Li+ |  |  |  |  | 2 | 1 | prototype | Human | Human studies | NA | NA | Miniaturized Thermo-Pneumatic Actuators. | NO | NA | NA | NA | 2018 | Sweilam, M. N.; Varcoe, J. R.; Crean, C. | Fabrication and Optimization of Fiber-Based Lithium Sensor: A Step toward Wearable Sensors for Lithium Drug Monitoring in Interstitial Fluid | ACS Sens |  | 2379-3694 (Electronic) 2379-3694 (Linking) | A miniaturized, flexible fiber-based lithium sensor was fabricated from low-cost cotton using a simple, repeatable dip-coating technique. This lithium sensor is highly suited for ready-to-use wearable applications and can be used directly without the preconditioning steps normally required with traditional ion-selective electrodes. The sensor has a stable, rapid, and accurate response over a wide Li(+) concentration range that spans over the clinically effective and the toxic concentration limits for lithium in human serum. The sensor is selective to Li(+) in human plasma even in the presence of a high concentration of Na(+) ions. This novel sensor concept represents a significant advance in wearable sensor technology which will target lithium drug monitoring from under the skin. | 43371 | 2018
6T4KEKPL | 10.1016/j.bios.2021.113764 | https://linkinghub.elsevier.com/retrieve/pii/S0956566321008010 | YES | wrist-worn | band | NO | detection and monitoring of phenylalanine (PHE) levels in biofluids | electrochemical | Electrochemical | In Buffer Solution: LOD = 3.0 μM In Diluted Saliva: LOD = 20.9 μM In Whole Saliva (using paper-based sampling): LOD = 39.8 μM In Serum: LOD = 114.2 μM | PKU | phenylketonuria (PKU) management | - | yes | - | saliva | serum | [9 empty] | yes |  | Phenylalanine (PHE) | Phenylalanine (PHE) |  |  |  | 2 | 1 | Prototype/Validation |  | laboratory testing + in-vitro tests | NA | NA | Wearable Electrochemical Detection: successfully miniaturizes and integrates lab-grade electrochemical sensing components, like potentiostats and SPEs, into a compact wearable device, enabling real-time monitoring. Modified Screen-Printed Electrodes: utilizes a Nafion/NQS-modified membrane for phenylalanine detection, simplifying in-situ analysis by eliminating complex sample pre-treatment or enzymatic reactions, improving ease of use and sensor stability. Paper-Based Sampling: Incorporates pre-soaked paper-based elements for direct saliva and serum analysis, bypassing dilution steps and streamlining biofluid collection and analysis. PKU-Specific Application: demonstrates feasibility and practicality for phenylketonuria (PKU) management, emphasizing user-friendly monitoring through an integrated sensing approach. | NO | NA | NO | NA | 2022 | Parrilla, Marc; Vanhooydonck, Andres; Watts, Regan; De Wael, Karolien | Wearable wristband-based electrochemical sensor for the detection of phenylalanine in biofluids | Biosensors and Bioelectronics |  | 9565663 | Wearable electrochemical sensors are driven by the user-friendly capability of on-site detection of key biomarkers for health management. Despite the advances in biomolecule monitoring such as glucose, still, several unmet clinical challenges need to be addressed. For example, patients suffering from phenylketonuria (PKU) should be able to monitor their phenylalanine (PHE) level in a rapid, decentralized, and affordable manner to avoid high levels of PHE in the body which can lead to a profound and irreversible mental disability. Herein, we report a wearable wristband electrochemical sensor for the monitoring of PHE tackling the necessity of controlling PHE levels in PHE hydroxylase deficiency patients. The proposed electrochemical sensor is based on a screen-printed electrode (SPE) modified with a membrane consisting of Nafion, to avoid interferences in biofluids. The membrane also consists of sodium 1,2-naphthoquinone-4-sulphonate for the in situ derivatization of PHE into an electroactive product, allowing its electrochemical oxidation at the surface of the SPE in alkaline conditions. Importantly, the electrochemical sensor is integrated into a wristband configuration to enhance user interaction and engage the patient with PHE self-monitoring. Besides, a paper-based sampling strategy is designed to alkalinize the real sample without the need for sample pretreatment, and thus simplify the analytical process. Finally, the wearable device is tested for the determination of PHE in saliva and blood serum. The proposed wristband-based sensor is expected to impact the PKU self-monitoring, facilitating the daily lives of PKU patients toward optimal therapy and disease management. | 44593 | 2022
IAKUZPEC | 10.1002/smll.202205061 | https://onlinelibrary.wiley.com/doi/10.1002/smll.202205061 | YES | skin-attached |  | NO | real-time, in situ monitoring of multiple biomarkers in epidermal sweat and wound exudate. | electrochemical | Electrochemical | For Electrolyte Sensing (Na+, K+, and pH): The Na+ sensor exhibits a linear relationship between the open circuit potential (OCP) and the logarithm of Na+ concentration from 5 to 320 mM with a slope of 58.11 mV decade−1. The K+ sensor shows a linear relationship between the OCP and the logarithm of K+ concentration from 1 to 64 mM with a slope of 58.34 mV decade−1. The pH sensor demonstrates a linear relationship between the OCP and the pH from 4.5 to 9.0 with a slope of -57.85 mV pH−1. All three sensors exhibit high selectivity over other common ions in biofluids, ensuring accurate measurements of Na+, K+, and pH in raw sweat and wound exudate. For Uric Acid (UA) Sensing: The UA sensor, utilizing highly ordered mesoporous carbons (HOMCs) as the electrode material, achieves a detection limit down to 0.30 µM. The sensor displays a wide linear relationship between the current and the UA content with dynamic ranges of 1–100 µM and 100–900 µM. The UA sensor also demonstrates high selectivity over common analytes in human biofluids. | Hyperuricemia | hyperuricemia management and wound care | yes | - | - | sweat | wound exudate | [6 empty] | yes |  |  | yes |  | Na+, K+, pH, Uric Acid (UA) | Na+, K+, pH, Uric Acid (UA) |  |  |  | 2 | 4 | Validation | Human | laboratory testing + human subject studies + animal models | NA | NA | Flexible and Skin-Conforming Design: incorporates PET and PI films for seamless wearability on diverse body locations, including forehead, forearm, and wounds. Advanced Microfluidic Module: enables efficient sampling, storage, and excretion of sweat and wound exudate, ensuring continuous delivery to the sensing module. Multiplexed Biomarker Sensing: simultaneously detects Na⁺, K⁺, pH, and uric acid (UA), providing a comprehensive analysis of physiological and wound conditions. High-Sensitivity UA Sensor with HOMCs: utilizes highly ordered mesoporous carbons for enhanced electron transfer, significantly improving UA detection sensitivity and selectivity. Biofluid-Specific Calibration: employs Na⁺/K⁺ co-assistance and pH calibration to ensure accurate UA readings for both sweat and wound exudate. Wireless Real-Time Data Transmission: features Bluetooth-enabled signal processing for convenient monitoring and data sharing. | NO | NA | NA | NA | 2022 | Pei, Xinyi; Sun, Mimi; Wang, Jingjuan; Bai, Jing; Bo, Xiangjie; Zhou, Ming | A Bifunctional Fully Integrated Wearable Tracker for Epidermal Sweat and Wound Exudate Multiple Biomarkers Monitoring | Small |  | 1613-6810, 1613-6829 | Abstract Fully integrated wearable electronics that combine the extraordinary feature of incessant and on‐body operation with the distinctive external equipment‐free trait are the ultimate goal of modern wearables. Epidermal sweat and wound exudate, as two noninvasively accessible biofluids on/surrounding the skin, reflect underlying health conditions. However, the design of universal wearable sensors with the bifunctional capability to monitor both epidermal secretions is still a challenge. Here, a single bifunctional fully integrated wearable tracker for wirelessly, simultaneously, and dynamically in situ measuring multiple epidermal sweat or wound exudate biomarkers is propos. Considering the electrolytes (e.g., Na + , K + , and H + ) and metabolites (e.g., uric acid (UA)) levels in sweat or wound exudate may correlate with health or wound conditions, the dynamic and skin‐on tracking of the biomarkers of Na + , K + , pH, and UA levels in sweat under subjects’ exercise and in wound exudate during subjects’ wound healing are performed through the seamless integration of microfluidic, sensing, and electronic modules. Its applicability is evaluated for noninvasive hyperuricemia management in hyperuricemia/healthy subjects through a purine‐rich intake test and for wound management in subjects’ infected wounds through a control medical treatment. | 44866 | 2022
JAXTNRXS | 10.1038/s41551-021-00685-1 | https://www.ncbi.nlm.nih.gov/pubmed/33589782 | YES | patch | skin-attached | NO | monitoring of blood pressure and heart rate | electrochemical | Electrochemical | Glucose: The device successfully captured the physiological effects of food intake, specifically the production of glucose after digestion and its subsequent consumption via glycolysis. This indicates the glucose sensor's sensitivity to changes in glucose levels. Lactate: The device detected increases in lactate production during exercise, reflecting the body's response to oxygen depletion. This demonstrates the lactate sensor's ability to detect changes in lactate levels associated with physiological processes. Blood pressure and heart rate: The device accurately measured increases in blood pressure and heart rate during exercise, which compensate for oxygen depletion and lactate generation. |  | NA | yes | - | yes | interstitial fluid | sweat |  |  |  | yes | yes | [6 empty] | glucose | Lactate, Caffeine, Alcohol |  |  |  | 2 | 4 | Validation | Human | Human studies | NA | NA | The device combines ultrasonic transducers for haemodynamic monitoring (blood pressure and heart rate) with electrochemical sensors for metabolic biomarker detection (glucose, lactate, caffeine, alcohol, pH, and Cl-) | NO | NA | NA | NA | 2021 | Sempionatto, J. R.; Lin, M.; Yin, L.; De la Paz, E.; Pei, K.; Sonsa-Ard, T.; de Loyola Silva, A. N.; Khorshed, A. A.; Zhang, F.; Tostado, N.; Xu, S.; Wang, J. | An epidermal patch for the simultaneous monitoring of haemodynamic and metabolic biomarkers | Nat Biomed Eng |  | 2157-846X (Electronic) 2157-846X (Linking) | Monitoring the effects of daily activities on the physiological responses of the body calls for wearable devices that can simultaneously track metabolic and haemodynamic parameters. Here we describe a non-invasive skin-worn device for the simultaneous monitoring of blood pressure and heart rate via ultrasonic transducers and of multiple biomarkers via electrochemical sensors. We optimized the integrated device so that it provides mechanical resiliency and flexibility while conforming to curved skin surfaces, and to ensure reliable sensing of glucose in interstitial fluid and of lactate, caffeine and alcohol in sweat, without crosstalk between the individual sensors. In human volunteers, the device captured physiological effects of food intake and exercise, in particular the production of glucose after food digestion, the consumption of glucose via glycolysis, and increases in blood pressure and heart rate compensating for oxygen depletion and lactate generation. Continuous and simultaneous acoustic and electrochemical sensing via integrated wearable devices should enrich the understanding of the body's response to daily activities, and could facilitate the early prediction of abnormal physiological changes. | 44378 | 2021
V4DT66WY | 10.1016/j.aca.2022.339518 | https://linkinghub.elsevier.com/retrieve/pii/S0003267022000897 | YES | skin-attached |  | NO | on-site monitoring of nickel levels in biological fluids | electrochemical | Electrochemical | The sensor exhibits a Nernstian response, meaning its potential difference changes linearly with the logarithm of the Ni(II) ion concentration. The reported slope value is 29.5 ± 0.09 mV decade-1. This indicates a highly sensitive response to changes in Ni(II) ion concentration. The sensor demonstrates a linear response over a wide concentration range of 1.0 × 10-5 to 1.0 × 10-1 mol L-1. This wide range makes the sensor suitable for detecting Ni(II) ions in various biological and environmental samples. The sensor achieves a detection limit of 2.7 × 10-6 mol L-1. This value represents the lowest concentration of Ni(II) ions that the sensor can reliably detect. The researchers tested the sensor's accuracy by measuring Ni(II) ion concentrations in spiked samples of human sweat, saliva, and tap water. The sensor achieved excellent recovery values, ranging from 99.56% to 100.9%. The sensor's performance was assessed for repeatability (measurements within one day) and intermediate precision (measurements over multiple days). The results demonstrate high repeatability and precision, with low standard deviation (SD) and relative standard deviation (RSD%) values. |  | NA | yes | yes | - | sweat | saliva | [11 empty] | Nickel (II) ions | Nickel (II) ions |  |  |  | 2 | 1 | Prototype/Validation |  | in-vitro tests + laboratory testing | NA | NA | The flexible Ni(II) ion sensor incorporates several cutting-edge innovations, including: - A novel, flexible, and highly porous substrate (AFCC-PNP) that enhances sensitivity, conductivity, and wearability. - The use of 2D Ni-MOF nanosheets as a highly sensitive and selective electroactive material. - Incorporation of antibacterial properties for safe handling in wearable applications. - A miniaturized design suitable for integration into compact devices. - Rapid response time and good reversibility for real-time and continuous monitoring. | NO | NA | NA | NA | 2022 | Elashery, Sally E.A.; Attia, Nour F.; Oh, Hyunchul | Design and fabrication of novel flexible sensor based on 2D Ni-MOF nanosheets as a preliminary step toward wearable sensor for onsite Ni (II) ions detection in biological and environmental samples | Analytica Chimica Acta |  | 32670 | Herein, for the ﬁrst time, an innovative and sensitive ﬂexible sensor for efﬁcient potentiometric monitoring of Ni (II) ions has been designed and developed. The developed ﬂexible sensor is constructed from highly porous activated ﬂexible carbon cloth decorated with nitrogen and spherical porous carbon nanoparticles derived from low-cost cotton doped with polypyrrole nanoparticles via simple carbonization-activation process followed by dip-coating in membrane cocktail containing 2D Ni-MOF nanosheets as an electroactive material. The developed ﬂexible sensor affords rapid, accurate and stable response for the Ni (II) ions monitoring at its trace level in the biological ﬂuids including human saliva and sweat samples in addition to tap water as an environmental sample without any preconditioning steps over pH range of 2e8 with detection limit of 2.7 10 6. Additionally, the ﬂexible sensor shows good antibacterial properties against both of Gram-positive (Staphylococcus aureus) and Gram-negative (Escherichia coli) bacteria presenting safe handling for human skin as a ﬂexible sensor for wearable system applications. This study presents a forward step for the fabrication of a wearable sensor for rapid onsite monitoring of Ni (II) ions in biological and environmental samples. | 44621 | 2022
MQNLH9ID | 10.1021/acsnano.6b04005 | https://www.ncbi.nlm.nih.gov/pubmed/27380446 | YES | skin-attached |  | NO | monitoring of ionized calcium and pH in body fluids | electrochemical | Electrochemical | The sensor shows fast response to changes in Ca2+ level with a 3.0% relative standard deviation (RSD) of sensitivity. Deviation due to potential drift is 0.7 mV/h which corresponds to 1.1% error in pH value over an hour of continuous measurement. | Metabolic Syndrome | Primary hyperparathyroidism and kidney stones. | yes | - | - | sweat | urine | tears | [8 empty] | yes |  | Ionized Calcium, pH | Ionized Calcium, pH | Ionized Calcium, pH |  |  | 3 | 2 | prototype | Human | Human studies | NA | NA | The device features a disposable and flexible array of sensors specifically designed for detecting calcium and pH. | NO | NA | NA | NA | 2016 | Nyein, H. Y.; Gao, W.; Shahpar, Z.; Emaminejad, S.; Challa, S.; Chen, K.; Fahad, H. M.; Tai, L. C.; Ota, H.; Davis, R. W.; Javey, A. | A Wearable Electrochemical Platform for Noninvasive Simultaneous Monitoring of Ca(2+) and pH | ACS Nano |  | 1936-086X (Electronic) 1936-0851 (Linking) | Homeostasis of ionized calcium in biofluids is critical for human biological functions and organ systems. Measurement of ionized calcium for clinical applications is not easily accessible due to its strict procedures and dependence on pH. pH balance in body fluids greatly affects metabolic reactions and biological transport systems. Here, we demonstrate a wearable electrochemical device for continuous monitoring of ionized calcium and pH of body fluids using a disposable and flexible array of Ca(2+) and pH sensors that interfaces with a flexible printed circuit board. This platform enables real-time quantitative analysis of these sensing elements in body fluids such as sweat, urine, and tears. Accuracy of Ca(2+) concentration and pH measured by the wearable sensors is validated through inductively coupled plasma-mass spectrometry technique and a commercial pH meter, respectively. Our results show that the wearable sensors have high repeatability and selectivity to the target ions. Real-time on-body assessment of sweat is also performed, and our results indicate that calcium concentration increases with decreasing pH. This platform can be used in noninvasive continuous analysis of ionized calcium and pH in body fluids for disease diagnosis such as primary hyperparathyroidism and kidney stones. | 42577 | 2016
U3WC8QC6 | 10.1002/adma.202006444 | https://onlinelibrary.wiley.com/doi/10.1002/adma.202006444 | YES | skin-attached |  | NO | monitoring vitamin C levels in biofluids | electrochemical | Electrochemical | Sensitivity Variation: the researchers tested five different sensors to assess reproducibility. They found a sensitivity variation of approximately 12% for vitamin C concentrations ranging from 0 to 5000 µm. This suggests that different sensors might produce slightly different readings for the same concentration of vitamin C. Calibration for Accuracy: the authors acknowledge that the sensitivity variation is smaller than typical vitamin C concentration changes throughout the day. However, they suggest that calibrating the sensor's signal before measurement can improve accuracy. This implies that individual sensors can be adjusted to minimize variations and enhance the reliability of the data. Long-term Stability: he sensor was tested for 48 hours to evaluate its long-term performance. The measured current recovered consistently when the vitamin C concentration was switched between 100 and 200 µm. This indicates that the sensor maintains its stability and ability to accurately measure vitamin C levels over extended periods. Storage Stability: Further testing showed that the sensitivity variation of a single sensor over five days was less than 2%. This demonstrates the sensor's reliability and ability to provide consistent measurements even after storage. |  | NA | yes | - | - | sweat | urine | blood | [10 empty] | Vitamin C | Vitamin C | Vitamin C |  |  | 3 | 1 | Prototype | Human | laboratory testing + in-vitro tests + human studies (preliminary). | NA | NA | Gold Nanodendrites: enhance sensitivity by increasing the electrode's surface area, improving electron transfer. Electroactive Conducting Polymer (PEDOT:LiClO4): improves electron transfer and electrochemical performance for better sensor functionality. Enzyme-Immobilized Membrane: ensures selectivity by using l-ascorbate oxidase to specifically detect vitamin C. Nafion Encapsulation: provides stability, preventing enzyme leakage and fouling in biofluids. Flexible PET Substrate: enhances wearability and comfort for continuous, conformal skin contact. | NO | NA | NA | NA | 2021 | Zhao, Jiangqi; Nyein, Hnin Yin Yin; Hou, Lei; Lin, Yuanjing; Bariya, Mallika; Ahn, Christine Heera; Ji, Wenbo; Fan, Zhiyong; Javey, Ali | A Wearable Nutrition Tracker | Advanced Materials |  | 0935-9648, 1521-4095 | Abstract Nutrients are essential for the healthy development and proper maintenance of body functions in humans. For adequate nourishment, it is important to keep track of nutrients level in the body, apart from consuming sufficient nutrition that is in line with dietary guidelines. Sweat, which contains rich chemical information, is an attractive biofluid for routine non‐invasive assessment of nutrient levels. Herein, a wearable sensor that can selectively measure vitamin C concentration in biofluids, including sweat, urine, and blood is developed. Detection through an electrochemical sensor modified with Au nanostructures, LiClO 4 ‐doped conductive polymer, and an enzymes‐immobilized membrane is utilized to achieve wide detection linearity, high selectivity, and long‐term stability. The sensor allows monitoring of temporal changes in vitamin C levels. The effect of vitamin C intake on the sweat and urine profile is explored by monitoring concentration changes upon consuming different amounts of vitamin C. A longitudinal study of sweat's and urine's vitamin C correlation with blood is performed on two individuals. The results suggest that sweat and urine analysis can be a promising method to routinely monitor nutrition through the sweat sensor and that this sensor can facilitate applications such as nutritional screening and dietary intervention. | 44197 | 2021
8UP6AP3D | 10.1016/j.bios.2019.03.053 | https://linkinghub.elsevier.com/retrieve/pii/S0956566319302696 | YES | skin-attached |  | NO | detection of cytokine levels | electrochemical | Electrochemical | IFN-γ: LOD of 476 × 10⁻¹⁵ m; IL-6: LOD of 608 × 10⁻¹⁵ m; TNF-α: LOD of 611 × 10⁻¹⁵ m. The results indicate high accuracy for all three cytokines across different biofluids: IFN-γ: Accuracy consistently above 95.48% in serum, saliva, and urine; TNF-α: Accuracy exceeding 95.2% in serum, saliva, and urine; IL-6: Accuracy greater than 95.8% in serum, saliva, and urine. The electrical properties of the flexible sensor remain consistent even when subjected to significant bending and twisting. The transconductance, a measure of the sensor's sensitivity, shows minimal variation under these conditions, with deviations less than 5%. The flexible sensor maintains a high level of consistency in measuring cytokine levels even when bent or twisted. For TNF-α detection, deviations are less than 8.44% across different concentrations and deformation states. The sources demonstrate the feasibility of using the flexible device for real-time monitoring of cytokines in human sweat. They show that the sensor can effectively detect and track changes in IFN-γ, TNF-α, and IL-6 concentrations in sweat samples. | Inflammation | cytokine storm syndrome (CSS) management | yes | yes | - | serum | saliva | urine | sweat | [9 empty] | IFN-γ, TNF-α, IL-6 | IFN-γ, TNF-α, IL-6 | IFN-γ, TNF-α, IL-6 | IFN-γ, TNF-α, IL-6 |  | 4 | 3 | Validation |  | in-vitro tests | NA | NA | The aptameric DGTFET biosensing device features a dual-channel design to reduce interference by comparing a sensing channel (functionalized with aptamers) to a reference channel, enhancing accuracy in complex biofluids. TWEEN 80 passivation minimizes non-specific biomolecule adsorption, improving specificity. Aptamers for cytokines like IFN-γ, TNF-α, and IL-6 ensure highly sensitive and selective detection through measurable electrical shifts. The sensor is fabricated on flexible PET substrates for wearable use, maintaining performance under bending conditions. Integrated wireless transmission and a mobile app enable real-time cytokine monitoring, supporting telemedicine and wearable healthcare applications. | NO | NA | YES | NA | 2019 | Hao, Zhuang; Pan, Yunlu; Shao, Wenwen; Lin, Qiao; Zhao, Xuezeng | Graphene-based fully integrated portable nanosensing system for on-line detection of cytokine biomarkers in saliva | Biosensors and Bioelectronics |  | 9565663 | Saliva has been reported to contain various cytokine biomarkers which are associated with some severe diseases such as cancers. Non-invasive saliva diagnosis using wearable or portable devices may pave a new avenue for monitoring conditions of the high risk population. Here, a graphene-based fully integrated portable nanosensing system, the entire size of which is smaller than a smart-phone and can be handheld, is presented for on-line detection of cytokine biomarkers in saliva. This miniaturized system employs an aptameric graphene-based ﬁeld eﬀect transistor (GFET) using a buried-gate geometry with HfO2 as the dielectric layer and on-line signal processing circuits to realize the transduction and processing of signals which reﬂect cytokine concentrations. The signal can be wirelessly transmitted to a smart-phone or cloud sever through the Wi-Fi connection for visualizing the trend of the cytokine concentration change. Interleukin-6 (IL-6) is used as a representative to examine the sensing capability of the system. Experimental results demonstrate that the nanosensing system responds to the change of IL-6 concentration within 400s in saliva with a detection limit down to 12 pM. Therefore, this portable system oﬀers the practicality to be potentially used for non-invasive saliva diagnosis of diseases at early stage. | 43617 | 2019
LMJIY823 | 10.1039/d0an00846j | https://www.ncbi.nlm.nih.gov/pubmed/32500895 | YES | only for sweat from skin surface |  | NO | detection of psychoactive substances within biological fluids | Other | Electrochemiluminescence (ECL) | The sensor demonstrated a Limit of Detection (LOD) of 0.1 µM in human pooled serum. In artificial saliva, the LOD was approximately 1 µM. In artificial sweat, a concentration of 1 µM scopolamine yielded a signal 3.4 times greater than the blank matrix signal, allowing for clear differentiation. The sensor was also tested by collecting sweat from synthetic skin spiked with scopolamine, resulting in a signal intensity 47 times greater than blank sweat. |  | NA | yes | yes | - | Human pooled serum | Urine | artificial saliva | sweat | sweat collected from the skin's surface | [8 empty] | scopolamine | scopolamine | scopolamine | scopolamine |  | 5 | 1 | Proof-of-concept | Human | Human studies | NA | NA | The core technology of the sensor relies on electrochemiluminescence, a process where light emission is triggered by an electrochemical reaction. While ECL itself is not a brand-new concept, its application in sensor development for psychoactive substances in biological matrices represents a potentially novel approach. | NO | NA | NA | NA | 2020 | Brown, K.; Jacquet, C.; Biscay, J.; Allan, P.; Dennany, L. | Electrochemiluminescent sensors as a screening strategy for psychoactive substances within biological matrices | Analyst |  | 1364-5528 (Electronic) 0003-2654 (Linking) | With the rapid growth and appearance of novel psychoactive substances (NPS) onto the global drug market, the need for alternative screening methodologies for implementation within clinical environments is substantial. The immunoassay methods currently in use are inadequate for this new drug trend with the potential for misdiagnosis and subsequent administration of incorrect patient treatment increased. This contribution illustrates a strong proof-of-concept for the use of electrochemiluminescence (ECL) as a screening methodology for NPS within biological fluids, using the hallucinogen scopolamine as a model compound. A low cost, easy-to-use and portable sensor has been developed and successfully employed for the detection of scopolamine at clinically relevant concentrations within a variety of biological matrices, including human pooled serum, urine, artificial saliva and sweat, without any prior sample preparation required. Moreover, assessment of the sensor's potential as a point-of-care wearable device was performed with sample collection from the surface of skin, demonstrating its capability for the qualitative identification of scopolamine despite collection of only minimal volumes off the skins surface. The developed sensor described herein exhibits a strong proof-of-concept for the employment of such ECL sensors as point-of-care devices, where the sensors ease of use and removal of time-consuming and complex sample preparation methods will ultimately increase its usability by physicians, widening the avenues where ECL sensors could be employed. | 43997 | 2020
3XC7JRTS | 10.3390/s20071904 | http://www.ncbi.nlm.nih.gov/pubmed/32235429 | YES | textile |  | NO | Measuring relative humidity (RH) changes in real-time during sports activities (NO MOLECULAR/IONIC BIOMARKERS) | Optical | Optical |  |  |  | - | - | - | [30 empty] | 2020 | He, Chenyang; Korposh, Serhiy; Hernandez, Francisco Ulises; Liu, Liangliang; Correia, Ricardo; Hayes-Gill, Barrie R.; Morgan, Stephen P. | Real-Time Humidity Measurement during Sports Activity using Optical Fibre Sensing | Sensors (Basel, Switzerland) |  | 1424-8220 | An optical fibre sensor for monitoring relative humidity (RH) changes during exercise is demonstrated. The humidity sensor comprises a tip coating of poly (allylamine hydrochloride) (PAH)/silica nanoparticles (SiO2 NPs) deposited using the layer-by-layer technique. An uncoated fibre is employed to compensate for bending losses that are likely to occur during movement. A linear fit to the response of the sensing system to RH demonstrates a sensitivity of 3.02 mV/% (R2 = 0.96), hysteresis ± 1.17% RH when 11 bilayers of PAH/SiO2 NPs are coated on the tip of the fibre. The performance of two different textiles (100% cotton and 100% polyester) were tested in real-time relative humidity measurement for 10 healthy volunteers. The results demonstrate the moisture wicking properties of polyester in that the relative humidity dropped more rapidly after cessation of exercise compared to cotton. The approach has the potential to be used to monitor sports performance and by clothing developers for characterising different garment designs. | 43920 | 2020
D3ZYCFEH | 10.2196/17326 | https://www.ncbi.nlm.nih.gov/pubmed/32432552 | YES | ECG device |  | NO | CARDIOVASCULAR MONITORING (NO MOLECULAR/IONIC BIOMARKERS) | Electrical | Electrode-Based Sensing (ECG technology) |  |  |  | - | - | - | [30 empty] | 2020 | De Canniere, H.; Smeets, C. J. P.; Schoutteten, M.; Varon, C.; Van Hoof, C.; Van Huffel, S.; Groenendaal, W.; Vandervoort, P. | Using Biosensors and Digital Biomarkers to Assess Response to Cardiac Rehabilitation: Observational Study | J Med Internet Res |  | 1438-8871 (Electronic) 1439-4456 (Print) 1438-8871 (Linking) | BACKGROUND: Cardiac rehabilitation (CR) is known for its beneficial effects on functional capacity and is a key component within current cardiovascular disease management strategies. In addition, a larger increase in functional capacity is accompanied by better clinical outcomes. However, not all patients respond in a similar way to CR. Therefore, a patient-tailored approach to CR could open up the possibility to achieve an optimal increase in functional capacity in every patient. Before treatment can be optimized, the differences in response of patients in terms of cardiac adaptation to exercise should first be understood. In addition, digital biomarkers to steer CR need to be identified. OBJECTIVE: The aim of the study was to investigate the difference in cardiac response between patients characterized by a clear improvement in functional capacity and patients showing only a minor improvement following CR therapy. METHODS: A total of 129 patients in CR performed a 6-minute walking test (6MWT) at baseline and during four consecutive short-term follow-up tests while being equipped with a wearable electrocardiogram (ECG) device. The 6MWTs were used to evaluate functional capacity. Patients were divided into high- and low-response groups, based on the improvement in functional capacity during the CR program. Commonly used heart rate parameters and cardiac digital biomarkers representative of the heart rate behavior during the 6MWT and their evolution over time were investigated. RESULTS: All participating patients improved in functional capacity throughout the CR program (P<.001). The heart rate parameters, which are commonly used in practice, evolved differently for both groups throughout CR. The peak heart rate (HR(peak)) from patients in the high-response group increased significantly throughout CR, while no change was observed in the low-response group (F(4,92)=8.321, P<.001). Similar results were obtained for the recovery heart rate (HR(rec)) values, which increased significantly over time during every minute of recuperation, for the high-response group (HR(rec1): P<.001, HR(rec2): P<.001, HR(rec3): P<.001, HR(rec4): P<.001, and HR(rec5): P=.02). The other digital biomarkers showed that the evolution of heart rate behavior during a standardized activity test differed throughout CR between both groups. These digital biomarkers, derived from the continuous measurements, contribute to more in-depth insight into the progression of patients' cardiac responses. CONCLUSIONS: This study showed that when using wearable sensor technology, the differences in response of patients to CR can be characterized by means of commonly used heart rate parameters and digital biomarkers that are representative of cardiac response to exercise. These digital biomarkers, derived by innovative analysis techniques, allow for more in-depth insights into the cardiac response of cardiac patients during standardized activity. These results open up the possibility to optimized and more patient-tailored treatment strategies and to potentially improve CR outcome. | 43971 | 2020
DLES74LR | 10.1038/s41598-018-19239-8 | https://www.ncbi.nlm.nih.gov/pubmed/29352237 | YES | wrist-worn |  | NO | monitor human thermal comfort (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2018 | Sim, J. K.; Yoon, S.; Cho, Y. H. | Wearable Sweat Rate Sensors for Human Thermal Comfort Monitoring | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | We propose watch-type sweat rate sensors capable of automatic natural ventilation by integrating miniaturized thermo-pneumatic actuators, and experimentally verify their performances and applicability. Previous sensors using natural ventilation require manual ventilation process or high-power bulky thermo-pneumatic actuators to lift sweat rate detection chambers above skin for continuous measurement. The proposed watch-type sweat rate sensors reduce operation power by minimizing expansion fluid volume to 0.4 ml through heat circuit modeling. The proposed sensors reduce operation power to 12.8% and weight to 47.6% compared to previous portable sensors, operating for 4 hours at 6 V batteries. Human experiment for thermal comfort monitoring is performed by using the proposed sensors having sensitivity of 0.039 (pF/s)/(g/m(2)h) and linearity of 97.9% in human sweat rate range. Average sweat rate difference for each thermal status measured in three subjects shows (32.06 +/- 27.19) g/m(2)h in thermal statuses including 'comfortable', 'slightly warm', 'warm', and 'hot'. The proposed sensors thereby can discriminate and compare four stages of thermal status. Sweat rate measurement error of the proposed sensors is less than 10% under air velocity of 1.5 m/s corresponding to human walking speed. The proposed sensors are applicable for wearable and portable use, having potentials for daily thermal comfort monitoring applications. | 43119 | 2018
I39HDFH6 | 10.3390/molecules26123570 | http://www.ncbi.nlm.nih.gov/pubmed/34208039 | YES | textile |  | NO | Personal thermal management (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2021 | Gorji, Mohsen; Mazinani, Saeedeh; Faramarzi, Abdol-Rahim; Ghadimi, Saeedeh; Kalaee, Mohammadreza; Sadeghianmaryan, Ali; Wilson, Lee D. | Coating Cellulosic Material with Ag Nanowires to Fabricate Wearable IR-Reflective Device for Personal Thermal Management: The Role of Coating Method and Loading Level | Molecules (Basel, Switzerland) |  | 1420-3049 | Textiles coated with silver nanowires (AgNWs) are effective at suppressing radiative heat loss without sacrificing breathability. Many reports present the applicability of AgNWs as IR-reflective wearable textiles, where such studies partially evaluate the parameters for practical usage for large-scale production. In this study, the effect of the two industrial coating methods and the loading value of AgNWs on the performance of AgNWs-coated fabric (AgNWs-CF) is reported. The AgNWs were synthesized by the polyol process and applied onto the surface of cotton fabric using either dip- or spray-coating methods with variable loading levels of AgNWs. X-ray diffraction, scanning electron microscopy (SEM), infrared (IR) reflectance, water vapor permeability (WVP), and electrical resistance properties were characterized. The results report the successful synthesis of AgNWs with a 30 μm length. The results also show that the spray coating method has a better performance for reflecting the IR radiation to the body, which increases with a greater loading level of the AgNWs. The antibacterial results show a good inhibition zone for cotton fabric coated by both methods, where the spray-coated fabric has a better performance overall. The results also show the coated fabric with AgNWs maintains the level of fabric breathability similar to control samples. AgNWs-CFs have potential utility for cold weather protective clothing in which heat dissipation is attenuated, along with applications such as wound dressing materials that provide antibacterial protection. | 44358 | 2021
E7DJX78A | 10.1038/s41467-018-06759-0 | http://www.ncbi.nlm.nih.gov/pubmed/30323200 | YES | textile | skin-touch-actuated | NO | HARVESTING MECHANICAL ENERGY (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2018 | Xiong, Jiaqing; Cui, Peng; Chen, Xiaoliang; Wang, Jiangxin; Parida, Kaushik; Lin, Meng-Fang; Lee, Pooi See | Skin-touch-actuated textile-based triboelectric nanogenerator with black phosphorus for durable biomechanical energy harvesting | Nature Communications |  | 2041-1723 | Textiles that are capable of harvesting biomechanical energy via triboelectric effects are of interest for self-powered wearable electronics. Fabrication of conformable and durable textiles with high triboelectric outputs remains challenging. Here we propose a washable skin-touch-actuated textile-based triboelectric nanogenerator for harvesting mechanical energy from both voluntary and involuntary body motions. Black phosphorus encapsulated with hydrophobic cellulose oleoyl ester nanoparticles serves as a synergetic electron-trapping coating, rendering a textile nanogenerator with long-term reliability and high triboelectricity regardless of various extreme deformations, severe washing, and extended environmental exposure. Considerably high output (~250-880 V, ~0.48-1.1 µA cm-2) can be attained upon touching by hand with a small force (~5 N) and low frequency (~4 Hz), which can power light-emitting diodes and a digital watch. This conformable all-textile-nanogenerator is incorporable onto cloths/skin to capture the low output of 60 V from subtle involuntary friction with skin, well suited for users' motion or daily operations. | 43388 | 2018
P3PIP8LX | 10.1016/j.jcis.2022.10.129 | https://linkinghub.elsevier.com/retrieve/pii/S0021979722018999 | YES | skin-attached |  | NO | triboelectric nanogenerator (TENG) for energy harvesting from human motion (NO MOLECULAR/IONIC BIOMARKERS) | electrochemical | Electrochemical |  |  | NA | yes | - | - | sweat | [22 empty] | NA | NA |  | NO | NA | YES | NA | 2023 | Cheng, Kuan; Huang, Zixu; Wang, Pengcheng; Sun, Li; Ghasemi, Hadi; Ardebili, Haleh; Karim, Alamgir | Antibacterial flexible triboelectric nanogenerator via capillary force lithography | Journal of Colloid and Interface Science |  | 219797 | Increasing the surface contact area (SCA) between active materials of triboelectric nanogenerator (TENG) can effectively enhance energy harvesting performance. In the meantime, owing to potential intimate contacts between skin and wearables, hygienic aspects of TENG devices with high SCA become important. Herein, we present a robust route to fabricate ﬂexible TENGs with antimicrobial capability realized by nano-patterning thermoplastic polyurethane (TPU) thin ﬁlms. Without involving sophisticated equipment and time-consuming technique, simply tuning curing conditions of capillary force lithography could precisely control the nanostructure geometry. Our topographically designed TENGs could promote higher power generation while preventing bioﬁlm formation without adding any chemical additives. Analysis of pattern amplitude and wavelength correlation to output power is uniquely provided for a deeper understanding of how patterns enable peak performance of TENGs. Furthermore, a prototype TENG was innovatively applied to a smart device as a transparent screen sensor. Speciﬁc ﬁnger trajectories could be recognized by different electric responses, and as-generated electricity can be directly charged into comercial capacitors. Our proposed TENGs with critical hygienic advancements are expected to open avenues for TENG-incorporated wearables in the post-pandemic era. | 44927 | 2023
E4SGQ7AY | 10.1109/JTEHM.2022.3175361 | https://www.ncbi.nlm.nih.gov/pubmed/35711336 | YES | skin-attached |  | NO | HEART-RATE EVALUATION (NO MOLECULAR/IONIC BIOMARKERS) |  |  |  |  | NA | - | - | - | [23 empty] | NA | NA |  | NO | NA | NA | NA | 2022 | Alharbi, E. A.; Jones, J. M.; Alomainy, A. | Non-Invasive Solutions to Identify Distinctions Between Healthy and Mild Cognitive Impairments Participants | IEEE J Transl Eng Health Med |  | 2168-2372 (Electronic) 2168-2372 (Linking) | Mild cognitive impairment (MCI) is a condition characterized by impairment in a single cognitive domain or mild deficit in several cognitive domains. MCI patients are at increased risk of progression to dementia with almost 50% of MCI patients developing dementia within five years. Early detection can play an important role in early intervention, prevention, and appropriate treatments. In this study, we examined heart rate variability (HRV) as a novel physiological biomarker for identifying individuals at higher risk of MCI. We investigated if measuring HRV using non-invasive sensors might offer reliable, non-invasive techniques to distinguish MCI patients from healthy controls. Twenty-one MCI patients were recruited to examine this possibility. HRV was assessed using CorSense wearable device. HRV indices were analyzed and compared in rest between MCI and healthy controls. The significance of difference of numerical data between two groups was assessed using parametric unpaired t-test or non-parametric Wilcoxon rank sum test based on the fulfilment of unpaired t-test assumptions. Multiple linear regression models were performed to assess the association between individual HRV parameter with the cognitive status adjusting for gender and age. Time-domain parameters i.e., the standard deviation of NN intervals (SDNN), and the root mean square of successive differences between normal heartbeats (RMSSD) were significantly lower in MCI patients compared with healthy controls. Prediction accuracy for the logistic regression using 10-fold cross-validation was 76.5%, Specificity was 0.8571, while sensitivity was 0.8095. Our study demonstrated that healthy participants have higher HRV indices compared to older adults with MCI using non-invasive biosensors technologies. Our results are of clinical importance in terms of showing the possibility that MCI of older people can be predicted using only HRV PPG-based data. | 2022 | 2022
FLFTAIIX | 10.2196/60925 | https://formative.jmir.org/2024/1/e60925 | YES | wrist-worn |  | NO | Monitoring nocturnal heart rate and sleep. The device discussed in the sources does not analyze any fluids to detect biomarkers. This device uses photoplethysmography (PPG) to track metrics like heart rate and sleep. The study also investigated the use of saliva and sweat to detect biomarkers, but these fluids were analyzed independently and not using the ŌURA ring. | [5 empty] | - | - | - | [30 empty] | 2024 | Pojednic, Rachele; Welch, Amy; Thornton, Margaret; Garvey, Meghan; Grogan, Tara; Roberts, Walter; Ash, Garrett | Intensive, Real-Time Data Collection of Psychological and Physiological Stress During a 96-Hour Field Training Exercise at a Senior Military College: Feasibility and Acceptability Cohort Study | JMIR Formative Research |  | 2561-326X | Background: Poor physical fitness, stress, and fatigue are factors impacting military readiness, national security, and economic burden for the United States Department of Defense. Improved accuracy of wearable biosensors and remote field biologic sample collection strategies could make critical contributions to understanding how physical readiness and occupational stressors result in on-the-job and environment-related injury, sleep impairments, diagnosis of mental health disorders, and reductions in performance in war-fighters. | 45583 | 2024
GFEL8UXX | 10.1038/s41598-023-40339-7 | https://www.ncbi.nlm.nih.gov/pubmed/37591885 | YES | wrist-worn | smartwatch-like | NO | monitoring of human skin hydration and body sweat loss (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2023 | Volkova, E.; Perchik, A.; Pavlov, K.; Nikolaev, E.; Ayuev, A.; Park, J.; Chang, N.; Lee, W.; Kim, J. Y.; Doronin, A.; Vilenskii, M. | Multispectral sensor fusion in SmartWatch for in situ continuous monitoring of human skin hydration and body sweat loss | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | Post-pandemic health operations have become a near-term reality, discussions around wearables are on the rise. How do wearable health solutions effectively deploy and use this opportunity to fill the gap between wellness and healthcare? In this paper, we will talk about wearable healthcare diagnosis, with a particular focus on monitoring skin hydration using optical multi-wavelength sensor fusion. Continuous monitoring of human skin hydration is a task of paramount importance for maintaining water loss dynamics for fitness lovers as well as for skin beauty, integrity and the health of the entire body. Preserving the appropriate levels of hydration ensures consistency of weight, positively affects psychological state, and proven to result in a decrease in blood pressure as well as the levels of "bad" cholesterol while slowing down the aging processes. Traditional methods for determining the state of water content in the skin do not allow continuous and non-invasive monitoring, which is required for variety of consumer, clinical and cosmetic applications. We present novel sensing technology and a pipeline for capturing, modeling and analysis of the skin hydration phenomena and associated changes therein. By expanding sensing capabilities built into the SmartWatch sensor and combining them with advanced modeling and Machine Learning (ML) algorithms, we identified several important characteristics of photoplethysmography (PPG) signal and spectral sensitivity corresponding to dynamics of skin water content. In a hardware aspect, we newly propose the expansion of SmartWatch capabilities with InfraRed light sources equipped with wavelengths of 970 nm and 1450 nm. Evaluation of the accuracy and characteristics of PPG sensors has been performed with biomedical optics-based simulation framework using Monte Carlo simulations. We performed rigorous validation of the developed technology using experimental and clinical studies. The developed pipeline serves as a tool in the ongoing studies of the next generation of optical sensing technology. | 45155 | 2023
8P7PUXU6 | 10.3390/s21144934 | http://www.ncbi.nlm.nih.gov/pubmed/34300678 | YES | textile |  | NO | SMART SENSING TECHNOLOGY (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2021 | Ibanez Labiano, Isidoro; Arslan, Dilan; Ozden Yenigun, Elif; Asadi, Amir; Cebeci, Hulya; Alomainy, Akram | Screen Printing Carbon Nanotubes Textiles Antennas for Smart Wearables | Sensors (Basel, Switzerland) |  | 1424-8220 | Electronic textiles have become a dynamic research field in recent decades, attracting attention to smart wearables to develop and integrate electronic devices onto clothing. Combining traditional screen-printing techniques with novel nanocarbon-based inks offers seamless integration of flexible and conformal antenna patterns onto fabric substrates with a minimum weight penalty and haptic disruption. In this study, two different fabric-based antenna designs called PICA and LOOP were fabricated through a scalable screen-printing process by tuning the conductive ink formulations accompanied by cellulose nanocrystals. The printing process was controlled and monitored by revealing the relationship between the textiles' nature and conducting nano-ink. The fabric prototypes were tested in dynamic environments mimicking complex real-life situations, such as being in proximity to a human body, and being affected by wrinkling, bending, and fabric care such as washing or ironing. Both computational and experimental on-and-off-body antenna gain results acknowledged the potential of tunable material systems complimenting traditional printing techniques for smart sensing technology as a plausible pathway for future wearables. | 44397 | 2021
3FBMHT9N | 10.1007/s10916-021-01709-8 | http://www.ncbi.nlm.nih.gov/pubmed/33608815 | YES | textile |  | NO | MONITORING HUMAN BODY MOVEMENT (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2021 | Kang, Da-Hye; Lee, Joo-Hyeon; Lee, Jeong-Whan; Cho, Hyun-Seung; Park, Seon-Hyung; Lee, Kang-Hwi; Kang, Seung-Jin | Conditions for CNT - Coated Textile Sensors Applied to Wearable Platforms to Monitor Limb Joint Motion | Journal of Medical Systems |  | 1573-689X | Despite recent research on joint motion measurement to monitor human body movement, current measurement techniques and tools have significant limitations, including requiring large space for measurement and causing discomfort in test subjects wearing motion sensors. Our study aims, first, to develop carbon nanotube (CNT)-based textile joint motion sensors. Second, ours study aims to identify the most suitable CNT-based sensor structure and attachment method for use on a wearable platform during general exercise speeds. Lastly, we used these sensors on the human body, using sleeves and legs to find the most stable location, and we used the CNT-based sensor condition to monitor joint motions. We utilized our CNT-based sensor, which has proper elasticity as well as conductivity, and applied it to the elbow and knee joints. Based on the strain gauge principle, we monitored the variance of electric resistance that occurred when the CNT-based sensor was stretched due to limb motion. Our study tested 48 types of sensors. These sensors were applied to the CNT using different base knit textiles as well as different attachment methods, layers, sensor lengths, and sensor widths. The four most successful sensor types, which showed superior efficacy over the others in joint motion measurement, were selected for further study. These four sensors were then used to measure the elbow and knee joint motions of human subjects by placing them on different locations on sleeves and legs. The CNT knit textile sensors best suited to measuring joint motions are those with a double-layered CNT knit and 5 cm long × 0.5 cm or 1 cm wide sensors attached to a polyester¬-based knit using a welding method. The best position for the sensor to more stably monitor joint motions was the "below hinge position" from the elbow or knee hinge joint. Our study suggests an alternative strategy for joint-motion measurement that could contribute to the development of more comfortable and human-friendly methods of human limb motion measurement. | 44246 | 2021
98HS67W5 | 10.3390/s21124038 | http://www.ncbi.nlm.nih.gov/pubmed/34208280 | YES | textile |  | NO | TEXTILE CHARACTERIZATION (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2021 | Tabaczyńska, Agnieszka; Dąbrowska, Anna; Słoma, Marcin | Printed Graphene, Nanotubes and Silver Electrodes Comparison for Textile and Structural Electronics Applications | Sensors (Basel, Switzerland) |  | 1424-8220 | Due to the appearance of smart textiles and wearable electronics, the need for electro-conductive textiles and electro-conductive paths on textiles has become clear. In this article the results of a test of developed textile electro-conductive paths obtained by applying the method of screen printing pastes containing silver nanoparticles and carbon (graphene, nanotubes, graphite) are presented. Conducted research included analysis of the adhesion test, as well as evaluation of the surface resistance before and after the washing and bending cycles. Obtained results indicated that the samples with the content of carbon nanotubes 3% by weight in PMMA on substrate made of aramid fibers (surface mass of 260 g/m2) were characterized by the best adhesion and the best resistance to washing and bending cycles. Such electro-conductive paths have potential to be used in smart clothing applications. | 44358 | 2021
GCB7UWKY | 10.1038/s41467-021-25075-8 | https://www.ncbi.nlm.nih.gov/pubmed/34385436 | YES | textile |  | NO | TEXTILE CHARACTERIZATION (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2021 | Yang, Y.; Wei, X.; Zhang, N.; Zheng, J.; Chen, X.; Wen, Q.; Luo, X.; Lee, C. Y.; Liu, X.; Zhang, X.; Chen, J.; Tao, C.; Zhang, W.; Fan, X. | A non-printed integrated-circuit textile for wireless theranostics | Nat Commun |  | 2041-1723 (Electronic) 2041-1723 (Linking) | While the printed circuit board (PCB) has been widely considered as the building block of integrated electronics, the world is switching to pursue new ways of merging integrated electronic circuits with textiles to create flexible and wearable devices. Herein, as an alternative for PCB, we described a non-printed integrated-circuit textile (NIT) for biomedical and theranostic application via a weaving method. All the devices are built as fibers or interlaced nodes and woven into a deformable textile integrated circuit. Built on an electrochemical gating principle, the fiber-woven-type transistors exhibit superior bending or stretching robustness, and were woven as a textile logical computing module to distinguish different emergencies. A fiber-type sweat sensor was woven with strain and light sensors fibers for simultaneously monitoring body health and the environment. With a photo-rechargeable energy textile based on a detailed power consumption analysis, the woven circuit textile is completely self-powered and capable of both wireless biomedical monitoring and early warning. The NIT could be used as a 24/7 private AI "nurse" for routine healthcare, diabetes monitoring, or emergencies such as hypoglycemia, metabolic alkalosis, and even COVID-19 patient care, a potential future on-body AI hardware and possibly a forerunner to fabric-like computers. | 44420 | 2021
8ANCTSZH | 10.1039/d2lc00670g | https://www.ncbi.nlm.nih.gov/pubmed/36268642 | YES | skin-attached |  | NO | autonomous and continuous monitoring of sweat rate (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2022 | Lin, H.; Yu, W.; Suarez, J. E. D.; Athavan, H.; Wang, Y.; Yeung, C.; Lin, S.; Sankararaman, S.; Milla, C.; Emaminejad, S. | Autonomous wearable sweat rate monitoring based on digitized microbubble detection | Lab Chip |  | 1473-0189 (Electronic) 1473-0197 (Print) 1473-0189 (Linking) | Advancements in wearable bioanalytical microsystems have enabled diurnal and (semi)continuous monitoring of physiologically-relevant indices that are accessible through probing sweat. To deliver an undistorted and physiologically-meaningful interpretation of these readings, tracking the sweat secretion rate is essential, because it allows for calibrating the biomarker readings against variations in sweat secretion and inferring the body's hydration/electrolyte homeostasis status. To realize an autonomous wearable solution with intrinsically high signal-to-noise ratio sweat rate sensing capabilities, here, we devise a digitized microbubble detection mechanism-delivered by a hybrid microfluidic/electronic system with a compact footprint. This mechanism is based on the intermittent generation of microliter-scale bubbles via electrolysis and the instantaneous measurement of their time-of-flight (and thus, velocity) via impedimetric sensing. In this way, we overcome the limitations of previously proposed sweat rate sensing modalities that are inherently susceptible to non-targeted secretion characteristics (pH, conductivity, and temperature), constrained by volume, or lack system integration for autonomous on-body operation. By deploying our solution in human subject trials, we validate the utility of our solution for seamless monitoring of exercise- and iontophoretically-induced sweat secretion profiles. | 44873 | 2022
CZKK3ZHM | 10.1038/s41598-022-18271-z | https://www.ncbi.nlm.nih.gov/pubmed/36064877 | YES | skin-attached |  | NO | A biosensor to measure physiological signals related to epileptic seizures. The device monitors electrodermal activity (EDA), peripheral body temperature (TEMP), and heart rate (HR) (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2022 | Vieluf, S.; El Atrache, R.; Cantley, S.; Jackson, M.; Clark, J.; Sheehan, T.; Bosl, W. J.; Zhang, B.; Loddenkemper, T. | Seizure-related differences in biosignal 24-h modulation patterns | Sci Rep |  | 2045-2322 (Electronic) 2045-2322 (Linking) | A seizure likelihood biomarker could improve seizure monitoring and facilitate adjustment of treatments based on seizure risk. Here, we tested differences in patient-specific 24-h-modulation patterns of electrodermal activity (EDA), peripheral body temperature (TEMP), and heart rate (HR) between patients with and without seizures. We enrolled patients who underwent continuous video-EEG monitoring at Boston Children's Hospital to wear a biosensor. We divided patients into two groups: those with no seizures and those with at least one seizure during the recording period. We assessed the 24-h modulation level and amplitude of EDA, TEMP, and HR. We performed machine learning including physiological and clinical variables. Subsequently, we determined classifier performance by cross-validated machine learning. Patients with seizures (n = 49) had lower EDA levels (p = 0.031), EDA amplitudes (p = 0.045), and trended toward lower HR levels (p = 0.060) compared to patients without seizures (n = 68). Averaged cross-validated classification accuracy was 69% (AUC-ROC: 0.75). Our results show the potential to monitor and forecast risk for epileptic seizures based on changes in 24-h patterns in wearable recordings in combination with clinical variables. Such biomarkers might be applicable to inform care, such as treatment or seizure injury risk during specific periods, scheduling diagnostic tests, such as admission to the epilepsy monitoring unit, and potentially other neurological and chronic conditions. | 44809 | 2022
3S6SVEZA | 10.1039/d3lc00109a | https://www.ncbi.nlm.nih.gov/pubmed/37043225 | YES | skin-attached |  | NO | a versatile artificial skin platform that recapitulates skin topography, active sweat pores, skin wetting behaviour and sweat rate, and that can be tuned to mimic the specifications of the targeted body location and sweating characteristics (NO MOLECULAR/IONIC BIOMARKERS). |  |  |  |  | NA | yes | - | - | sweat | [22 empty] | NA | NA |  | NO | NA | NA | NA | 2023 | Moonen, E. J. M.; Ul Islam, T.; van Kemenade, S.; Pelssers, E.; Heikenfeld, J.; den Toonder, J. M. J. | A versatile artificial skin platform for sweat sensor development | Lab Chip |  | 1473-0189 (Electronic) 1473-0189 (Linking) | Research targeting the development of on-body sensors has been significantly growing in recent years - an example is on-skin sweat sensing. However, the wide inter and intra person variability of skin characteristics make in vivo testing of these sensors and included materials such as skin adhesives difficult, which hampers especially the initial development phase of such wearables. Besides the development of wearable sweat sensors, companies developing deodorants, cosmetics, medical adhesives and wearable textiles now need to perform expensive human subjects testing with little control over the exact sweat mechanisms. Hence, there is a need for a realistic, adaptable and stable test platform, or artificial skin. We present a versatile artificial skin platform that faithfully recapitulates skin topography, active sweat pores, skin wetting behaviour and sweat rate, and that can be tuned to mimic the specifications of the targeted body location and sweating characteristics. The developed artificial skin is capable of generating sweat rates as low as 0.1 nL min(-1) pore(-1) and as high as 100 nL min(-1) pore(-1), spanning the whole range of physiological sweat rates. Specifically, the platform can be used for the development of sweat sensors for sedentary persons whose sweat rates are commonly lower than currently delivered by any other artificial skin platform. | 45048 | 2023
6WGJ6W6Y | 10.1016/j.jcis.2022.07.011 | https://linkinghub.elsevier.com/retrieve/pii/S002197972201181X | YES | patch |  | NO | non-woven fabric treated with silver nanoplates (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2022 | Jiang, Shan; Cui, Ce; Bai, Wenhao; Wang, Weijie; Ren, Erhui; Xiao, Hongyan; Zhou, Mi; Cheng, Cheng; Guo, Ronghui | Shape-controlled silver nanoplates colored fabric with tunable colors, photothermal antibacterial and colorimetric detection of hydrogen sulfide | Journal of Colloid and Interface Science |  | 219797 |  | 44866 | 2022
XRUPNCGR | 10.3390/s24144710 | https://www.ncbi.nlm.nih.gov/pubmed/39066108 | YES | skin-attached |  | NO | measuring digital biomarkers like muscle oxygen levels and heart rate (NO MOLECULAR/IONIC BIOMARKERS) |  |  |  |  | NA | yes | - | - | sweat | [22 empty] | NA | NA |  | NO | NA |  |  | 2024 | Gorti, B.; Stephenson, C.; Sethi, M.; Gross, K.; Ramos, M.; Seshadri, D.; Drummond, C. K. | Overarm Training Tolerance: A Pilot Study on the Use of Muscle Oxygen Saturation as a Biomarker | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | Ulnar collateral ligament (UCL) tears occur due to the prolonged exposure and overworking of joint stresses, resulting in decreased strength in the flexion and extension of the elbow. Current rehabilitation approaches for UCL tears involve subjective assessments (pain scales) and objective measures such as monitoring joint angles and range of motion. The main goal of this study is to find out if using wearable near-infrared spectroscopy technology can help measure digital biomarkers like muscle oxygen levels and heart rate. These measurements could then be applied to athletes who have been injured. Specifically, measuring muscle oxygen levels will help us understand how well the muscles are using oxygen. This can indicate improvements in how the muscles are healing and growing new blood vessels after reconstructive surgery. Previous research studies demonstrated that there remains an unmet clinical need to measure biomarkers to provide continuous, internal data on muscle physiology during the rehabilitation process. This study's findings can benefit team physicians, sports scientists, athletic trainers, and athletes in the identification of biomarkers to assist in clinical decisions for optimizing training regimens for athletes that perform overarm movements; the research suggests pathways for possible earlier detection, and thus earlier intervention for injury prevention. | 45493 | 2024
QDDWUY8Z | 10.1016/j.jcis.2020.08.037 | https://linkinghub.elsevier.com/retrieve/pii/S0021979720310808 | YES | skin-attached |  | NO | multifunctional fabric (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2021 | Gao, Ya-Nan; Wang, Ye; Yue, Tian-Ning; Weng, Yun-Xuan; Wang, Ming | Multifunctional cotton non-woven fabrics coated with silver nanoparticles and polymers for antibacterial, superhydrophobic and high performance microwave shielding | Journal of Colloid and Interface Science |  | 219797 | Multifunctional cotton fabrics have attracted signiﬁcant attention as next-generation wearable materials. Herein, we report a facile method for the fabrication of ﬂexible and wearable cotton fabrics with ultrahigh electromagnetic interference (EMI) shielding, antibacterial, and superhydrophobic properties. Cotton fabrics were ﬁrst coated chemically with silver nanoparticles using polydopamine as adhesive and then with hydrophobic polydimethylsiloxane or polyimide. The introduction of polydopamine significantly increased the bond between silver nanoparticles and cotton ﬁbers, thereby preventing silver nanoparticles from falling off the surface. The composite fabrics exhibited a high conductivity of ~1000 S/cm, and their EMI shielding effectiveness increased up to ~110 dB. The composite fabrics exhibited excellent self-cleaning performance and acid-alkali corrosion resistance because of their superhydrophobicity. Notably, the fabric composites showed a signiﬁcant antibacterial action against Staphylococcus aureus and Escherichia coli. | 44197 | 2021
NPH2WXUI | 10.3390/s16111833 | https://www.ncbi.nlm.nih.gov/pubmed/27809270 | YES | skin-attached |  | NO | MONITORING HEART RATE (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2016 | Lou, C.; Li, R.; Li, Z.; Liang, T.; Wei, Z.; Run, M.; Yan, X.; Liu, X. | Flexible Graphene Electrodes for Prolonged Dynamic ECG Monitoring | Sensors (Basel) |  | 1424-8220 (Electronic) 1424-8220 (Linking) | This paper describes the development of a graphene-based dry flexible electrocardiography (ECG) electrode and a portable wireless ECG measurement system. First, graphene films on polyethylene terephthalate (PET) substrates and graphene paper were used to construct the ECG electrode. Then, a graphene textile was synthesized for the fabrication of a wearable ECG monitoring system. The structure and the electrical properties of the graphene electrodes were evaluated using Raman spectroscopy, scanning electron microscopy (SEM), and alternating current impedance spectroscopy. ECG signals were then collected from healthy subjects using the developed graphene electrode and portable measurement system. The results show that the graphene electrode was able to acquire the typical characteristics and features of human ECG signals with a high signal-to-noise (SNR) ratio in different states of motion. A week-long continuous wearability test showed no degradation in the ECG signal quality over time. The graphene-based flexible electrode demonstrates comfortability, good biocompatibility, and high electrophysiological detection sensitivity. The graphene electrode also combines the potential for use in long-term wearable dynamic cardiac activity monitoring systems with convenience and comfort for use in home health care of elderly and high-risk adults. | 42675 | 2016
5QIYN8EM | 10.1152/ajpregu.00192.2021 | https://www.ncbi.nlm.nih.gov/pubmed/34843411 | YES | head-mounted |  | NO | monitoring oxygenated haemoglobin (Oxy-Hb) levels as an indicator of brain activity (NO MOLECULAR/IONIC BIOMARKERS) | [5 empty] | - | - | - | [30 empty] | 2022 | Asahara, R.; Ishii, K.; Liang, N.; Hatanaka, Y.; Hihara, K.; Matsukawa, K. | Regional difference in prefrontal oxygenation before and during overground walking in humans: a wearable multichannel NIRS study | Am J Physiol Regul Integr Comp Physiol |  | 1522-1490 (Electronic) 0363-6119 (Linking) | Using wireless multichannel near-infrared spectroscopy, regional difference in cortical activity over the prefrontal cortex (PFC) was examined before and during overground walking and in response to changes in speed and cognitive demand. Oxygenated-hemoglobin concentration (Oxy-Hb) as index of cortical activity in ventrolateral PFC (VLPFC), dorsolateral PFC (DLPFC), and frontopolar cortex (FPC) was measured in 14 subjects, whereas heart rate was measured as estimation of exercise intensity in six subjects. The impact of mental imagery on prefrontal Oxy-Hb was also explored. On both sides, Oxy-Hb in VLPFC, DLPFC, and lateral FPC was increased before the onset of normal-speed walking, whereas Oxy-Hb in medial FPC did not respond before walking onset. During the walking, Oxy-Hb further increased in bilateral VLPFC, whereas Oxy-Hb was decreased in DLPFC and lateral and medial FPC. Increasing walking speed did not alter the increase in Oxy-Hb in VLPFC but counteracted the decrease in Oxy-Hb in DLPFC (but not in lateral and medial FPC). Treadmill running evoked a greater Oxy-Hb increase in DLPFC (n = 5 subjects). Furthermore, increasing cognitive demand during walking, by deprivation of visual feedback, counteracted the decrease in Oxy-Hb in DLPFC and lateral and medial FPC, but it did not affect the increase in Oxy-Hb in VLPFC. Taken together, the profound and localized Oxy-Hb increase is a unique response for the VLPFC. The regional heterogeneity of the prefrontal Oxy-Hb responses to natural overground walking was accentuated by increasing walking speed or cognitive demand, suggesting functional distinction within the PFC. | 44562 | 2022