Disagreement between anterior segment optical coherence tomography devices in eyes with a shallow anterior chamber
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Supplemental Table 2 Gonioscopy, Visante® and DRI-Triton® OCT agreement on angle closure detection.
	
	Gonioscopy and 
Visante® OCT
	
	Gonioscopy and DRI OCT Triton®
	
	Visante® OCT and DRI OCT Triton®

	Quadrant
	AC1 (95% CI)
	
	AC1 (95% CI)
	
	AC1 (95% CI)

	All
	0.52 (0.33 – 0.70)
	
	0.12 (-0.09 – 0.33)
	
	0.44 (0.25 – 0.63)

	Superior
	0.66 (0.33 – 0.99)
	
	0.34 (-0.10 – 0.77)
	
	0.41 (0.00 – 0.81)

	Nasal
	0.07 (-0.39 – 0.53)
	
	-0.22 (-0.65 – 0.22)
	
	0.63 (0.29 – 0.98)

	Inferior
	0.66 (0.33 – 0.99)
	
	0.30 (-0.15 – 0.76) 
	
	0.49 (0.10 – 0.88)

	Temporal
	0.67 (0.34 – 1.00)
	
	0.15 (-0.30 – 0.59)
	
	0.34 (-0.10 – 0.77)


Abbreviations: AC1 = first-order agreement coefficient; 95% CI = 95% Confidence Interval OCT = optical coherence tomography;

