Supplementary

# Methodology

## Imaging and Localization Overview

In medical image-guided navigation, C-arm X-ray imaging can be abstracted as a pinhole camera system. The pinhole camera model is one of the most widely used geometric models in computer vision, describing the mapping from three-dimensional space to the two-dimensional image plane through linear projection. Its core lies in estimating the intrinsic and extrinsic parameters of the imaging system.

Fig. S1 Overview of Camera Imaging Model and Localization

As illustrated in Fig. S1, the X-ray pinhole camera model maps a 3D feature point expressed in the object coordinate system to the X-ray image coordinate system via a projective transformation. The mapping equation is given by:

(1)

Where  denotes the homogeneous 2D coordinates of a feature point in the X-ray image coordinate system,  denotes the homogeneous 3D coordinates of the corresponding feature point in the object coordinate system, is a scale factor, represents the image coordinates,  is the transformation (projection) matrix, and  denotes the 3D coordinates in the target coordinate system. The transformation matrix  contains 12 unknown parameters. Substituting the coordinates into the mapping equation yields:

Here  denotes the homogeneous 2D image coordinates of the feature point in the X-ray image,  represents the homogeneous 3D coordinates of the corresponding feature point in the object coordinate system, is a scale factor, and  is the projection matrix. The projection matrix  contains 12 unknown parameters and encapsulates both intrinsic and extrinsic imaging geometry.

(2)

That is,  and . Using  point correspondences, a system of  nonlinear equations can be constructed, i.e.,

(3)

Where,, , (3) can be solved using a least-squares formulation.

To achieve three-dimensional localization of a target point, the system requires X-ray images acquired from at least two distinct viewing angles. In practice, a biplanar configuration is typically adopted, consisting of anteroposterior (AP) and lateral (LAT) views. Under ideal, noise-free conditions, two projection matrices can be estimated by solving Eq. (3) for each view. Given a corresponding point observed in both images, its three-dimensional position can then be reconstructed through the intersection of the associated projection rays.

(4)

(5)

(6)

(7)

The matrix  is decomposed using singular value decomposition (SVD), , , The solution vector  is given by the right singular vector associated with the smallest singular value, i.e., the last column of . After normalization,  represents the reconstructed three-dimensional coordinates of the target point.

## Structural Error Propagation Mechanism

Error Analysis

In practical image-guided surgical applications, various sources of uncertainty are inevitably introduced throughout the system, including measurement noise, numerical estimation error, and coordinate transformation inaccuracies. These errors do not exist in isolation within individual modules. Instead, they propagate and accumulate along the entire localization pipeline—spanning perception, geometric reconstruction, and coordinate mapping—and are progressively amplified through multi-stage geometric estimation. As a result, such system-level error propagation can substantially degrade three-dimensional target localization accuracy and ultimately compromise the execution accuracy of the robotic manipulator.

Fig.S2 Reference-based geometric perception pipeline in biplanar X-ray–guided robotic navigation. The C-arm X-ray imaging system acquires images of a localization reference structure rigidly attached to the robotic end effector. The reference structure consists of a set of fiducial points with known three-dimensional geometry. By detecting the two-dimensional pixel coordinates of these fiducials in the X-ray images and associating them with their pre-calibrated three-dimensional coordinates expressed in the connection.

As illustrated in Fig. S2, the C-arm X-ray imaging system acquires images of a localization reference structure rigidly attached to the end effector of the robotic manipulator. The reference structure contains a set of fiducial points with well-defined geometric features. By detecting the two-dimensional pixel coordinates of these fiducials in the X-ray images and associating them with their pre-calibrated three-dimensional coordinates expressed in the connector component coordinate frame, the system can estimate the transformation matrix from the connector component frame to the C-arm X-ray world coordinate frame by solving Eq. (3).

Under ideal, error-free conditions, this transformation matrix is denoted as . However, in practical systems, installation-induced structural perturbations between the connector component and the localization reference structure introduce systematic errors into the input three-dimensional coordinates. As a result, the estimated transformation deviates from the ideal case and is instead represented as a perturbed transformation .

Under this condition, Eq. (1) can be rewritten as:

(8)

It follows directly that the estimated transformation matrix becomes .

Under this condition, Eq. (4) can be rewritten as, . Compared with , we can obtain that the perturbation in  directly affects the estimation of . As a consequence, the resulting three-dimensional target position is obtained as =.

Since the estimated position  is expressed in the connector component coordinate frame, execution of the robotic motion requires transforming this position into the robot tool center point (TCP) coordinate frame.

(9)

In practice, installation-induced structural perturbations exist not only between the connector component and the localization reference structure, but also between the connector component and the robotic end effector. Consequently, the actual target position expressed in the TCP coordinate frame can be written as

(10)

Under rigid-body kinematics, the installation-induced error  can be decomposed into a rotational perturbation about the -, -, and -axes and a translational offset.

(11)

(12)

Since the rotation angles are typically small, a first-order small-angle approximation can be applied. Specifically, the cosine terms in the transformation matrix are approximated as unity, while the sine terms are approximated by the corresponding rotation angles. By neglecting higher-order terms, the perturbed transformation can be expressed as:

(13)

Under this approximation, we obtain:

(14)

It follows that both  and  directly influence the final localization of the target point.

## Uncertainty and Tail-Risk Analysis

For Eq. (3), the three-dimensional coordinates of the control points  and the corresponding two-dimensional image observations  are modeled as mutually independent random variables following zero-mean Gaussian distributions.

Let  denote the covariance matrix of the output , and let  denote the covariance matrix of the input . According to the principle of uncertainty propagation, the following relationship holds:

(15)

Where, ,

Here,  denotes the Jacobian matrix of the function with respect to the input vector , which consists of the three-dimensional control point coordinates x, y, z and the corresponding two-dimensional image observations u, v. ,  where

In addition,  denotes the Jacobian of the function  with respect to the twelve parameters of the transformation matrix , which map the parameter vector  to the output . ,

We can obtain.

For the computation of the three-dimensional target point coordinates, considering Eq. (5), we denote  as the mapping function, where  represents the associated input variables. Let  denote the covariance matrix of the output vector  of , and let denote the covariance matrix of the output vector . According to the uncertainty propagation principle, the resulting covariance is given by  .

Where, , , , , ,

The pseudoinverse can thus be obtained as

(16)

# Experiments and Analysis

## Simulation Setup

3.2.1 Biplanar X-ray Imaging and Reconstruction

In simulation, biplanar X-ray imaging is modeled using a pinhole camera formulation. For each control point,

(17)

where  denotes homogeneous pixel coordinates,  is the projection matrix, and represents the potentially perturbed 3D control point. Projection matrices are estimated using DLT and least squares from multiple 3D–2D correspondences, and the target point is reconstructed via biplanar triangulation to obtain .

3.2.2 Error Source Modeling

Two primary error sources are considered. Pixel detection noise is modeled as zero-mean Gaussian noise added to 2D observations. Installation error is treated as a structural perturbations, where the 3D control point input is perturbed by small rotations and translations:

(18)

To reflect the execution chain in robotic navigation, the transformation from the reconstructed target to the robot TCP frame is included, with an additional installation error introduced between the mounting component and the end effector.

3.2.3 Evaluation Metrics

Performance is evaluated using the 3D localization error, TCP execution error, and average reprojection error. Tail-risk behavior is characterized using the 95th percentile (P95) and worst-case error to capture extreme failure modes arising from nonlinear geometric estimation.

## Simulation III: Analytic Uncertainty Propagation versus Monte Carlo

The third simulation evaluates the predictive capability of the proposed error propagation model. Under a representative operating condition, output error covariances are computed using both Monte Carlo sampling and first-order uncertainty propagation.

The comparison shows that analytic propagation accurately predicts dominant error directions and relative magnitudes, while underestimating absolute variance under strongly nonlinear conditions. This suggests that the analytic model is effective for sensitivity analysis, whereas statistical methods are required for reliable tail-risk assessment.

# Simulation Results

To systematically validate the proposed system-level error propagation mechanism in biplanar X-ray–guided navigation, three simulation studies are conducted to analyze:

(1) the primary effect of installation error

(2) the coupled amplification between installation error and pixel noise, and

(3) the predictive capability of first-order analytic uncertainty propagation.

All simulations are based on the biplanar pinhole imaging model described in Section III and evaluated using Monte Carlo sampling with 2000 trials per configuration. Both 3D localization error and TCP execution error are reported, together with tail-risk metrics to characterize worst-case behavior.

## Simulation Study I: Primary Effect of Installation Error

We first analyze the direct impact of reference structure installation error on localization and execution accuracy. Rotational and translational installation errors are independently scanned while pixel noise is fixed at a low level, so as to isolate the dominant effect of installation-induced structural perturbations.

### Rotational Installation Error

Quantitative results are summarized in Table S1. Increasing the rotation misalignment from 0^∘to 2^∘amplifies the mean localization error from approximately 18 mm to over 75 mm. Meanwhile, the P95 tail risk of the TCP execution error nearly doubles, rising from about 45 mm to close to 90 mm.

It is important to note that, in this experiment, the TCP transformation is assumed to be an ideal rigid-body mapping without additional execution uncertainty. Consequently, the execution error magnitude matches the localization error magnitude. This observation does not imply that the execution chain is insensitive to error, but rather indicates that structural installation error introduced at the perception and reconstruction stages alone is sufficient to determine the upper bound and tail-risk behavior of navigation accuracy.

Table S1. Effect of installation rotation and translation errors on localization accuracy (Sim1)

| Installation Error | (mm) | (mm) | P95 (mm) |
| --- | --- | --- | --- |
| Rotation 0.0° | 0.018 | 0.018 | 0.045 |
| Rotation 0.5° | 0.027 | 0.027 | 0.048 |
| Rotation 1.0° | 0.042 | 0.042 | 0.058 |
| Rotation 1.5° | 0.058 | 0.058 | 0.070 |
| Rotation 2.0° | 0.075 | 0.075 | 0.086 |
| Translation 0 mm | 0.019 | 0.019 | – |
| Translation 1 mm | 0.018 | 0.018 | – |
| Translation 2 mm | 0.019 | 0.019 | – |
| Translation 3 mm | 0.018 | 0.018 | – |
| Translation 4 mm | 0.019 | 0.019 | – |
| Translation 5 mm | 0.019 | 0.019 | – |

### Translational Installation Error

As reported in Table S1, translational misalignment within the range of 0–5 mm results in only marginal variations in both localization and execution error, with error magnitudes remaining nearly constant.

## Simulation Study II: Coupled Amplification of Installation Error and Pixel Noise

In practical systems, installation error and pixel-level perception noise coexist. To investigate their interaction, we perform a coupled simulation study by jointly varying installation error levels (L0–L2) and pixel noise magnitude.

This behavior is quantitatively summarized in Table S2. For example, under moderate pixel noise, the P95 TCP error increases from 90.39 mm at L0 to 108.56 mm at L2. Under high pixel noise, tail risk exceeds 200 mm regardless of whether installation error is present, indicating that the system has entered a high-risk regime.

These results demonstrate that system error is not a simple additive combination of perception noise and installation error. Instead, installation error fundamentally alters the system’s sensitivity to pixel-level uncertainty, transforming otherwise tolerable noise into a navigation-level risk factor through structural coupling in the geometric estimation pipeline.

Table S2. Coupled amplification of pixel noise and installation error (P95 of execution error)

| Installation Level |  |  |  |
| --- | --- | --- | --- |
| L0 (0° + 0 mm) | 0.00 mm | 0.090 mm | 0.240 mm |
| L1 (1° + 2 mm) | 0.036 mm | 0.098 mm | 0.242 mm |
| L2 (2° + 5 mm) | 0.072 mm | 0.108 mm | 0.219 mm |

## Simulation Study III: Analytic Uncertainty Propagation versus Monte Carlo Statistics

To further evaluate the predictive capability of the proposed error propagation model, we compare first-order analytic uncertainty propagation with Monte Carlo statistics under a representative operating condition.

Table S3 provides a quantitative comparison between analytic and Monte Carlo results. While the analytic model predicts the correct ordering and anisotropy of error components, the depth-direction standard deviation is underestimated by approximately a factor of two. This behavior is consistent with the linearization assumptions underlying first-order propagation and highlights its inherent limitations under strong nonlinear geometric conditions.

Rather than constituting a deficiency, this comparison clarifies the complementary roles of analytic and statistical approaches: analytic propagation provides interpretable sensitivity analysis and directional insight, while Monte Carlo sampling is necessary for reliable tail-risk estimation in nonlinear regimes.

Table S3. Comparison of analytic and Monte Carlo uncertainty propagation (Sim3)

| Axis | MC Mean (mm) | ANA Mean (mm) | MC Std (mm) | ANA Std (mm) |
| --- | --- | --- | --- | --- |
| X | 0.602 | 0.602 | 0.015 | 0.010 |
| Y | 0.515 | 0.515 | 0.010 | 0.008 |
| Z | 0.221 | 0.220 | 0.043 | 0.019 |

Across the three simulation studies, the results consistently indicate that installation error acts as a primary driver of system-level error propagation in biplanar X-ray navigation. Rotational misalignment governs localization accuracy and tail-risk behavior, installation error and pixel noise exhibit coupled amplification, and first-order analytic propagation captures dominant sensitivity structure while requiring statistical evaluation for worst-case risk assessment.

# Real-System Validation

Table S4 summarizes the localization error statistics under different installation error levels. Regardless of whether 3D localization error or reprojection error is used, the error magnitude exhibits a clear monotonic increase across installation levels, i.e., L0 < L1 < L2. Under ideal installation conditions (L0), the localization error remains at a relatively low level of 0. Introducing moderate installation error (L1) increases the mean error to 0.551, while larger misalignment (L2) further degrades accuracy to 1.380. These results indicate that installation perturbations significantly deteriorate localization accuracy in a real biplanar navigation pipeline.

Beyond mean error, tail-risk metrics show an even more pronounced dependence on installation error. As the installation level increases, the growth of P95 substantially exceeds that of the mean error, indicating that installation error not only elevates overall error magnitude but also markedly amplifies worst-case behavior. This observation is consistent with the simulation results, which identified structural installation error as the dominant contributor to tail-risk amplification.

For quantitative comparison, Table S4 reports the mean error, standard deviation, and P95 values under each installation level. As shown in the table, both 3D localization error and reprojection error increase consistently with installation misalignment, with P95 being particularly sensitive to changes in installation state. This further confirms that installation error is systematically amplified through the geometric reconstruction chain in real systems.

Table S4. Real-system validation results under different installation error levels

| Level | Nominal misalignment | Mean (mm) | Std (mm) | P95 (mm) | Mean (px) |
| --- | --- | --- | --- | --- | --- |
| L0 | 0.000 | 0.000 | 0.000 | 0.000 | 0.000 |
| L1 | ~1° / 2 mm | 0.551 | 0.0046 | 0.556 | 2.0 |
| L2 | ~2° / 5 mm | 1.380 | 0.0110 | 1.391 | 5.0 |