Detailed Methodology for the Quantification of Serum Bisphenols via LLE-UHPLC-MS/MS

1. Detection of Serum Bisphenols in Pregnant Women

Serum BPA, BPS, BPF and BPAF concentrations were determined using liquid-liquid extraction coupled with ultra-high-performance liquid chromatography-tandem mass spectrometry (LLE-UHPLC-MS/MS) [1-5]. .

| Table S1.1 The 4 BPs detected in this study |  |  |  |
| --- | --- | --- | --- |
| Bisphenols | Bisphenols | Molecular Formula | Molecular Weight |
| Bisphenol A | BPA | C₁₅H₁₆O₂ | 228.29 |
| Bisphenol S | BPS | C1₂H₁₀O₄S | 250.28 |
| Bisphenol F | BPF | C₁₃H₁₂O₂ | 200.24 |
| Bisphenol AF | BPAF | C₁₅H₁₀F₆O₂ | 336.23 |

2. Laboratory Instruments and Consumables

Table S1.2 Main Instruments

| Laboratory Instruments | Model | Manufacturer | Country |
| --- | --- | --- | --- |
| Ultra-High Performance Liquid Chromatograph | Nexera X2 | SHIMADZU | Japan |
| Triple Quadrupole Mass Spectrometer | AB SCIEX Triple Quad 5500 | AB SCIEX | USA |
| Chromatographic Column | ZORBAX SB-Phenyl | Agilent | USA |
| LC Guard Column | 1290 Infinity II | Agilent | USA |
| Electronic Analytical Balance | MS105DU | Mettler Toledo | Switzerland |
| High-Speed Centrifuge | / | Eppendorf | Germany |
| Milli-Q Ultrapure Water System | Milli-Q Advantage A10 | Millipore | USA |
| Dry Nitrogen Evaporator | NDK200-2N | Hangzhou Miou Instrument Co., Ltd. | China |

Table S1.3  Main Experimental Consumables

| Consumables | Specification | Manufacturer | Country |
| --- | --- | --- | --- |
| BPA standard (for calibration) | CAS: 80-05-7; ZC-60176 | Zzstandard | China |
| BPS standard (for calibration) | CAS: 80-09-1; ULM-9320-S | CIL | USA |
| BPF standard (for calibration) | CAS: 620-92-8; ULM-9827-S | CIL | USA |
| BPAF standard (for calibration) | CAS: 1478-61-1; 90477 | Sigma | USA |
| ¹³C-labeled BPA internal standard | CLM-9776-1.2 | Cambridge Isotope Laboratories | USA |
| ¹³C-labeled BPS internal standard |  | Cambridge Isotope Laboratories | USA |
| ¹³C-labeled BPF internal standard |  | Cambridge Isotope Laboratories | USA |
| ¹³C-labeled BPAF internal standard |  | Cambridge Isotope Laboratories | USA |
| MS-grade Methanol & Acetonitrile | / | Honeywell | USA |
| MS-grade Ethyl Acetate & Formic Acid | / | Sigma | USA |
| Concentrated Sulfuric Acid & Potassium Dichromate | / | Sinopharm Chemical Reagent Co., Ltd. | China |
| Fetal Bovine Serum | / | Zhejiang Tianhang Biotechnology Co., Ltd. | China |

- Main Experimental Procedures

- Gradual Thawing: Serum samples were removed from the -80°C freezer and placed in a -20°C freezer overnight, then transferred to a 4°C refrigerator for storage until use.

- Addition of Internal Standard and Extraction Solvent (First Extraction): A 200 μL aliquot of serum was mixed with 20 μL of the ¹³C-labeled internal standard (BPA, BPS, BPF) working solution. Subsequently, 600 μL of 2% ethyl acetate (solvent composition to be confirmed, e.g., in hexane or another specified solvent) was added to initiate the first liquid-liquid extraction. The mixture was then processed sequentially as follows: vortex mixing for 20-30 seconds, continuous agitation on a vortex mixer at 2500 rpm for 40 minutes, ultrasonication for 5 minutes, and centrifugation at 12,000 × g for 10 minutes.

(3) Second Extraction: A 500 μL aliquot of the supernatant from the first extraction was transferred to a new tube. An additional 600 μL of 2% ethyl acetate was added, and the second extraction was performed following the identical procedure described in step (2).

(4) Nitrogen Evaporation and Reconstitution: The combined organic extracts were evaporated to complete dryness under a gentle stream of nitrogen at 37°C. The residue was reconstituted in 60 μL of a 50:50 (v/v) methanol/acetonitrile solution. This reconstituted solution was then vortex-mixed for 1 minute, agitated at 2500 rpm for 40 minutes, ultrasonicated for 5 minutes, and centrifuged at 12,000 × g for 10 minutes.

(5) Instrumental Analysis: A 40 μL aliquot of the resulting supernatant was transferred to a glass insert vial for subsequent analysis by UHPLC-MS/MS.

- Ion Conditions

Table S1.4  Optimized MRM parameters for Bisphenols and the internal standard

| Bisphenols | Precursor Ion (m/z) | Product Ion (m/z) | Declustering Potential (V) | Collision Energy (V) | Retention Time (min) |
| --- | --- | --- | --- | --- | --- |
| BPA | 227.2 | 212.0 | 140.0 | 24.0 | 3.80 |
| BPS | 249.0 | 156.0 | 170.0 | 30.0 | 4.20 |
| BPF | 199.0 | 104.8 | 130.0 | 28.0 | 3.73 |
| BPAF | 335.0 | 315.0 | 100.0 | 25.0 | 4.25 |
| ¹³C-labeled BPA(IS) | 241.1 | 142.2 | 146.0 | 35.0 | 4.01 |
| ¹³C-labeled BPS(IS) | 261.0 | 162.0 | 67.0 | 30.0 | 4.02 |
| ¹³C-labeled BPAF(IS) | 347.0 | 326.9 | 70.0 | 27.0 | 4.50 |

IS: internal standard.

- Quality Control

(1) Procedural Blanks: Reagent blanks were analyzed at the beginning of each sequence to assess potential instrumental carryover or background contamination. The response from these blanks served as the background value for subsequent quantification.

(2) In-Run Quality Control: For each analytical batch (comprising up to 15 samples), one procedural blank and one matrix-spiked quality control (QC) sample were incorporated. The procedural blank was used to monitor contamination from solvents or the sample preparation process, while the matrix-spiked QC sample was used to verify the stability and performance of the instrument throughout the run.

(3) Recovery and Precision (Accuracy and Precision): Method accuracy was determined by recovery experiments. Blank fetal bovine serum was spiked with bisphenols (for example BPF) at two concentration levels: a low level of 200 pg/mL and a high level of 1000 pg/mL (n=6 replicates per level). The mean recoveries were 91.89% and 100.23%, respectively. 
    Method precision was evaluated by analyzing the 200 pg/mL QC sample over three consecutive days (6 replicates per day). Precision was expressed as the relative standard deviation (RSD). The intra-day (within-day) RSD was 3.85%, and the inter-day (between-day) RSD was 7.57%.

(4) Matrix Effect: The matrix effect was assessed by comparing the analyte response in post-extraction spiked serum (matrix-matched) to the response in pure solvent (neat standard solution) at the same concentrations (200 pg/mL and 1000 pg/mL; n=6 each). The matrix effect, calculated as (response in matrix / response in solvent) × 100%, was 94.77% and 100.29% at the low and high concentration levels, respectively.

(5) Calibration: A matrix-matched calibration curve was constructed using blank fetal bovine serum spiked with known concentrations of Bisphenols. The calibration standards were processed identically to the study samples to account for procedural variations. The calibration curve exhibited excellent linearity across the working range.

Reference:

- Li YL, Mao HY, Yan LM, Liu ZY, Sun S, Jiang TT, Xu LY, Zhang XY, Liu BL, Zhu M, Xie JA, Hao JH, Zhang C. Association between prenatal exposure to bisphenols and preterm birth: the interaction effect of anemia. Environ Int. 2025 Aug 11;203:109724. pii: S0160-4120(25)00475-1. doi: 10.1016/j.envint.2025.109724. PubMed PMID: 40829515.

- Li Y, Yuan M, Zhang C, Zhang X, Hao J, Tao F, Wang G, Su P. Bisphenol AF exposure synergistically increases the risk for suicidality among early adolescents with child maltreatment: A prospective cohort study. Ecotoxicol Environ Saf. 2024 May 28;279:116511. pii: S0147-6513(24)00587-6. doi: 10.1016/j.ecoenv.2024.116511. PubMed PMID: 38810289.

- Dou L, Sun S, Chen L, Lv L, Chen C, Huang Z, Zhang A, He H, Tao H, Yu M, Zhu M, Zhang C, Hao J. The association between prenatal bisphenol F exposure and infant neurodevelopment: The mediating role of placental estradiol. Ecotoxicol Environ Saf. 2024 Jan 25;271:116009. pii: S0147-6513(24)00084-8. doi: 10.1016/j.ecoenv.2024.116009. PubMed PMID: 38277971.

- Zhang C, Zhou L, Wu XC, Guan TY, Zou XM, Chen C, Yuan MY, Li YH, Wang S, Tao FB, Hao JH, Su PY. Association of serum bisphenol AF concentration with depressive symptoms in adolescents: A nested case-control study in China. Ecotoxicol Environ Saf. 2022 Aug;241:113734. pii: S0147-6513(22)00574-7. doi: 10.1016/j.ecoenv.2022.113734. PubMed PMID: 35679728.

- Huang Z, Fu W, Dou L, Bao H, Wu W, Su P, Huang K, Zhu P, Sheng J, Xu Y, Tao F, Hao J. Prenatal Bisphenol A Exposure and Early Childhood Behavior and Cognitive Function: A Chinese Birth Cohort Study. Neuroendocrinology. 2022;112(4):311-323. pii: 000516881. doi: 10.1159/000516881. PubMed PMID: 33910209.