Standard Operating Procedure (SOP)

5-Hour Temperature-Dependent Raman Study of Monolayer Graphene (Heating–Cooling Hysteresis)

1. Sample Information

- Material: CVD-grown monolayer graphene on 300 nm SiO₂/Si

- Supplier Verification:

  - Monolayer confirmed

  - Low defect density: I(D)/I(G) < 0.1

  - Monolayer signature: I(G)/I(2D) < 1

  - Areal coverage > 95%

Purpose:
High-quality, manufacturer-verified graphene ensures that observed effects arise from thermal irreversibility, not defects or sample inhomogeneity.

2. Instrumentation

- Raman System: Confocal micro-Raman

- Laser: 532 nm CW

- Laser Power (at sample): 0.5 mW

- Objective: 50×, NA ≥ 0.75

- Spectral Range: 1200–3200 cm⁻¹

- Spectral Resolution: ≤ 1.5 cm⁻¹

- Detector: CCD

- Temperature Stage: 300–500 K

- Temperature Stability: ±1 K

- Atmosphere: Ambient (note if N₂ or vacuum)

3. Calibration & Pre-Checks (45 minutes)

3.1 Wavenumber Calibration (15 min)

- Calibrate using Si peak at 520.0 cm⁻¹

- Acceptable drift: ±0.5 cm⁻¹

3.2 Focus, Alignment & Spot Selection (15 min)

- Identify uniform graphene region

- Mark primary measurement spot

- Mark secondary sacrificial spot (for power checks)

3.3 Laser Heating Verification (15 min)

- At RT, collect spectra at:

  - 0.3 mW

  - 0.5 mW

  - 0.7 mW

- Confirm linear phonon shift vs power

4. Measurement Protocol (Total Experimental Time ≈ 5 Hours)

Hour 1 | Room-Temperature Baseline (60 min)

- Measure three distinct spots

- Integration time: 30–60 s per spectrum

- Acquire Stokes + Anti-Stokes

- Purpose:

  - Establish baseline Raman parameters

  - Confirm spatial uniformity

Hour 2 | Heating Ramp (300 → 500 K) (75 min)

Temperatures: 300, 350, 400, 450, 500 K

At each temperature:

- Stabilization: 5 min

- Acquisition: 2 × 60 s (Stokes + Anti-Stokes)

- Measurement at fixed primary spot

Total per temperature ≈ 15 min

Hour 3 | High-Temperature Hold & Control (45 min)

- Hold at 500 K for 20 min

- Acquire:

  - Repeated spectra at fixed power

  - One high-T power check (0.3 / 0.5 / 0.7 mW) at sacrificial spot

- Purpose:

  - Confirm thermal equilibrium

  - Rule out power-induced artifacts

Hour 4 | Cooling Ramp (500 → 300 K) (75 min)

Temperatures: 500, 450, 400, 350, 300 K

- Same stabilization and acquisition protocol as heating

- Measure exact same spot used in heating ramp

Hour 5 | Post-Cycle Verification & Mapping (45 min)

- RT spectrum at original spot

- Small RT Raman map (5×5 points, 1–2 µm pitch)

- Verify:

  - No permanent damage

  - No defect generation

- Backup all data

5. Extracted Parameters

- G peak position and FWHM

- 2D peak position and FWHM

- I(2D)/I(G), I(D)/I(G)

- Anti-Stokes/Stokes ratio (G mode)

6. Data Analysis

6.1 Local Temperature Validation

IASIS=νASνS4exp⁡ ⁣ℏΩkBTlocal

Acceptance:
|T_local − T_stage| ≤ 5 K

6.2 Hysteresis Quantification

Δωhyst(T)=ωheat(T)-ωcool(T)
ΔΓhyst(T)=Γheat(T)-Γcool(T)

Applied to G and 2D modes.

7. Acceptance Criteria

- Si calibration within ±0.5 cm⁻¹

- No permanent RT peak shift

- Stable I(D)/I(G) < 0.1

- Reproducible Anti-Stokes temperature

- No visible degradation

8. Expected Results

- Temperature-dependent phonon softening

- Clear heating–cooling hysteresis

- Linewidth asymmetry indicating irreversible coupling

- Absence of defect-driven artifacts

9. Data Handling

- Raw spectra archived by temperature and ramp direction

- Fitting outputs stored as CSV

- Scripts version-controlled