Table 1 Composition of Simulated Menstrual Blood Solutions
Volume-based (1A–1D) and hemoglobin-based (2A–2D) porcine blood–saline dilution schemes used for benchtop calibration. For hemoglobin-based solutions, target hemoglobin concentrations were achieved within ±5% tolerance.

| Solution | Concentration (%) | Blood Volume (mL) | Saline Volume (mL) |
| --- | --- | --- | --- |
| 1A | 25 | 125 | 375 |
| 1B | 50 | 250 | 250 |
| 1C | 75 | 375 | 125 |
| 1D | 100 | 500 | 0 |

| Solution | Concentration (%) | 14 g/dL Blood Volume (mL) | Saline Volume (mL) | Target Hgb Concentration (g/dL) Tolerance ± 5% |
| --- | --- | --- | --- | --- |
| 2A | 25 | 125 | 375 | 3.5 |
| 2B | 50 | 250 | 250 | 7 |
| 2C | 75 | 375 | 125 | 10.5 |
| 2D | 100 | 500 | 0 | 14 |

Table 2 Quantiflow Algorithm Performance Summary for Solution 1D(<20mL)
Performance metrics for undiluted porcine whole blood (solution 1D) at volumes ≤20 mL, including linear regression parameters, Bland–Altman agreement statistics, error metrics, inter-batch consistency, and repeatability. All predefined acceptance criteria were met.

| Category | Metric | Result | Threshold / Interpretation | Outcome |
| --- | --- | --- | --- | --- |
| Linear Regression (Zero-Intercept) | R2, slope | R2 = 0.9757 Slope = 0.9040 | >= 0.80 0.9-1.10 | PASS |
| Bland-Altman Analysis | Mean Bias, 90%LoA (mL), 90% LoA(% of mean volume | -0.5486mL (4.48%), -4.32 to 3.22 mL, 35.23% | <= 5%, Closer to 0 is better, <= 40% | PASS, ---, PASS |
| Error Metrics | MAE, MAPE | 1.8638 mL (15.21%), 18.05% | <= 20%, <= 20% | PASS, PASS |
| Inter-Batch Consistency | Slope Variation, R2 Variation Batch 1 Batch 2 Batch 3 | 7.05%, 1/04%, Slope 0.9533, R2 0.9670 Slope 0.8069, R2 0.9682 Slope 0.9228, R2 0.9465 | <= 10%, <= 10%, -- -- -- | PASS, PASS, -- -- -- |
| Repeatability (Coefficient of Variation) | 5 mL 10 mL 15 mL 20 mL | 15.92%, 7.48%, 8.28%, 10.07% | <= 20%, <= 20%, <= 20%, <= 20% | PASS PASS PASS PASS |
| Overall Assessment | All criteria met | PASS | PASS | PASS |