Additional file 1.

Additional Methods and Tables

This additional file provides detailed MRI acquisition parameters and supplementary tables supporting the main manuscript.

Supplementary Methods S1. Phantom MRI acquisition

3D and 4D-MRI: The 4D-MRI for the phantom was simulated by acquiring five static 3D T2W-MRI volumes at predefined positions within one motion cycle. A reference 3D T2W FSE scan was performed (TR = 2100 ms, TE = 205.585 ms, slice thickness = 1.2 mm; total scan time = 3.3 minutes), and the five phase volumes (e.g., 0%, 25%, 50%, 75%, and 100%) were assembled as the simulated 4D-MRI dataset for contouring.

Cine-MRI: Sagittal cine-MRI was acquired for the phantom using the default thorax protocol employing a T2 3D Tra sequence with the following parameters: TR = 2100 ms, TE = 206 ms, voxel dimensions = 2.0 x 2.0 x 2.4 mm, field-of-view = 320 x 448 x 300 mm, scan time = 200-228 s, matrix size = 160 x 224 x 250 slices, slice thickness = 2.4 mm, and number of signal averages = 2.

Supplementary Methods S2. Patient MRI acquisition

Cine-MRI: During patient positioning, patients were instructed to position both arms above the head, consistent with the radiotherapy delivery position. Sagittal cine-MRI was acquired using the Unity MR-Linac system (Elekta AB, Stockholm, Sweden). All MRI scans were performed on the integrated 1.5T wide-bore MRI system using an 8-channel body array coil. Sequence parameters were as follows: repetition time = 3.6 ms, echo time = 1.5 ms, flip angle = 70 degrees, temporal resolution = 0.5 s per frame (2 images per second), matrix size = 224 x 256 (interpolated to 512 x 512), field-of-view = 40 x 40 cm, and bandwidth = 125 kHz.

Additional file 1:

| Table S1. Motion Pattern Parameters for the 4D Dynamic Thorax Phantom Study |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Motion pattern | Time(s) | Max Speed (mm/s) | Amplitude Direction |  |  | Direction |
|  |  |  | IS | AP | LR |  |
| Sin-10A | 4 | 5 | 10 | 0 | 0 | IS |
| Sin-25A | 4.9 | 10.2 | 25 | 0 | 0 | IS |
| Cos4AP | 4 | 5 | 10 | 2.5 | 1 | IS/AP/LR |
| Heart | 1 | 10 | 5 | 8 | 6 | IS/AP/LR |
| Sawtooth | 4 | 5 | 10 | 2.5 | 1 | IS/AP/LR |
| Note: IS = Superior-Inferior, AP = Anterior-Posterior, LR = Left-Right. |  |  |  |  |  |  |

| Table S2. Summary of tumor motion amplitude measurements |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| A. Patient study: Motion amplitudes measured by 4D-CT and cine-MRI |  |  |  |  |  |
| Cohort | Modality | Geometric Center Amplitude (mm, SI direction) |  |  |  |
|  |  | Mean ± SD | Range |  |  |
| Lung tumors (n=10) | 4D-CT | 16.7 ± 4.2 | 5.5 – 27.8 |  |  |
|  | cine-MRI | 16.1 ± 3.4 | 6.7 – 28.5 |  |  |
| Liver tumors (n=10) | 4D-CT | 20.2 ± 5.6 | 7.8 – 29.4 |  |  |
|  | cine-MRI | 24.2 ± 6.8 | 8.2 – 31.5 |  |  |
| SI: Superior-Inferior; SD: Standard Deviation. |  |  |  |  |  |
| B. Phantom study: Accuracy of motion amplitude capture across different patterns |  |  |  |  |  |
| Motion Pattern | Programmed Amplitude (mm, Ground Truth) | Measured Amplitude (mm, Mean ± SD) |  | Absolute Error (mm, Mean) |  |
|  |  | 4D-CT | cine-MRI | 4D-CT | cine-MRI |
| Sin-10A | 10 | 9.8 ± 0.3 | 10.1 ± 0.4 | 0.2 | 0.1 |
| Sin-25A | 25 | 24.5 ± 0.5 | 25.2 ± 0.6 | 0.5 | 0.2 |
| Cos4AP | 10.0 (SI) / 2.5 (AP) / 1.0 (LR) | 9.7 ± 0.4 | 10.3 ± 0.5 | 0.3 | 0.3 |
| Heart | 5.0 (SI) / 8.0 (AP) / 6.0 (LR) | 4.9 ± 0.6 | 5.2 ± 0.7 | 0.1 | 0.2 |
| Sawtooth | 10.0 (SI) / 2.5 (AP) / 1.0 (LR) | 9.6 ± 0.5 | 10.4 ± 0.6 | 0.4 | 0.4 |
| Overall Mean | - | 11.7 ± 6.2 | 12.2 ± 6.4 | 0.3 | 0.2 |
| SI: Superior-Inferior; AP: Anterior-Posterior; LR: Left-Right. The overall mean amplitude is calculated across all patterns and directions. The absolute error is the mean of the absolute differences between the measured and programmed amplitudes for each pattern. |  |  |  |  |  |